## Data Objects
PQS E006 DS01.1 Annex 1 Cold Chain Data Objects
Data Objects
| List of all required and optional data objects that must be generated and stored by EMS | [5 empty] | Logger | EMD
| N | Data Category | Data Object | Setting, Measurement or Logic | Description | Data Source | Required - Must record data object | Required - Must record data object Transmit only applicable for remote EMDs | Programmed, Time or Event Monitoring | Object ID | Unit of Measure | Data Type | Data Format (1) | Min Value | Max Value | Comments
| 1 | Administrative | Appliance manufacturer | Setting | Manufacturer name | Manufacturer input in logger | Required | Accept and transmit | Programmed | AMFR | N/A | String | Alphanumeric
| 2 | Administrative | Appliance model | Setting | Manufacturer model number | Manufacturer input in logger | Required | Accept and transmit | Programmed | AMOD | N/A | String | Alphanumeric
| 3 | Administrative | Appliance manufacturer serial number | Setting | Manufacturer serial number | Manufacturer input in logger | Required | Accept and transmit | Programmed | ASER | N/A | String | Alphanumeric
| 4 | Administrative | Appliance indentifier | Setting | Identifier for program asset tracking | Program input in database | N/A | Accept and transmit | Programmed | AID | N/A | String | Alphanumeric
| 5 | Administrative | Appliance date of production | Setting | Date of manufacture | Manufacturer input in logger | Required | Accept and transmit | Programmed | ADOP | YYYY-MM-DD | Date string | ISO 8601 | 1970-01-01 | 2958465
| 6 | Administrative | Appliance PQS code | Setting | PQS code (E00X/XXX) | Manufacturer input in logger | Required | Accept and transmit | Programmed | APQS | N/A | String | E0XX/XXX
| 7 | Administrative | Appliance PQS device type | Setting | PQS device type or category (e.g. RF03) | Manufacturer input in logger | Required | Accept and transmit | Programmed | ACAT | N/A | String | Alphanumeric |  |  | Device category, which E-EMDs can use to set alarm thresholds.
| 8 | Administrative | Logger manufacturer | Setting | Manufacturer name | Manufacturer input in logger | Required | Accept and transmit | Programmed | LMFR | N/A | String | Alphanumeric
| 9 | Administrative | Logger model | Setting | Manufacturer model number | Manufacturer input in logger | Required | Accept and transmit | Programmed | LMOD | N/A | String | Alphanumeric
| 10 | Administrative | Logger serial number | Setting | Manufacturer serial number | Manufacturer input in logger | Required | Accept and transmit | Programmed | LSER | N/A | String | Alphanumeric
| 11 | Administrative | Logger indentifier | Setting | Identifier for program asset tracking | Program input in database | Optional | Accept and transmit | Programmed | LID | N/A | String | Alphanumeric
| 12 | Administrative | Logger date of production | Setting | Date of manufacture | Manufacturer input in logger | Required | Accept and transmit | Programmed | LDOP | YYYY-MM-DD | Date string | ISO 8601 | 1970-01-01 | 2958465
| 13 | Administrative | Logger software version | Setting | Version number of Logger software/firmware installed | Software version on Logger | Required | Accept and transmit | Programmed | LSV | N/A | String | Alphanumeric
| 14 | Administrative | Logger PQS code | Setting | PQS code (E00X/XXX) | Manufacturer input in logger | Required | Accept and transmit | Programmed | LPQS | N/A | String | E0XX/XXX
| 15 | Administrative | EMD identifier | Setting | Identifier for program asset tracking | Program input in database | N/A | Optional | Programmed | EID | N/A | String | Alphanumeric
| 16 | Administrative | EMD manufacturer | Setting | Manufacturer name | Manufacturer input in EMD | N/A | Required | Programmed | EMFR | N/A | String | Alphanumeric
| 17 | Administrative | EMD model | Setting | Manufacturer model number | Manufacturer input in EMD | N/A | Required | Programmed | EMOD | N/A | String | Alphanumeric
| 18 | Administrative | EMD serial number | Setting | Manufacturer serial number | Manufacturer input in EMD | N/A | Required | Programmed | ESER | N/A | String | Alphanumeric
| 19 | Administrative | EMD date of production | Setting | Date of manufacture | Manufacturer input in EMD | N/A | Required | Programmed | EDOP | YYYY-MM-DD | Date string | ISO 8601 | 1970-01-01 | 2958465
| 20 | Administrative | EMD software version | Measurement | Version number of EMD software/firmware installed | Software version on EMD | N/A | Required | Programmed | EMSV | N/A | String | Alphanumeric
| 21 | Administrative | EMD PQS code | Setting | PQS code (E00X/XXX) | Manufacturer input in EMD | N/A | Required | Programmed | EPQS | N/A | String | E0XX/XXX
| 22 | Administrative | Compressor Electronic Unit Manufacturer | Setting | Compressor electronic unit admin information | Manufacturer input | Required, if applicable | Accept and transmit | Programmed | CNAM | N/A | String | Alphanumeric |  |  | Appliances may have more then one cooling system. This field refers to the refrigerator system in combined refrigerator/freezer appliances, the primary system in a dual-system cold room, and the first stage in a two-stage ultra-cold freezer.
| 23 | Administrative | Compressor Electronic Unit Product Code | Measurement | Compressor electronic unit admin information | Compressor electronic unit | Required, if applicable | Accept and transmit | Programmed | CSER | N/A | String | Alphanumeric |  |  | Appliances may have more then one cooling system. This field refers to the refrigerator system in combined refrigerator/freezer appliances, the primary system in a dual-system cold room, and the first stage in a two-stage ultra-cold freezer.
| 24 | Administrative | Compressor Electronic Unit Software version | Measurement | Compressor electronic unit admin information | Compressor electronic unit | Required, if applicable | Accept and transmit | Programmed | CSOF | N/A | String | Alphanumeric |  |  | Appliances may have more then one cooling system. This field refers to the refrigerator system in combined refrigerator/freezer appliances, the primary system in a dual-system cold room, and the first stage in a two-stage ultra-cold freezer.
| 25 | Administrative | Compressor Electronic Unit production date | Measurement | Compressor electronic unit admin information | Compressor electronic unit | Required, if applicable | Accept and transmit | Programmed | CDAT | YYYY-MM-DD | Date string | ISO 8601 | 1970-01-01 | 2958465 | Appliances may have more then one cooling system. This field refers to the refrigerator system in combined refrigerator/freezer appliances, the primary system in a dual-system cold room, and the first stage in a two-stage ultra-cold freezer.
| 26 | Administrative | Secondary Compressor Electronic Unit Manufacturer | Setting | Compressor electronic unit admin information | Manufacturer input | Required, if applicable | Accept and transmit | Programmed | CNAM2 | N/A | String | Alphanumeric |  |  | Appliances may have more then one cooling system. This field refers to the freezer system in combined refrigerator/freezer appliances, the secondary system in a dual-system cold room, and the second stage in a two-stage ultra-cold freezer.
| 27 | Administrative | Secondary Compressor Electronic Unit Product Code | Measurement | Compressor electronic unit admin information | Compressor electronic unit | Required, if applicable | Accept and transmit | Programmed | CSER2 | N/A | String | Alphanumeric |  |  | Appliances may have more then one cooling system. This field refers to the freezer system in combined refrigerator/freezer appliances, the secondary system in a dual-system cold room, and the second stage in a two-stage ultra-cold freezer.
| 28 | Administrative | Secondary Compressor Electronic Unit Software version | Measurement | Compressor electronic unit admin information | Compressor electronic unit | Required, if applicable | Accept and transmit | Programmed | CSOF2 | N/A | String | Alphanumeric |  |  | Appliances may have more then one cooling system. This field refers to the freezer system in combined refrigerator/freezer appliances, the secondary system in a dual-system cold room, and the second stage in a two-stage ultra-cold freezer.
| 29 | Administrative | Secondary Compressor Electronic Unit production date | Measurement | Compressor electronic unit admin information | Compressor electronic unit | Required, if applicable | Accept and transmit | Programmed | CDAT2 | YYYY-MM-DD | Date string | ISO 8601 | 1970-01-01 | 2958465 | Appliances may have more then one cooling system. This field refers to the freezer system in combined refrigerator/freezer appliances, the secondary system in a dual-system cold room, and the second stage in a two-stage ultra-cold freezer.
| 30 | Administrative | Relative Time | Measurement | Time since start of relative time keeping at PT0S | Logger clock | Required | Accept and transmit | N/A | RELT | PnDTnHnMnS | Alphanumeric | ISO 8601 Duration PnDTnHnMnS | PT0S | P9999D
| 31 | Administrative | RTC Wakeup | Measurement | Relative timestamp of when logger timer started its most recent continuous operation | Logger clock | Required | Accept and transmit | N/A | RTCW | PnDTnHnMnS | Alphanumeric | ISO 8601 Duration PnDTnHnMnS | PT0S | P9999D
| 32 | Administrative | Absolute time in UTC | Measurement | Absolute time measured in UTC | Synchronized server clocks | Optional | Required, for remote EMDs | N/A | ABST | YYYYMMDDThhmmssZ | Datetime string | ISO 8601
| 33 | Administrative | Region name |  | Name of region where device is located | Program input | N/A | N/A | Programmed in remote data systems | RNAM | N/A | String | Alphanumeric
| 34 | Administrative | District name |  | Name of district where device is located | Program input | N/A | N/A | Programmed in remote data systems | DNAM | N/A | String | Alphanumeric
| 35 | Administrative | Facility name |  | Name of facility where device is located | Program input | N/A | N/A | Programmed in remote data systems | FNAM | N/A | String | Alphanumeric
| 36 | Administrative | Country ID | Setting | Country where monitored appliance is located | Program input | N/A | N/A | Programmed | CID | N/A | String | Unique country code in ISO 3166-1 alpha-2 format |  |  | For EMD, the location data object must be available, but the source of this information may be either on-board GPS or manual input by programs through on-site or online programming of the EMD
| 37 | Administrative | Facility ID | Setting | Facility where monitored appliance is located | Program input | N/A | N/A | Programmed | FID | N/A | String | Alphanumeric
| 38 | Administrative | Location / GIS Coordinates / Latitude | Setting or Measurement | Physical location of asset | Sensor / Program input | N/A | Optional | Time | LAT | Degrees Decimal (D.DDDDD) | Numeric | ISO 6709 | -90.00000 | 90.00000 | 1 or 2 digits before decimal place. 5 digits after decimal gives about 1m resolution.
| 39 | Administrative | Location / GIS Coordinates / Longitude | Setting or Measurement | Physical location of asset | Sensor / Program input | N/A | Optional | Time | LNG | Degrees Decimal (D.DDDDD) | Numeric | ISO 6709 | -180.00000 | 180.00000 | 1 to 3 digits before decimal place. 5 digits after decimal gives about 1m resolution.
| 40 | Administrative | Location accuracy | Setting or Measurement | Accuracy of location information, in meters | Sensor / Program input | N/A | Optional | Time | LACC | meters | Numeric | 000.0 | 0 | 999.9
| 41 | Administrative | File integrity check | Logic | Integrity check on file contents, could be CRC, hash, or signature | Logger or EMD | Optional | Optional | Time | SIGN | N/A | String | alphanumeric characters |  |  | To be handled in a manufacturer-specific way.
| 42 | Environmental | AC supply voltage to the appliance | Measurement | Average AC supply voltage to appliance within each 15 minute period. Average value of raw samples collected at intervals not longer than 10 seconds. | Compressor electronic unit | Optional | Accept and transmit | Time | ACSV | V ac | Numeric | 000.0 | 0 | 600
| 43 | Environmental | AC supply voltage availability | Measurement | Data indicating the number of seconds within each 15-minute period when the AC voltage is within acceptable bounds to operate the appliance; | Compressor electronic unit | Required, for mains | Accept and transmit | Event | SVA | N/A | Numeric | 000 | 0 | 900 (max seconds of AC supply voltage availability in 15-min period))
| 44 | Environmental | DC supply voltage to the appliance | Measurement | Average DC supply voltage to appliance within each 15 minute period. Average value of raw samples collected at intervals not longer than 10 seconds. | Compressor electronic unit | Required, for solar | Accept and transmit | Time | DCSV | V dc | Numeric | 000.0 | 0 | 999.9
| 45 | Environmental | Ambient relative humidity | Measurement | RH% | Sensor | Optional | Optional | Time | HAMB | % | Numeric | 000.0 | 0 | 100
| 46 | Environmental | Ambient temperature | Measurement | Temperature of the immediate appliance surroundings; should not be affected by appliance operation | Sensor | Required | Optional | Time | TAMB | °C | Numeric | 00.0 | -55 | 60
| 47 | Performance | AC current drawn by the appliance | Measurement | Average AC current being drawn from supply by the operation of the appliance within each 15 minute period. Average value of raw samples collected at intervals not longer than 10 seconds. | Compressor electronic unit | Optional | Accept and transmit | Time | ACCD | A | Numeric | 00.00 | 0 | 50
| 48 | Performance | DC current drawn by the appliance | Measurement | Average DC current being drawn from supply by the operation of the appliance within each 15 minute period. Average value of raw samples collected at intervals not longer than 10 seconds. | Compressor electronic unit | Required, for solar | Accept and transmit | Time | DCCD | A | Numeric | 00.0 | 0 | 99.9
| 49 | Performance | Compartment relative humidity | Measurement | RH% | Sensor | Optional | Accept and transmit | Time | HCOM | % | Numeric | 000.0 | 0 | 100
| 50 | Performance | Condenser temperature | Measurement | Temperature of the condenser | Sensor | Optional | Accept and transmit | Time | TCON | °C | Numeric | 000.0 | -55 | 150 | Appliances may have more then one cooling system. This field refers to the refrigerator system in combined refrigerator/freezer appliances, the primary system in a dual-system cold room, and the first stage in a two-stage ultra-cold freezer.
| 51 | Performance | Secondary Condenser temperature | Measurement | Temperature of the condenser | Sensor | Optional | Accept and transmit | Time | TCON2 | °C | Numeric | 000.0 | -55 | 150 | Appliances may have more then one cooling system. This field refers to the freezer system in combined refrigerator/freezer appliances, the secondary system in a dual-system cold room, and the second stage in a two-stage ultra-cold freezer.
| 52 | Performance | Compressor speed | Measurement | Maximum operating speed of the compressor in revolutions per minute within each 15-minute period, based on samples collected at intervals not longer than 10 seconds. | Compressor electronic unit | Optional | Accept and transmit | Time | CMPS | rpm | Numeric | 00000 | 0 | 20000 | Appliances may have more then one cooling system. This field refers to the refrigerator system in combined refrigerator/freezer appliances, the primary system in a dual-system cold room, and the first stage in a two-stage ultra-cold freezer.
| 53 | Performance | Compressor runtime | Measurement | Total duration compressor operated within each 15 minute period, measured in minutes | Compressor electronic unit | Required, if applicable | Accept and transmit | Time | CMPR | min. | Numeric | 00 | 0 | 15 | Appliances may have more then one cooling system. This field refers to the refrigerator system in combined refrigerator/freezer appliances, the primary system in a dual-system cold room, and the first stage in a two-stage ultra-cold freezer.
| 54 | Performance | Compressor Electronic Unit temperature | Measurement | Temperature of Compressor Electronic Unit | Compressor electronic unit | Optional | Accept and transmit | Time | TPCB | °C | Numeric | 000.0 | -55 | 150 | Appliances may have more then one cooling system. This field refers to the refrigerator system in combined refrigerator/freezer appliances, the primary system in a dual-system cold room, and the first stage in a two-stage ultra-cold freezer.
| 55 | Performance | Secondary Compressor speed | Measurement | Maximum operating speed of the compressor in revolutions per minute within each 15-minute period, based on samples collected at intervals not longer than 10 seconds. | Compressor electronic unit | Optional | Accept and transmit | Time | CMPS2 | rpm | Numeric | 00000 | 0 | 20000 | Appliances may have more then one cooling system. This field refers to the freezer system in combined refrigerator/freezer appliances, the secondary system in a dual-system cold room, and the second stage in a two-stage ultra-cold freezer.
| 56 | Performance | Secondary Compressor runtime | Measurement | Total duration compressor operated within each 15 minute period, measured in minutes | Compressor electronic unit | Required, if applicable | Accept and transmit | Time | CMPR2 | min. | Numeric | 00 | 0 | 15 | Appliances may have more then one cooling system. This field refers to the freezer system in combined refrigerator/freezer appliances, the secondary system in a dual-system cold room, and the second stage in a two-stage ultra-cold freezer.
| 57 | Performance | Secondary Compressor Electronic Unit temperature | Measurement | Temperature of Compressor Electronic Unit | Compressor electronic unit | Optional | Accept and transmit | Time | TPCB2 | °C | Numeric | 000.0 | -55 | 150 | Appliances may have more then one cooling system. This field refers to the freezer system in combined refrigerator/freezer appliances, the secondary system in a dual-system cold room, and the second stage in a two-stage ultra-cold freezer.
| 58 | Performance | Fan speed | Measurement | Min. 0% / Max. 100% speed | Compressor electronic unit | Optional | Accept and transmit | Time | FANS | % | Numeric | 000.0 | 0 | 100
| 59 | Performance | Logger battery remaining | Measurement | Estimated number of days of battery life remaining | Logger | Required | Accept and transmit | Time | BLOG | DDDD | Numeric | 0000 | 0.0 | 9999
| 60 | Performance | Vaccine Compartment Temperature | Measurement | Temperature sensor located in coldest location of storage compartment | Sensor | Required | Accept and transmit | Time | TVC | °C | Numeric | 00.0 | -55 | 60
| 61 | Performance | Freezer Compartment Temperature | Measurement | Temperature sensor located in a freezer compartment | Sensor | Required, if applicable | Accept and transmit | Time | TFRZ | °C | Numeric | 00.0 | -99.9 | 60
| 62 | Performance | EMD battery remaining | Measurement | Estimated number of days of battery life remaining | EMD | N/A | Required | Time | BEMD | DDDD.D | Numeric | 000.0 | 0 | 9999.9
| 63 | Performance | Holdover, autonomy, or independence time | Measurement | Estimated time in days during which all points in the vaccine compartment remain between +2°C and +8°C if supply power disconnected | Sensors, data analysis | Optional | Optional | Time | HOLD | DDD.D | Numeric | 000.0 | 0.0 | 999.9
| 64 | Performance | Logger Error codes | Measurement | Alphanumeric codes corresponding to conditions that may impair normal operation of the Logger (e.g. broken or disconnected sensors, self-test failure) | Logger, various inputs | Required | Accept and transmit | Event | LERR | N/A | String | pre-defined error codes with 1 to 7 characters; if multiple error conditions exist, all relevant codes to be included as single string within data object, separated by spaces, e.g. "COMM PWR" |  |  | Mfr. may dictate code conventions, but must make a dictionary of codes and relevant explanation available to the Employer and their designated partners
| 65 | Performance | EMD Error codes | Measurement | Alphanumeric codes corresponding to conditions that may impair normal operation of the EMD (e.g. broken or disconnected sensors, self-test failure) | EMD, various inputs | N/A | Required | Event | EERR | N/A | String | pre-defined error codes with 1 to 7 characters; if multiple error conditions exist, all relevant codes to be included as single string within data object, separated by spaces, e.g. "COMM PWR" |  |  | Mfr. may dictate code conventions, but must make a dictionary of codes and relevant explanation available to the Employer and their designated partners
| 66 | Alarms | Alarm condition | Logic | Presence of defined alarm conditions | Logger | Optional | Required | Event | ALRM | N/A | String | pre-defined alarm codes with 4 to 7 characters; if multiple alarm conditions exist, all relevant codes to be included as single string within data object, separated by spaces, e.g. "HEAT DOOR" |  |  | Codes referenced in Data Standard section 4.2 and include: -Heat alarm (HEAT), acknowledged (HEATACK) -Freeze alarm (FRZE), acknowledged (FRZEACK) -Door open alarm (DOOR), acknowledged (DOORACK) -Power outage alarm (POWR), acknowledged (POWRACK) -EMD disconnect alarm (DCNT), acknowledged (DCNTACK)
| 67 | Use | Door/lid opening (vaccine compartment) | Measurement | Total duration vaccine compartment door was opened within each 15 minute period, measured in seconds | Sensor | Required | Accept and transmit | Event | DORV | seconds | Numeric | 000 | 0 | 900 (max seconds of door opening in 15-min period)
| 68 | Use | Instantaneous door/lid opening (vaccine compartment) | Measurement | Total continuous duration vaccine compartment door has been open at time of USB mount, measured in seconds | Sensor | Required in sync file | Accept and use for alarm | Event | IDRV | seconds | Numeric | 000 | 0 | 900 (max seconds of door opening in 15-min period)
| 69 | Use | Number of door/lid openings (vaccine compartment) | Measurement | Number of door openings within each 15-minute sample period | Sensor | Optional | Accept and transmit | Event | DRCV | N/A | Numeric | 000 | 0
| 70 | Use | Door/lid opening (freezer/other compartment) | Measurement | Total duration freezer compartment or other compartment door was opened within each 15 minute period, measured in seconds | Sensor | Required, if applicable | Accept and transmit | Event | DORF | seconds | Numeric | 000 | 0 | 900 (max seconds of door opening in 15-min period)
| 71 | Use | Instantaneous door/lid opening (freezer compartment) | Measurement | Total continuous duration freezer compartment door has been open at time of USB mount, measured in seconds | Sensor | Required in sync file | Accept and use for alarm | Event | IDRF | seconds | Numeric | 000 | 0 | 900 (max seconds of door opening in 15-min period)
| 72 | Use | Number of door/lid openings (freezer/other compartment) | Measurement | Number of door openings within each 15-minute sample period | Sensor | Optional | Accept and transmit | Event | DRCF | N/A | Numeric | 000 | 0
| 73 | Use | Main ON/OFF switch | Setting | OFF (false) / ON (true) of main control switch of appliance | Sensor or compressor electronic unit | Optional except required for refrigerated vehicles and transportable powered devices | Accept and transmit | Event | MSW | N/A | Boolean | true OR false | false | true
1. Data format column shows maximum expected number of digits before a decimal place. Leading zeroes are not needed, though a single leading zero is required if the value is < 1.0.

## Alarms (hidden)
Alarms |  | For discussion see: | Teams Discussion |  |  | Please use Teams for discussion; use this space for specific suggestions/comments
>>>Use these columns to populate specific comments
N | Alarm Code | Alarm Name | Definition | Source Data | Reference | I-WG Comments | I-WG Comments | I-WG Comments | I-WG Comments | I-WG Comments | I-WG Comments | I-WG Comments
1 | A0 | No alarm | No alarm conditions met | Various
2 | A1 | Heat alarm | Any vaccine compartment temperature sensor has recorded temperatures >8°C for 10 or more consecutive hours | Vaccine compartment temperature sensors / logger
3 | A2 | Freeze alarm | Any vaccine compartment temperature sensor has recorded temperatures below -0.5°C for 1 or more consecutive hours | Vaccine compartment temperature sensors / logger
4 | A3 | Door alarm | Door sensor has recorded 'OPEN' for X consecutive minutes | Door sensors / logger
5 | A4 | Power outage | Supply voltage availability' data registers the power as 'OFF' for X hours or more | Voltage sensor / compressor electronic unit
6 | A5 optional? | Clean Solar Array | Solar array power is insufficient to start the appliance and the cause is attributed to accumulated dust, dirt or debris on the solar array | Various: power supply, ambient and compartment temperatures, etc.
7
8
9 |  | Add others
10 |  | Add others
11 |  | Add others
12 |  | Add others
13 |  | Add others
14 |  | Add others
15 |  | Add others
16 |  | Add others
17 |  | Add others
18 |  | Add others
19 |  | Add others
20 |  | Add others
21 |  | Add others
OUT OF DATE - TO BE UPDATED IN LINE WITH SPEC

## KPIs (hidden)
KPIs |  | For discussion see: | Teams Discussion |  |  | Please use Teams for discussion; use this space for specific suggestions/comments
>>>Use these columns to populate specific comments
N | KPI code | KPI name | Definition | Source Data | Reference | I-WG Comments | I-WG Comments | I-WG Comments | I-WG Comments | I-WG Comments | I-WG Comments | I-WG Comments
1 | HOLD | Holdover time remaining | For an ILR, the estimated time remaining before the vaccine compartment is no longer expected to be within safe storage temperatures of +2-8C | Various: power supply, ambient and compartment temperatures, etc.
2 | AUTO | Autonomy time remaining | For an SDD, the estimated time remaining before the vaccine compartment is no longer expected to be within safe storage temperatures of +2-8C | Various: power supply, ambient and compartment temperatures, etc.
3 | UPTM | Appliance Uptime | % of time the Appliance's vaccine storage temperature has remained withing +2-8C for a user specified period of time | Vaccine Compartment Temperature Sensor - Cold
4 | CMPD | Compressor duty cycle | Hours of compressor operation per 24 hour cycle; possible to calculate based on user input of specified time period | Compressor Runtime (CMPR)
5 | AVS | Average storage temp
6 | STAT | Status
7 | MINT | Min. temperature
8 | TTBL | Total time below limit
9 | TTLA | Total low alarm time
10 | MAXT | Max. temperature
11 | TTAL | Total time above limit
12 | TTHA | Total high alarm time
13 | CUMD | Cumulative door open time
14 | PWRA | Power Availability | Time power was recorded as available during a given time period; represented as % of given time period
15 | CUMC | Cumulative compressor runtime
16 | MAXA | Max. ambient temperature
17 | NHEAT | Number of heat alarms in past 30 days
18
19
20
21
22
23
24
25
26
27
28
29
30
31
32
33
34
35
36
37
38
39
40
41
42
43
44
45
46
47
48
OUT OF DATE - TO BE UPDATED BASED ON INDUSTRY FEEDBACK AND LATEST SPEC

## Error Codes (hidden)
Error Codes |  | For discussion see: | Teams Discussion |  |  | Please use Teams for discussion; use this space for specific suggestions/comments
>>>Use these columns to populate specific comments
N | Error Code | Error Name | Source | Description | Reference | I-WG Comments | I-WG Comments | I-WG Comments | I-WG Comments | I-WG Comments | I-WG Comments | I-WG Comments
1 | 0 | No error | Compressor electronic unit |  | 1
2 | 1 | Voltage failure | Compressor electronic unit | The voltage is outside the cut-out setting | 1
3 | 2 | Fan failure | Compressor electronic unit | The rotor is blocked or the differential pressure in the refrigeration system is too high (>5 bar) | 1
4 | 3 | Motor failure | Compressor electronic unit | Too many compressor or fan starts in short time or fan current higher thhan 0.5A | 1
5 | 4 | Speed failure | Compressor electronic unit | If the refrigeration system is too heavily loaded, the motor cannot maintain minimum speed at approximately 1,850 rpm | 1
6 | 5 | Thermal failure | Compressor electronic unit | If the refrigeration system has been too heavily loaded, or if the ambient temperature is high, the electronic unit will run too hot. | 1
7 | 6 | NTC sensor failure | Compressor electronic unit | If the thermistor is short-circuit or has no connection | 1
8 | 7 | Communication error | Compressor electronic unit |  | 1
9 |  | Communication error | EMD | Communication connection lost with monitored appliance
10 |  | Add others
11 |  | Add others
12 |  | Add others
13 |  | Add others
14 |  | Add others
15 |  | Add others
16 |  | Add others
17 |  | Add others
18 |  | Add others
19 |  | Add others
20 |  | Add others
21 |  | Add others
References:
1 | SECOP - CONTROLLERS FOR BD COMPRESSORS - OPERATING INSTRUCTIONS (BD P-Housing)
2 | Instructions - Electronic Unit (Solar Applications) for BD35F/35K Compressors, 101N0420, 10-45V DC

## Taxonomy (hidden)
Equipment Monitoring Device | EMD
Equipment Monitoring System | EMS