## PPE SARI INTRO
COVID19 Community Facilities and Treatment centers
Orientative Medical Devices Bill of quantities
-NEEDED CONSUMABLES FOR X PATIENTS PER MONTH
-NEEDED MEDICAL EQUIPMENT FOR X PATIENTS TYPOLOGY PER FACILITY
-NEEDED ACCESORIES FOR X PATIENTS TYPOLOGY PER FACILITY
Based on Electro-mechanical medical equipment for NCoV Case Management WHE / OSL V 1.0
| PPE KIT, nCoV, 100 PATIENTS + 40 STAFF for 30 Days | Based on Severe Acute Respiratory Infections Treatment Centre
Practical manual to set up and manage a SARI treatment centre and a SARI screening facility in health care facilities - WHO 2020
[8 empty] | MODEL S 100B | MODEL L 180B | MODEL L 260B | MODEL L 340B | MODEL XL 420B | MODEL XL 500 B
| A1 (kit nCoV 100 patients) MODULE, PPE |  |  | US$ |  |  |  | TOTAL | TOTAL | TOTAL | TOTAL | TOTAL | TOTAL
WHO Code | WHO Description | Qtity | Unit Cost USO | Total cost (USD) | Estim. Unit Weight (kg) | Estim. Total Weight (kg) |  | 100 patients
YMEQGLASWSl--Al | GOGGLES PROTECTIVE, wraparound, soft frame, indirectvent. | 300 | 13 | 3900 | 0.86 | 258 | 3 | 300 | 540 | 780 | 1020 | 1260 | 1500
PEXTALCOlG---Al | ALCOHOL-BASED HAND RUB,gel, lOOm L, bottle | 60 | 1.29 | 77.4 | 0.12 | 7.2 | 0.6 | 60 | 108 | 156 | 204 | 252 | 300
EWASBAGBR007 -Al | BAG BIOHAZARD, REFUSE,AUTOCLAVABLE , 30x50cm, yellow | 100 | 0.35 | 35 | 0.005 | 0.5 | 1 | 100 | 180 | 260 | 340 | 420 | 500
EWASYCHN5Gl--Al | CHLORINE NaDCC, 45-55%,gran., lkg, pot | 8 | 6 | 48 | 1 | 8 | 0.08 | 8 | 14.4 | 20.8 | 27.2 | 33.6 | 40
CPPEGOWl3L---Al | GOWN,AAMI level 3, non sterile, disp., size L | 540 | 0.8 | 432 | 0.11467 | 61.9218 | 5.4 | 540 | 972 | 1404 | 1836 | 2268 | 2700
CPPEGOWl3M ---Al | GOWN,AAMI level 3, non sterile, disp., size M | 630 | 0.8 | 504 | 0.11467 | 72.2421 | 6.3 | 630 | 1134 | 1638 | 2142 | 2646 | 3150
CPPEGOWl3XL--Al | GOWN,AAMI level 3, non sterile, disp., size XL | 450 | 0.8 | 360 | 0.11467 | 51.6015 | 4.5 | 450 | 810 | 1170 | 1530 | 1890 | 2250
CPPEGOWl3XXL-Al | GOWN,AAMI level 3, non sterile, disp., size XXL | 180 | 0.8 | 144 | 0.11467 | 20.6406 | 1.8 | 180 | 324 | 468 | 612 | 756 | 900
CMSUGLENlLl--Al | GLOVE EXAMINATION, nitrile, pi, size L | 2200 | 0.07 | 154 | 0.00707 | 15.554 | 22 | 2200 | 3960 | 5720 | 7480 | 9240 | 11000
CMSUGLENlMl--Al | GLOVE EXAMINATION, nitrile, pi, size M | 4200 | 0.07 | 294 | 0.00707 | 29.694 | 42 | 4200 | 7560 | 10920 | 14280 | 17640 | 21000
CMSUGLENlSl--Al | GLOVE EXAMINATION, nitrile, pi, size S | 4200 | 0.07 | 294 | 0.00707 | 29.694 | 42 | 4200 | 7560 | 10920 | 14280 | 17640 | 21000
CMSUGLENlXL--Al | GLOVE EXAMINATION, nitrile, pi, size XL | 1600 | 0.07 | 112 | 0.00707 | 11.312 | 16 | 1600 | 2880 | 4160 | 5440 | 6720 | 8000
CPPEMASS2RL--Al | MASKSURGICAL, type llR, level 2, s.u, non sterile, earloop, size L | 1100 | 0.66 | 726 | 0.00421 | 4.631 | 11 | 1100 | 1980 | 2860 | 3740 | 4620 | 5500
CPPEMASS2RM--Al | MASKSURGICAL, type llR, level 2, s.u, non sterile, earloop, size M | 1100 | 0.66 | 726 | 0.00421 | 4.631 | 11 | 1100 | 1980 | 2860 | 3740 | 4620 | 5500
CPPEMASS2RS--Al | MASKSURGICAL,type llR, level 2, s.u, non sterile, earloop, size S | 1100 | 0.66 | 726 | 0.00421 | 4.631 | 11 | 1100 | 1980 | 2860 | 3740 | 4620 | 5500
CPPEMASPF205-Al | RESPIRATOR,mask, FFP2/N95, type llR, s.u., unvalved,nosedip | 6000 | 0.66 | 3960 | 0.00421 | 25.26 | 60 | 6000 | 10800 | 15600 | 20400 | 25200 | 30000
CPPEFSHIE002-Al | FACESHIELD,clear plastic, disp. | 2700 | 0.43 | 1161 | 0.01 | 27 | 27 | 2700 | 4860 | 7020 | 9180 | 11340 | 13500
CMSUTHERIOl--Al | THERMOMETER, INFRARED, no contact, handheld | 30 | 25 | 750 | 0.02 | 0.6 | 0.3 | 30 | 54 | 78 | 102 | 126 | 150
CINSCONTC51--Al | SAFETY BOX, needles/syringes,51, cardboard for incineration | 40 | 0.82 | 32.8 | 0.33 | 13.2 | 0.4 | 40 | 72 | 104 | 136 | 168 | 200
OPACUN62BS1--Al | BOX, triple packaging, biological substance UN3373 +pouch | 100 | 6.18 | 618 | 2 | 200 | 1 | 100 | 180 | 260 | 340 | 420 | 500
OPACUN621Sl--Al | BOX, triple packaging, infectious substance UN2814 | 20 | 30.28 | 605.6 | 2 | 40 | 0.2 | 20 | 36 | 52 | 68 | 84 | 100
CMSUBAGB4A04-Al | BAG BODY, 8handles, U-shaped zip, white,400 microns, adult, 230x100cm | 20 | 17.8 | 356 | 5 | 100 | 0.2 | 20 | 36 | 52 | 68 | 84 | 100
16015.8
NOTE:closed shoes are recommended, however, for low resourcessetting, it advisable to equip staff with scrubs and rubber boots (local purchases whenever possible).Underneath the recommended quantitiesfor 40 staffs per shift with 4 for shift, morning, afternoon, night and recovery. | [7 empty] | MODEL S 100B | MODEL L 180B | MODEL L 260B | MODEL L 340B | MODEL XL 420B | MODEL XL 500 B
| (kit nCoV 40 staffs/shift x 4 shift) MODULE, Uniform | [6 empty] | TOTAL | TOTAL | TOTAL | TOTAL | TOTAL | TOTAL
WHO Code | WHO Description | Qtlty | Unit Cost USD | Total cost (USD) | Estlm. Unit Weight (kg) |  |  | 50 staff per Swift + 3 swifts + 1 laundry | 80 staff per Swift + 3 swifts + 1 laundry | 112 staff per Swift + 3 swifts + 1 laundry | 142 staff per Swift + 3 swifts + 1 laundry | 172 staff per Swift + 3 swifts + 1 laundry | 200 staff per Swift + 3 swifts + 1 laundry
YPPESTUTROSS-Al | SET, TUNIC + TROUSERS SURGICAL. woven,reusable,green, size [S) | 40 | 10.45 | 418 | 0.541 |  | 0.25 | 50 | 80 | 112 | 142 | 172 | 200
YPPESTUTROSM-Al | SET, TUNIC+ TROUSERS SURGICAL. woven, reusable,green, size [M) | 70 | 10.45 | 731.5 | 0.541 |  | 0.4375 | 87.5 | 140 | 196 | 248.5 | 301 | 350
YPPESTUTROSL-Al | SET,TUNIC+ TROUSERS SURGICAL. woven,reusable,green, size [L] | 60 | 10.45 | 627 | 0.541 |  | 0.375 | 75 | 120 | 168 | 213 | 258 | 300
YPPESTUTROSXLAl | SET, TUNIC+ TROUSERS SURGICAL. woven, reusable,green, size [XL) | 30 | 10.45 | 313.5 | 0.541 |  | 0.1875 | 37.5 | 60 | 84 | 106.5 | 129 | 150
OUFBOOTW38--Al | BOOTS, rubber,size (38),dark color (green or black), pair | 25 | 6.1 | 152.5 | 1.437 |  | 0.15625 | 31.25 | 50 | 70 | 88.75 | 107.5 | 125
OUFBOOTW40--A1 | BOOTS, rubber,size (40),dark color (green or black), pair | 50 | 6.1 | 305 | 1.437 |  | 0.3125 | 62.5 | 100 | 140 | 177.5 | 215 | 250
OLIFBOOTW42--Al | BOOTS, rubber,size (42],dark color (green or black), pair | 40 | 6.1 | 244 | 1.437 |  | 0.25 | 50 | 80 | 112 | 142 | 172 | 200
OUFBOOTW44--Al | BOOTS, rubber, size (44], dark color (green or black), pair | 30 | 6.1 | 183 | 1.437 |  | 0.1875 | 37.5 | 60 | 84 | 106.5 | 129 | 150
OLIFBOOTW46--A1 | BOOTS, rubber,size (46],dark color (green or black), pair | 15 | 6.1 | 91.5 | 1.437 |  | 0.09375 | 18.75 | 30 | 42 | 53.25 | 64.5 | 75
3066
[8 empty] | MODEL S 100B | MODEL L 180B | MODEL L 260B | MODEL L 340B | MODEL XL 420B | MODEL XL 500 B
| (kit nCoV 10 Cleaning staffs/shift x 4 shift) MODULE, Uniform and PPE | [6 empty] | TOTAL | TOTAL | TOTAL | TOTAL | TOTAL | TOTAL
WHO Code | WHO Description | Qtlty | Unit Cost USD | Total cost (USD) | Estlm. Unit Weight (kg) |  |  | 12 cleaning staff per Swift + 3 swifts + 1 laundry | 14 cleaning staff per Swift + 3 swifts + 1 laundry | 16 cleaning staff per Swift + 3 swifts + 1 laundry | 18 cleaning staff per Swift + 3 swifts + 1 laundry | 20 cleaning staff per Swift + 3 swifts + 1 laundry | 22 cleaning staff per Swift + 3 swifts + 1 laundry
YPPESTUTROSS-Al | SET, TUNIC + TROUSERS SURGICAL. woven,reusable,green, size [S) | 10 | 10.45 | 104.5 | 0.541 | 5.41 | 0.25 | 12 | 14 | 16 | 18 | 20 | 22
YPPESTUTROSM-Al | SET, TUNIC+ TROUSERS SURGICAL. woven, reusable,green, size [M) | 17.5 | 10.45 | 182.875 | 0.541 | 9.4675 | 0.4375 | 21 | 24.5 | 28 | 31.5 | 35 | 38.5
YPPESTUTROSL-Al | SET,TUNIC+ TROUSERS SURGICAL. woven,reusable,green, size [L] | 15 | 10.45 | 156.75 | 0.541 | 8.115 | 0.375 | 18 | 21 | 24 | 27 | 30 | 33
YPPESTUTROSXLAl | SET, TUNIC+ TROUSERS SURGICAL. woven, reusable,green, size [XL) | 7.5 | 10.45 | 78.375 | 0.541 | 4.0575 | 0.1875 | 9 | 10.5 | 12 | 13.5 | 15 | 16.5
OUFBOOTW38--Al | BOOTS, rubber,size (38),dark color (green or black), pair | 6.25 | 6.1 | 38.125 | 1.437 | 8.9813 | 0.15625 | 7.5 | 8.75 | 10 | 11.25 | 12.5 | 13.75
OUFBOOTW40--A1 | BOOTS, rubber,size (40),dark color (green or black), pair | 12.5 | 6.1 | 76.25 | 1.437 | 17.9625 | 0.3125 | 15 | 17.5 | 20 | 22.5 | 25 | 27.5
OLIFBOOTW42--Al | BOOTS, rubber,size (42],dark color (green or black), pair | 10 | 6.1 | 61 | 1.437 | 14.37 | 0.25 | 12 | 14 | 16 | 18 | 20 | 22
OUFBOOTW44--Al | BOOTS, rubber, size (44], dark color (green or black), pair | 7.5 | 6.1 | 45.75 | 1.437 | 10.7775 | 0.1875 | 9 | 10.5 | 12 | 13.5 | 15 | 16.5
OLIFBOOTW46--A1 | BOOTS, rubber,size (46],dark color (green or black), pair | 3.75 | 6.1 | 22.875 | 1.437 | 5.3887 | 0.09375 | 4.5 | 5.25 | 6 | 6.75 | 7.5 | 8.25
CPPEMASPF205-Al | RESPIRATOR,mask, FFP2/N95, type llR, s.u., unvalved,nosedip | 40 | 0.66 | 26.4 | 0.00421 | 0.1684 | 1 | 48 | 56 | 64 | 72 | 80 | 88
YMEQGLASWSl--Al | GOGGLES PROTECTIVE, wraparound, soft frame, indirectvent. | 40 | 13 | 520 | 0.86 | 34.4 | 1 | 48 | 56 | 64 | 72 | 80 | 88
1312.9
Fill in the NEXT TAB " Calculator" YELLOW CELLS as corresponds
100
then click on the "Procurement List" TAB for the selected model
BACK TO MENU
1
21

## Calculator
WHE/OSL - v0.1 |  |  |  | Electro-mechanical medical equipment for NCoV Case Management | [6 empty] | Only fill yellow cells | [7 empty] | Electro-mechanical medical equipment for NCoV Case Management
[7 empty] | Suspected | Severe | ICU |  | Main equipment |  | * Consumables are quantified as unit on relation to total of suspected, confirmed and ICU cases estimated per month.
Location (site) | Country prepardness |  | Total No. of beds | 100 |  | % of total capacity for beds of | 0% | 80% | 20% |  | Consumables or accessories |  | * Quantities of spare accessories are calculated on relation to common of use of the equipment. | [5 empty] | 20
Date | 2020-03-10 |  | Period (Days) | 30 |  | Average number of days of stay | 0 | 21 | 14 |  | Not applicable |  | * The percentage of ICU beds is a "common agreement" in design of healthcare facilities (10 to 20% of total no. of beds). | [5 empty] | 30
Responsible |  |  | Nb site | 3 |  | No. of beds | 0 | 80 | 20 |  | Suspected cases: Triage and observation room / CENTRE |  | * The percentage for paediatric supplies is set at 20% (for which 5% is for neonate). | [5 empty] | 3
|  |  | Occupancy rates | 100% |  | Bed-Day | 0 | 2400 | 600 |  | Severe cases: IV fluids and Oxygen Therapy |  | * It is considered that the main equipment includes its set of accessories when purchased. | [5 empty] | 1
|  |  | Beds (Full) | 100 |  | Total cases (patients) |  | 114.2857 | 42.8571 |  | ICU cases: IV fluids and Mechanical Ventilation |  | * Be careful with accessories and consumables that are not brand related to be included in the first delivery; they should be calculated according to the health care facility needs. | [5 empty] | 20
* Oxygen Therapy devices may be chosen according to the required flow of oxygen (up to 15 L/min).
Medical Area | Medical Purpose |  | Type | Designation (EN) | Variable | Per bed | Per case | Bed-Day | Per equipment | Per centre (site) | No. of beds sharing equipment (for confirmed case only) | Ratio remarks | Suspected cases: Triage and observation room / CENTRE | Severe cases: IV fluids and Oxygen Therapy | ICU cases: IV fluids and Mechanical Ventilation | Total Quantity | Approx. Cost (USD) | Total Apprx. Cost (USD) | Designation (EN)
Severe | Monitoring | Patient Monitor | Equipment | MONITOR PATIENT, NIBP, w/o ECG , battery, trolley, +acc. | Per bed | 1 |  |  |  |  | 1 | Per bed. |  | 80 |  | 80 | 2740.8636 |  | MONITOR PATIENT, NIBP, w/o ECG , battery, trolley, +acc.
Severe | Oxygen therapy | Pulse Oximeter | Equipment | PULSE OXYMETER - FINGERTIP | Per bed | 0.8 |  |  |  |  | 1 | Per adult patient. It is not recommended for paediatric patients. |  | 64 |  | 64 | 25 |  | PULSE OXYMETER - FINGERTIP
Severe | Oxygen therapy | Concentrator O2 | Equipment | CONCENTRATOR O2 10L, 230V, 50 Hz + acc. | Per bed | 1 |  |  |  |  | 1 | Per 2 beds (Split between 2 px. maximum because the O2 flow should be minimum 3L (infants) and 6 L (adults)). |  | 80 |  | 80 | 1486.4391 |  | CONCENTRATOR O2 10L, 230V, 50 Hz + acc.
Severe | Oxygen therapy | Concentrator O2 | Accessories | (humidifier) HOSE CONNECTOR | Per equipment |  |  |  | 0.1 |  |  | Two spare accessories per 20 eqt. |  | 8 |  | 8 |  |  | (humidifier) HOSE CONNECTOR
Severe | Oxygen therapy | Concentrator O2 | Accessories | HUMIDIFIER, autoclavable | Per equipment |  |  |  | 0.1 |  |  | Two spare accessories per 20 eqt. |  | 8 |  | 8 | 13 |  | HUMIDIFIER, autoclavable
Severe | Oxygen therapy | Concentrator O2 | Accessories | (concentr.O2) FLOW SPLITTER pediat., 5 flowmeters 0-2L/min | Per equipment |  |  |  | 0.1 |  |  | Two spare accessories per 20 eqt. |  | 8 |  | 8 |  |  | (concentr.O2) FLOW SPLITTER pediat., 5 flowmeters 0-2L/min
Severe | Oxygen therapy | Concentrator O2 | Consumables | NASAL OXYGEN CANNULA, 2.1m, 2 prongs + tube, adult | Per case |  | 0.8 |  |  |  |  | Per month. According to use per px. |  | 91.4286 |  | 92 | 0.29654 |  | NASAL OXYGEN CANNULA, 2.1m, 2 prongs + tube, adult
Severe | Oxygen therapy | Concentrator O2 | Consumables | NASAL OXYGEN CANNULA, 2 prongs + tube, paediatric | Per case |  | 0.15 |  |  |  |  | Per month. According to use per px. |  | 17.1429 |  | 18 | 0.37748 |  | NASAL OXYGEN CANNULA, 2 prongs + tube, paediatric
Severe | Oxygen therapy | Concentrator O2 | Consumables | NASAL OXYGEN CANNULA, 2 prongs + tube, neonate | Per case |  | 0.05 |  |  |  |  | Per month. According to use per px. |  | 5.7143 |  | 6 | 1.678 |  | NASAL OXYGEN CANNULA, 2 prongs + tube, neonate
Severe | Oxygen therapy | Concentrator O2 | Consumables | Mask, oxygen, with connection tube, reservoir bag and valve, high-concentration, adult, non-sterile, single use | Per case |  | 0.8 |  |  |  |  | Per month. According to use per px. |  | 91.4286 |  | 92 |  |  | Mask, oxygen, with connection tube, reservoir bag and valve, high-concentration, adult, non-sterile, single use
Severe | Oxygen therapy | Concentrator O2 | Consumables | Mask, oxygen, with connection tube, reservoir bag and valve, high-concentration, paediatric, non-sterile, single use | Per case |  | 0.15 |  |  |  |  | Per month. According to use per px. |  | 17.1429 |  | 18 |  |  | Mask, oxygen, with connection tube, reservoir bag and valve, high-concentration, paediatric, non-sterile, single use
Severe | Oxygen therapy | Concentrator O2 | Consumables | Venturi Mask, with percent O2 Lock + 2.1 m tubing , adult | Per case |  | 0.8 |  |  |  |  | Per month. According to use per px. |  | 91.4286 |  | 92 |  |  | Venturi Mask, with percent O2 Lock + 2.1 m tubing , adult
Severe | Oxygen therapy | Concentrator O2 | Consumables | Venturi Mask, with percent O2 Lock + 2.1 m tubing , paediatric | Per case |  | 0.15 |  |  |  |  | Per month. According to use per px. |  | 17.1429 |  | 18 |  |  | Venturi Mask, with percent O2 Lock + 2.1 m tubing , paediatric
Severe | Oxygen therapy | Concentrator O2 | Consumables | Catheter, nasal, 8 Fr, 40 cm, with lateral eyes, sterile, single use | Per case |  | 0.8 |  |  |  |  | Per month. According to use per px. |  | 91.4286 |  | 92 |  |  | Catheter, nasal, 8 Fr, 40 cm, with lateral eyes, sterile, single use
Severe | Oxygen therapy | Concentrator O2 | Consumables | CONNECTOR, biconical, symmetric, ext. diam.7-11 mm, autoclav. | Per case |  | 1 |  |  |  |  | Used for oxygen source device. |  | 114.2857 |  | 115 | 0.5 |  | CONNECTOR, biconical, symmetric, ext. diam.7-11 mm, autoclav.
Severe | Oxygen therapy | Concentrator O2 | Consumables | TUBE, silicone, autoclavable, int. diam.5 mm, 25 m | Per case |  | 2 |  |  |  |  | Used for oxygen source device. According to use per px. Calculation is per meter. |  | 228.5714 |  | 229 | 28.1879 |  | TUBE, silicone, autoclavable, int. diam.5 mm, 25 m
Severe | Oxygen therapy | O2 Wall | Equipment | Flowmeter, Thorpe tube, for oxygen 0-15L/min | Per bed | 2 |  |  |  |  | 1 | For wall oxygen. |  | 160 |  | 160 |  |  | Flowmeter, Thorpe tube, for oxygen 0-15L/min
Severe | Oxygen therapy | CPAP, infant | Equipment | CPAP 10 machine, w/twin flowmeters | Per bed | 0.01 |  |  |  |  | 1 | For infant use. Up to 5 L/min O2 and 5 L/min medical air. |  | 0.8 |  | 1 |  |  | CPAP 10 machine, w/twin flowmeters
Severe | Oxygen therapy | CPAP, adult | Equipment | CPAP unit w/nasal tubing and mask for adult | Per bed | 0.08 |  |  |  |  | 1 | For adult use. To treat hypercarbic respiratory failure (few incidence: 10% of adult). |  | 6.4 |  | 7 |  |  | CPAP unit w/nasal tubing and mask for adult
Severe | Oxygen therapy | Suction Pump | Equipment | SUCTION PUMP, MECHANICAL + collection bottles | Per centre |  |  |  |  | 4 |  | For emergency cleaning of airways. |  | 12 |  | 12 | 400 |  | SUCTION PUMP, MECHANICAL + collection bottles
Severe | Oxygen therapy | Suction Pump | Consumables | TUBE, silicone, autoclavable, int. diam.8 mm, 10 m | Per case |  | 0.025 |  |  |  |  | According to use, but equipment may have intermittent use. |  | 2.8571 |  | 3 |  |  | TUBE, silicone, autoclavable, int. diam.8 mm, 10 m
Severe | Drug administration | Drop Counter | Equipment | ELECTRONIC DROP COUNTER, IV fluids infu. gravity monitor, alarm, batt.AA | Per bed | 1 |  |  |  |  | 1 | Per bed. |  | 80 |  | 80 |  |  | ELECTRONIC DROP COUNTER, IV fluids infu. gravity monitor, alarm, batt.AA
Severe | Blood Chemistry | Chemistry Analyser | Equipment | CLINICAL CHEMISTRY ANALYSER | Per centre |  |  |  |  | 1 |  | Per centre. One cartridge is aimed to perform a test panel for one patient. |  | 3 |  | 3 |  |  | CLINICAL CHEMISTRY ANALYSER
Severe | Blood Chemistry | Chemistry Analyser | Consumables | CARTRIDGE for Chemistry anaylser (lactic acidosis and hyperlactatemia) | Per case |  | 2 |  |  |  |  | Calculation of this cartridge to be higher than the second one. |  | 228.5714 |  | 229 |  |  | CARTRIDGE for Chemistry anaylser (lactic acidosis and hyperlactatemia)
Severe | Blood Chemistry | Chemistry Analyser | Consumables | CARTRIDGE ( metabolic status and renal function) | Per case |  | 1 |  |  |  |  | Calculation of this cartridge to be lower than the first one. |  | 114.2857 |  | 115 |  |  | CARTRIDGE ( metabolic status and renal function)
Severe | Blood Chemistry | Chemistry Analyser | Consumables | CONTROL,solution 1 | Per equipment |  |  |  | 2 |  |  | Two spare per equipment. |  | 6 |  | 6 |  |  | CONTROL,solution 1
Severe | Blood Chemistry | Chemistry Analyser | Consumables | CONTROL,solution 2 | Per equipment |  |  |  | 2 |  |  | Two spare per equipment. |  | 6 |  | 6 |  |  | CONTROL,solution 2
Severe | Blood Chemistry | Chemistry Analyser | Consumables | CONTROL,solution 3 | Per equipment |  |  |  | 2 |  |  | Two spare per equipment. |  | 6 |  | 6 |  |  | CONTROL,solution 3
Severe | Blood Chemistry | Chemistry Analyser | Consumables | TUBE, VACUUM, plastic, Li-HEPARIN, 2 mL green | Per case |  | 2 |  |  |  |  | Drug list? |  | 228.5714 |  | 229 |  |  | TUBE, VACUUM, plastic, Li-HEPARIN, 2 mL green
Severe | Blood Chemistry | Chemistry Analyser | Consumables | TUBE, VACUUM, plastic, Li-HEPARIN, 4 mL green | Per case |  | 2 |  |  |  |  | Drug list? |  | 228.5714 |  | 229 |  |  | TUBE, VACUUM, plastic, Li-HEPARIN, 4 mL green
ICU | Monitoring | Patient Monitor | Equipment | MONITOR PATIENT, multiparamet, ECG/CAPNO/SpO2/NIBP/Temp,230V,+acc | Per bed | 1 |  |  |  |  | 1 | Per ICU bed. |  |  | 20 | 20 | 9210.2386 |  | MONITOR PATIENT, multiparamet, ECG/CAPNO/SpO2/NIBP/Temp,230V,+acc
ICU | Monitoring | Patient Monitor | Consumables | (ECG) ELECTRODE GEL, bottle | Per case |  | 0.05 |  |  |  |  | Per month. According to use. 5% use of disposable electrode (optional). |  |  | 2.1429 | 3 |  |  | (ECG) ELECTRODE GEL, bottle
ICU | Oxygen therapy | Pulse Oximeter | Equipment | PULSE OXYMETER - FINGERTIP | Per bed | 0.8 |  |  |  |  | 1 | Per adult patient. It is not recommended for paediatric patients. |  |  | 16 | 16 | 25 |  | PULSE OXYMETER - FINGERTIP
ICU | Oxygen therapy | Concentrator O2 | Equipment | CONCENTRATOR O2 10L, 230V, 50 Hz + acc. | Per bed | 1 |  |  |  |  | 1 | Per ICU bed connected to patient ventilator. |  |  | 20 | 20 | 1486.4391 |  | CONCENTRATOR O2 10L, 230V, 50 Hz + acc.
ICU | Oxygen therapy | Concentrator O2 | Accessories | (humidifier) HOSE CONNECTOR 9/16 | Per equipment |  |  |  | 0.1 |  |  | Two spare accessories per 20 eqt. |  |  | 2 | 2 |  |  | (humidifier) HOSE CONNECTOR 9/16
ICU | Oxygen therapy | Concentrator O2 | Accessories | HUMIDIFIER, autoclavable | Per equipment |  |  |  | 0.1 |  |  | Two spare accessories per 20 eqt. |  |  | 2 | 2 | 13 |  | HUMIDIFIER, autoclavable
ICU | Oxygen therapy | Concentrator O2 | Accessories | (concentrator O2) FLOW SPLITTER pediat., 5 flowmeters 0-2L/min | Per equipment |  |  |  | 0.1 |  |  | Two spare accessories per 20 eqt. |  |  | 2 | 2 |  |  | (concentrator O2) FLOW SPLITTER pediat., 5 flowmeters 0-2L/min
ICU | Oxygen therapy | Concentrator O2 | Consumables | NASAL OXYGEN CANNULA, 2.1m, 2 prongs + tube, adult | Per case |  | 0.8 |  |  |  |  | Per month. According to use per px. |  |  | 34.2857 | 35 | 0.29654 |  | NASAL OXYGEN CANNULA, 2.1m, 2 prongs + tube, adult
ICU | Oxygen therapy | Concentrator O2 | Consumables | NASAL OXYGEN CANNULA, 2 prongs + tube, paediatric | Per case |  | 0.15 |  |  |  |  | Per month. According to use per px. |  |  | 6.4286 | 7 | 0.37748 |  | NASAL OXYGEN CANNULA, 2 prongs + tube, paediatric
ICU | Oxygen therapy | Concentrator O2 | Consumables | NASAL OXYGEN CANNULA, 2 prongs + tube, neonate | Per case |  | 0.05 |  |  |  |  | Per month. According to use per px. |  |  | 2.1429 | 3 | 1.678 |  | NASAL OXYGEN CANNULA, 2 prongs + tube, neonate
ICU | Oxygen therapy | Concentrator O2 | Consumables | Mask, oxygen, with connection tube, reservoir bag and valve, high-concentration, adult, non-sterile, single use | Per case |  | 1 | [7 empty] | 42.8571 | 43 |  |  | Mask, oxygen, with connection tube, reservoir bag and valve, high-concentration, adult, non-sterile, single use
ICU | Oxygen therapy | Concentrator O2 | Consumables | Mask, oxygen, with connection tube, reservoir bag and valve, high-concentration, paediatric, non-sterile, single use | Per case |  | 1 | [7 empty] | 42.8571 | 43 |  |  | Mask, oxygen, with connection tube, reservoir bag and valve, high-concentration, paediatric, non-sterile, single use
ICU | Oxygen therapy | Concentrator O2 | Consumables | Venturi Mask, with percent O2 Lock + 2.1 m tubing , adult | Per case |  | 1 | [7 empty] | 42.8571 | 43 |  |  | Venturi Mask, with percent O2 Lock + 2.1 m tubing , adult
ICU | Oxygen therapy | Concentrator O2 | Consumables | Venturi Mask, with percent O2 Lock + 2.1 m tubing , paediatric | Per case |  | 1 | [7 empty] | 42.8571 | 43 |  |  | Venturi Mask, with percent O2 Lock + 2.1 m tubing , paediatric
ICU | Oxygen therapy | Concentrator O2 | Consumables | Catheter, nasal, 8 Fr, 40 cm, with lateral eyes, sterile, single use | Per case |  | 1 | [7 empty] | 42.8571 | 43 |  |  | Catheter, nasal, 8 Fr, 40 cm, with lateral eyes, sterile, single use
ICU | Oxygen therapy | Concentrator O2 | Consumables | CONNECTOR, biconical, symmetric, ext. diam.7-11 mm, autoclav. | Per case |  | 1 |  |  |  |  | Used for oxygen source device. |  |  | 42.8571 | 43 |  |  | CONNECTOR, biconical, symmetric, ext. diam.7-11 mm, autoclav.
ICU | Oxygen therapy | Concentrator O2 | Consumables | TUBE, silicone, autoclavable, int. diam.5 mm, 25 m | Per case |  | 2 |  |  |  |  | Per month. According to use per px. Calculation is per meter. |  |  | 85.7143 | 86 | 28.1879 |  | TUBE, silicone, autoclavable, int. diam.5 mm, 25 m
ICU | Oxygen therapy | O2 Wall | Equipment | FLOWMETER, Thorpe tube, for oxygen 0-15L/min | Per bed | 2 |  |  |  |  | 1 | For wall oxygen. |  |  | 40 | 40 |  |  | FLOWMETER, Thorpe tube, for oxygen 0-15L/min
ICU | Airway Management | Independent | Equipment | CRICOTHYROTOMY, SET, emergency, 6 mm, ster., s.u., w/acc. | Per centre |  |  |  |  | 4 |  | Supplied with: 1 - 6.0 mm, cuffed cricothyroidotomy tube 1 - cricothyroidotomy, tube holder 1 - dilator 1 - scalpel blade, No. 15, for handle No. 3, sterile, single use 1 - suture, surgical, synthetic, non absorbable, monofilament, DEC 3.5 (0), 45 cm, with needle, 3/8 circle, 29.9 mm, cutting point 1 - syringe, 10 mL, two or three pieces, Luer type, sterile, single use 1 - Thermovent T (HME) |  |  | 12 | 12 |  |  | CRICOTHYROTOMY, SET, emergency, 6 mm, ster., s.u., w/acc.
ICU | Airway Management | Independent | Consumables | AIRWAY, NASOPHARYNGEAL, sterile, set, single use (from Fr20 to Fr36) | Per case |  | 1 | [7 empty] | 42.8571 | 43 |  |  | AIRWAY, NASOPHARYNGEAL, sterile, set, single use (from Fr20 to Fr36)
ICU | Airway Management | Independent | Consumables | AIRWAY, OROPHARYNGEAL, Guedel, sterile, set, single use (00, 0, 1, 2, 3, 4, 5) | Per case |  | 1 |  |  |  |  | Recommended to use sterile; but it exists the possibility of being reusable and autoclavable. |  |  | 42.8571 | 43 |  |  | AIRWAY, OROPHARYNGEAL, Guedel, sterile, set, single use (00, 0, 1, 2, 3, 4, 5)
ICU | Airway Management | Independent | Consumables | ENDOTRACHEAL TUBE INTRODUCER, Bougie, ster., s.u., Fr15, 70cm | Per case |  | 1 | [7 empty] | 42.8571 | 43 |  |  | ENDOTRACHEAL TUBE INTRODUCER, Bougie, ster., s.u., Fr15, 70cm
ICU | Airway Management | Independent | Consumables | ENDOTRACHEAL TUBE INTRODUCER, Bougie, ster., s.u., Fr10, 60cm | Per case |  | 1 | [7 empty] | 42.8571 | 43 |  |  | ENDOTRACHEAL TUBE INTRODUCER, Bougie, ster., s.u., Fr10, 60cm
ICU | Airway Management | Independent | Consumables | ENDOTRACHEAL TUBE INTRODUCER, Stylet, ster., s.u., Fr14, 30 to 45 cm | Per case |  | 1 | [7 empty] | 42.8571 | 43 |  |  | ENDOTRACHEAL TUBE INTRODUCER, Stylet, ster., s.u., Fr14, 30 to 45 cm
ICU | Airway Management | Independent | Consumables | ENDOTRACHEAL TUBE INTRODUCER, Stylet, ster., s.u., Fr10, 30 to 45 cm | Per case |  | 1 | [7 empty] | 42.8571 | 43 |  |  | ENDOTRACHEAL TUBE INTRODUCER, Stylet, ster., s.u., Fr10, 30 to 45 cm
ICU | Airway Management | Independent | Consumables | TUBE, ENDOTRACHEAL, No.2, w/o cuff, ster., s.u. | Per case |  | 1 | [7 empty] | 42.8571 | 43 |  |  | TUBE, ENDOTRACHEAL, No.2, w/o cuff, ster., s.u.
ICU | Airway Management | Independent | Consumables | TUBE, ENDOTRACHEAL, No.2.5, w/o cuff, ster., s.u. | Per case |  | 1 | [7 empty] | 42.8571 | 43 |  |  | TUBE, ENDOTRACHEAL, No.2.5, w/o cuff, ster., s.u.
ICU | Airway Management | Independent | Consumables | TUBE, ENDOTRACHEAL, No.3, w/o cuff, ster., s.u. | Per case |  | 1 | [7 empty] | 42.8571 | 43 |  |  | TUBE, ENDOTRACHEAL, No.3, w/o cuff, ster., s.u.
ICU | Airway Management | Independent | Consumables | TUBE, ENDOTRACHEAL, No.3.5, w/o cuff, ster., s.u. | Per case |  | 1 | [7 empty] | 42.8571 | 43 |  |  | TUBE, ENDOTRACHEAL, No.3.5, w/o cuff, ster., s.u.
ICU | Airway Management | Independent | Consumables | TUBE, ENDOTRACHEAL, No.4, w/o cuff, ster., s.u. | Per case |  | 1 | [7 empty] | 42.8571 | 43 |  |  | TUBE, ENDOTRACHEAL, No.4, w/o cuff, ster., s.u.
ICU | Airway Management | Independent | Consumables | TUBE, ENDOTRACHEAL, No.5, w/o cuff, ster., s.u. | Per case |  | 1 | [7 empty] | 42.8571 | 43 |  |  | TUBE, ENDOTRACHEAL, No.5, w/o cuff, ster., s.u.
ICU | Airway Management | Independent | Consumables | TUBE, ENDOTRACHEAL, No.4, w/ cuff, ster., s.u. | Per case |  | 1 | [7 empty] | 42.8571 | 43 |  |  | TUBE, ENDOTRACHEAL, No.4, w/ cuff, ster., s.u.
ICU | Airway Management | Independent | Consumables | TUBE, ENDOTRACHEAL, No.5, w/ cuff, ster., s.u. | Per case |  | 1 | [7 empty] | 42.8571 | 43 |  |  | TUBE, ENDOTRACHEAL, No.5, w/ cuff, ster., s.u.
ICU | Airway Management | Independent | Consumables | TUBE, ENDOTRACHEAL, No.6, w/ cuff, ster., s.u. | Per case |  | 1 | [7 empty] | 42.8571 | 43 |  |  | TUBE, ENDOTRACHEAL, No.6, w/ cuff, ster., s.u.
ICU | Airway Management | Independent | Consumables | TUBE, ENDOTRACHEAL, No.7, w/ cuff, ster., s.u. | Per case |  | 1 | [7 empty] | 42.8571 | 43 |  |  | TUBE, ENDOTRACHEAL, No.7, w/ cuff, ster., s.u.
ICU | Airway Management | Independent | Consumables | TUBE, ENDOTRACHEAL, No.8, w/ cuff, ster., s.u. | Per case |  | 1 | [7 empty] | 42.8571 | 43 |  |  | TUBE, ENDOTRACHEAL, No.8, w/ cuff, ster., s.u.
ICU | Airway Management | Independent | Consumables | TUBE, ENDOTRACHEAL, No.9, w/ cuff, ster., s.u. | Per case |  | 1 | [7 empty] | 42.8571 | 43 |  |  | TUBE, ENDOTRACHEAL, No.9, w/ cuff, ster., s.u.
ICU | Airway Management | Independent | Consumables | LARYNGEAL MASK AIRWAY (LMA), size 2, sterile, single use | Per case |  | 1 |  |  |  |  | Added as an option of endotracheal tube (ETT). |  |  | 42.8571 | 43 |  |  | LARYNGEAL MASK AIRWAY (LMA), size 2, sterile, single use
ICU | Airway Management | Independent | Consumables | LARYNGEAL MASK AIRWAY (LMA), size 3, sterile, single use | Per case |  | 1 |  |  |  |  | Added as an option of endotracheal tube (ETT). |  |  | 42.8571 | 43 |  |  | LARYNGEAL MASK AIRWAY (LMA), size 3, sterile, single use
ICU | Airway Management | Independent | Consumables | LARYNGEAL MASK AIRWAY (LMA), size 4, sterile, single use | Per case |  | 1 |  |  |  |  | Added as an option of endotracheal tube (ETT). |  |  | 42.8571 | 43 |  |  | LARYNGEAL MASK AIRWAY (LMA), size 4, sterile, single use
ICU | Airway Management | Independent | Consumables | COLORIMETRIC END TIDAL CO2 DETECTOR, adult, s.u. | Per case |  | 1 |  |  |  |  | Used to confirm proper ETT placement. |  |  | 42.8571 | 43 |  |  | COLORIMETRIC END TIDAL CO2 DETECTOR, adult, s.u.
ICU | Airway Management | Independent | Consumables | COLORIMETRIC END TIDAL CO2 DETECTOR, paediatric, s.u. | Per case |  | 1 |  |  |  |  | Used to confirm proper ETT placement. |  |  | 42.8571 | 43 |  |  | COLORIMETRIC END TIDAL CO2 DETECTOR, paediatric, s.u.
ICU | Airway Management | Independent | Consumables | SYRINGE, Luer slip, 10mL, sterile, single use | Per case |  | 1 |  |  |  |  | Used to inflate the ETT cuff. |  |  | 42.8571 | 43 | 0.8567 |  | SYRINGE, Luer slip, 10mL, sterile, single use
ICU | Airway Management | Independent | Consumables | LUBRICATING Jelly, 5g, box of 150 | Per case |  | 1 | [7 empty] | 42.8571 | 43 |  |  | LUBRICATING Jelly, 5g, box of 150
ICU | Airway Management | Independent | Consumables | FORCEPS, MAGILL, paediatric, 15 cm | Per case |  | 1 | [7 empty] | 42.8571 | 43 |  |  | FORCEPS, MAGILL, paediatric, 15 cm
ICU | Airway Management | Independent | Consumables | FORCEPS, MAGILL, child, 19 cm | Per case |  | 1 | [7 empty] | 42.8571 | 43 |  |  | FORCEPS, MAGILL, child, 19 cm
ICU | Airway Management | Independent | Consumables | FORCEPS, MAGILL, adult, 24 cm | Per case |  | 1 | [7 empty] | 42.8571 | 43 |  |  | FORCEPS, MAGILL, adult, 24 cm
ICU | Airway Management | Laryngoscope, adult | Equipment | LARYNGOSCOPE,fib.opt, ad/ch, diam.28mm, blades(Macintosh2/3/4+Miller1), sp.bulbs | Per centre |  |  |  |  | 8 |  | Supplied with: * Blades, Macintosh type (curved): No. 2, length 90 - 110 mm, for child. No. 3, length 110 - 135 mm, for small adult. No. 4, length 135 - 155 mm, for adult. * Blades, Miller type (straight): No. 1, length 100 mm. * Heavy-walled plastic or metal case. * Six compatible batteries in total. * Four extra halogen bulbs. |  |  | 24 | 24 | 675 |  | LARYNGOSCOPE,fib.opt, ad/ch, diam.28mm, blades(Macintosh2/3/4+Miller1), sp.bulbs
ICU | Airway Management | Laryngoscope, infant | Equipment | LARYNGOSCOPE,fib.opt, neonate, diam.19mm, blades(Macintosh0/1/2), sp.bulbs,batt. | Per centre |  |  |  |  | 4 |  | Supplied with: * Blades, Macintosh type (curved): No. 0, length 55 mm, for newborn. No. 1, length 70 mm, for infant. No. 2, length 90 mm, for child. * Heavy-walled plastic or metal case. * Six compatible batteries in total. * Four extra halogen bulbs. |  |  | 12 | 12 | 600 |  | LARYNGOSCOPE,fib.opt, neonate, diam.19mm, blades(Macintosh0/1/2), sp.bulbs,batt.
ICU | Mechanical ventilation | Resuscitator, adult | Equipment | SELF-INFLATING BAG, ad./child + masks RH5 / RH2 (Ambu type) | Per centre |  |  |  |  | 4 |  |  |  |  | 12 | 12 | 80 |  | SELF-INFLATING BAG, ad./child + masks RH5 / RH2 (Ambu type)
ICU | Mechanical ventilation | Resuscitator, infant | Equipment | SELF-INFLATING BAG, child/neonate + masks RH2/S1 (Ambu type) | Per centre |  |  |  |  | 2 |  |  |  |  | 6 | 6 | 97 |  | SELF-INFLATING BAG, child/neonate + masks RH2/S1 (Ambu type)
ICU | Mechanical ventilation | Suction device | Equipment | SUCTION BULB, for newborn, reusable, autoclavable | Per centre |  |  |  |  | 2 |  |  |  |  | 6 | 6 |  |  | SUCTION BULB, for newborn, reusable, autoclavable
ICU | Oxygen therapy | Suction Pump | Equipment | SUCTION PUMP, MECHANICAL + collection bottles | Per centre |  |  |  |  | 4 |  | For emergency cleaning of airways. |  |  | 12 | 12 | 400 |  | SUCTION PUMP, MECHANICAL + collection bottles
ICU | Oxygen therapy | Suction Pump | Consumables | TUBE, silicone, autoclavable, int. diam.8 mm, 10 m | Per case |  | 0.025 |  |  |  |  | According to use, but equipment may have intermittent use. |  |  | 1.0714 | 2 |  |  | TUBE, silicone, autoclavable, int. diam.8 mm, 10 m
ICU | Mechanical ventilation | Independent | Consumables | FILTER, HEAT AND MOISTURE EXCHANGER (HMEF), high efficiency, with connectors, adult, single use | Per case |  | 0.8 |  |  |  |  | Normally for long term ventilated patients. |  |  | 34.2857 | 35 |  |  | FILTER, HEAT AND MOISTURE EXCHANGER (HMEF), high efficiency, with connectors, adult, single use
ICU | Mechanical ventilation | Independent | Consumables | FILTER, HEAT AND MOISTURE EXCHANGER (HMEF), high efficiency, with connectors, paediatric, single use | Per case |  | 0.2 |  |  |  |  | Normally for long term ventilated patients. |  |  | 8.5714 | 9 |  |  | FILTER, HEAT AND MOISTURE EXCHANGER (HMEF), high efficiency, with connectors, paediatric, single use
ICU | Gastro-enteral Feeding | Independent | Consumables | TUBE, FEEDING, NASOGASTRIC, 10 Fr, 50 cm, ENFit tip, sterile, single use | Per case |  | 1 | [7 empty] | 42.8571 | 43 |  |  | TUBE, FEEDING, NASOGASTRIC, 10 Fr, 50 cm, ENFit tip, sterile, single use
ICU | Gastro-enteral Feeding | Independent | Consumables | TUBE, FEEDING, NASOGASTRIC, 12 Fr, 90 cm, ENFit tip, sterile, single use | Per case |  | 1 | [7 empty] | 42.8571 | 43 |  |  | TUBE, FEEDING, NASOGASTRIC, 12 Fr, 90 cm, ENFit tip, sterile, single use
ICU | Gastro-enteral Feeding | Independent | Consumables | TUBE, FEEDING, NASOGASTRIC, 14 Fr, 90 cm, ENFit tip, sterile, single use | Per case |  | 1 | [7 empty] | 42.8571 | 43 |  |  | TUBE, FEEDING, NASOGASTRIC, 14 Fr, 90 cm, ENFit tip, sterile, single use
ICU | Gastro-enteral Feeding | Independent | Consumables | TUBE, FEEDING, NASOGASTRIC, 6 Fr, 50 cm, ENFit tip, sterile, single use | Per case |  | 1 | [7 empty] | 42.8571 | 43 |  |  | TUBE, FEEDING, NASOGASTRIC, 6 Fr, 50 cm, ENFit tip, sterile, single use
ICU | Gastro-enteral Feeding | Independent | Consumables | TUBE, FEEDING, NASOGASTRIC, 8 Fr, 50 cm, ENFit tip, sterile, single use | Per case |  | 1 | [7 empty] | 42.8571 | 43 |  |  | TUBE, FEEDING, NASOGASTRIC, 8 Fr, 50 cm, ENFit tip, sterile, single use
ICU | Gastro-enteral Feeding | Independent | Consumables | SYRINGE, FEEDING, 1 mL, LDT, ENFit, sterile, single use | Per case |  | 1 | [7 empty] | 42.8571 | 43 |  |  | SYRINGE, FEEDING, 1 mL, LDT, ENFit, sterile, single use
ICU | Gastro-enteral Feeding | Independent | Consumables | SYRINGE, FEEDING, 10 mL, ENFit, sterile, single use | Per case |  | 1 | [7 empty] | 42.8571 | 43 |  |  | SYRINGE, FEEDING, 10 mL, ENFit, sterile, single use
ICU | Gastro-enteral Feeding | Independent | Consumables | SYRINGE, FEEDING, 2.5 mL, LDT, ENFit, sterile, single use | Per case |  | 1 | [7 empty] | 42.8571 | 43 |  |  | SYRINGE, FEEDING, 2.5 mL, LDT, ENFit, sterile, single use
ICU | Gastro-enteral Feeding | Independent | Consumables | SYRINGE, FEEDING, 20 mL, ENFit, sterile, single use | Per case |  | 1 | [7 empty] | 42.8571 | 43 |  |  | SYRINGE, FEEDING, 20 mL, ENFit, sterile, single use
ICU | Gastro-enteral Feeding | Independent | Consumables | SYRINGE, FEEDING, 5 mL, LDT, ENFit, sterile, single use | Per case |  | 1 | [7 empty] | 42.8571 | 43 |  |  | SYRINGE, FEEDING, 5 mL, LDT, ENFit, sterile, single use
ICU | Gastro-enteral Feeding | Independent | Consumables | SYRINGE, FEEDING, 60 mL, ENFit, sterile, single use | Per case |  | 1 | [7 empty] | 42.8571 | 43 |  |  | SYRINGE, FEEDING, 60 mL, ENFit, sterile, single use
ICU | Gastro-enteral Feeding | Independent | Consumables | LUBRICATING Jelly, 50g, tube | Per bed | 1 | [5 empty] | Used for the allocation of the nasogastric feeding tube. |  |  | 20 | 20 |  |  | LUBRICATING Jelly, 50g, tube
ICU | Gastro-enteral Feeding | Independent | Consumables | PAD, absorbent | Per case |  | 1 |  |  |  |  | Used for the allocation of the nasogastric feeding tube. |  |  | 42.8571 | 43 |  |  | PAD, absorbent
ICU | Gastro-enteral Feeding | Independent | Consumables | BASIN KIDNEY, stainless steel, 825 mL | Per bed | 1 | [5 empty] | Used for the allocation of the nasogastric feeding tube. |  |  | 20 | 20 |  |  | BASIN KIDNEY, stainless steel, 825 mL
ICU | Gastro-enteral Feeding | Independent | Consumables | STETHOSCOPE, binaural, double cup, adult/child, single use | Per bed | 1 | [5 empty] | Used for the allocation of the nasogastric feeding tube. |  |  | 20 | 20 |  |  | STETHOSCOPE, binaural, double cup, adult/child, single use
ICU | Central line | Independent | Consumables | CENTRAL VENOUS CATHETERS KIT | Per case |  | 1 |  |  |  |  | Supplied with: Finder needle, syringe, wire, dilator, lidocaine, scapel, needle, thread (e.g., ARROW) |  |  | 42.8571 | 43 |  |  | CENTRAL VENOUS CATHETERS KIT
ICU | Central line | Independent | Consumables | Transparent adhesive plasters, washproof, 5x5 cm | Per case |  | 3 |  |  |  |  | (e.g., Tegaderm) |  |  | 128.5714 | 129 |  |  | Transparent adhesive plasters, washproof, 5x5 cm
ICU | Urine collection | Independent | Consumables | Bag, collecting, urine, with outlet tap, with non-return valve, 2000 mL, adult, non-sterile, single use | Per case |  | 1 | [7 empty] | 42.8571 | 43 |  |  | Bag, collecting, urine, with outlet tap, with non-return valve, 2000 mL, adult, non-sterile, single use
ICU | Urine collection | Independent | Consumables | CATHETER, URETHRAL, Foley, 2-way, 10 Fr, sterile, single use | Per case |  | 1 | [7 empty] | 42.8571 | 43 |  |  | CATHETER, URETHRAL, Foley, 2-way, 10 Fr, sterile, single use
ICU | Urine collection | Independent | Consumables | CATHETER, URETHRAL, Foley, 2-way, 12 Fr, sterile, single use | Per case |  | 1 | [7 empty] | 42.8571 | 43 |  |  | CATHETER, URETHRAL, Foley, 2-way, 12 Fr, sterile, single use
ICU | Urine collection | Independent | Consumables | CATHETER, URETHRAL, Foley, 2-way, 14 Fr, sterile, single use | Per case |  | 1 | [7 empty] | 42.8571 | 43 |  |  | CATHETER, URETHRAL, Foley, 2-way, 14 Fr, sterile, single use
ICU | Urine collection | Independent | Consumables | CATHETER, URETHRAL, Foley, 2-way, 16 Fr, sterile, single use | Per case |  | 1 | [7 empty] | 42.8571 | 43 |  |  | CATHETER, URETHRAL, Foley, 2-way, 16 Fr, sterile, single use
ICU | Urine collection | Independent | Consumables | CATHETER, URETHRAL, Foley, 2-way, 18 Fr, sterile, single use | Per case |  | 1 | [7 empty] | 42.8571 | 43 |  |  | CATHETER, URETHRAL, Foley, 2-way, 18 Fr, sterile, single use
ICU | Urine collection | Independent | Consumables | CATHETER, URETHRAL, Foley, 2-way, 20 Fr, sterile, single use | Per case |  | 1 | [7 empty] | 42.8571 | 43 |  |  | CATHETER, URETHRAL, Foley, 2-way, 20 Fr, sterile, single use
ICU | Urine collection | Independent | Consumables | CATHETER, URETHRAL, Foley, 2-way, 8 Fr, sterile, single use | Per case |  | 1 | [7 empty] | 42.8571 | 43 |  |  | CATHETER, URETHRAL, Foley, 2-way, 8 Fr, sterile, single use
ICU | General supplies | Independent | Consumables | Compress, gauze, 10 x 10 cm, 8 to 12 ply, sterile, single use | Per case |  | 1 | [7 empty] | 42.8571 | 43 |  |  | Compress, gauze, 10 x 10 cm, 8 to 12 ply, sterile, single use
ICU | General supplies | Independent | Consumables | Tape, surgical, hypoallergenic, 5 x 2.5 cm | Per case |  | 1 | [7 empty] | 42.8571 | 43 |  |  | Tape, surgical, hypoallergenic, 5 x 2.5 cm
ICU | General supplies | Independent | Consumables | Drape, surgical, non woven, sterile, single use | Per case |  | 1 | [7 empty] | 42.8571 | 43 |  |  | Drape, surgical, non woven, sterile, single use
ICU | General supplies | Independent | Consumables | Gloves, examination, nitrile, powder-free, pair-packed, large, sterile, single use | Per case |  | 1 |  |  |  |  | Sterile gloves. |  |  | 42.8571 | 43 |  |  | Gloves, examination, nitrile, powder-free, pair-packed, large, sterile, single use
ICU | General supplies | Independent | Consumables | Gloves, examination, nitrile, powder-free, pair-packed, medium, sterile, single use | Per case |  | 1 |  |  |  |  | Sterile gloves. |  |  | 42.8571 | 43 |  |  | Gloves, examination, nitrile, powder-free, pair-packed, medium, sterile, single use
ICU | General supplies | Independent | Consumables | Gloves, examination, nitrile, powder-free, pair-packed, small, sterile, single use | Per case |  | 1 |  |  |  |  | Sterile gloves. |  |  | 42.8571 | 43 |  |  | Gloves, examination, nitrile, powder-free, pair-packed, small, sterile, single use
ICU | General supplies | Independent | Consumables | Antiseptic Wipe with Alcohol & Chlorhexidine | Per case |  | 1 |  |  |  |  | Drug list? It is needed for many procedures in this list. |  |  | 42.8571 | 43 |  |  | Antiseptic Wipe with Alcohol & Chlorhexidine
ICU | Blood Chemistry | Independent | Consumables | Arterial Blood Sample Kits | Per case |  | 1 |  |  |  |  | Arterial blood sampling syringes with dry lithium heparin [23.5 iu/mL] |  |  | 42.8571 | 43 |  |  | Arterial Blood Sample Kits
ICU | Mechanical ventilation | Patient Ventilator | Equipment | VENTILATOR PATIENT, for adult, paediatric and neonate w/acc. | Per bed | 0.5 |  |  |  |  | 1 | Per ICU bed. It includes the CO2 detector. |  |  | 10 | 10 | 35000 |  | VENTILATOR PATIENT, for adult, paediatric and neonate w/acc.
ICU | Mechanical ventilation | Patient Ventilator | Consumables | Breathing circuit (tubes/balloon/valves/mask), ADULT, single use | Per case |  | 0.8 |  |  |  |  | Per month. According to use per px. Consider to change to reusable. |  |  | 34.2857 | 35 |  |  | Breathing circuit (tubes/balloon/valves/mask), ADULT, single use
ICU | Mechanical ventilation | Patient Ventilator | Consumables | Breathing circuit (tubes/balloon/valves/mask), PAEDIATRIC, single use | Per case |  | 0.15 |  |  |  |  | Per month. According to use per px. Consider to change to reusable. |  |  | 6.4286 | 7 |  |  | Breathing circuit (tubes/balloon/valves/mask), PAEDIATRIC, single use
ICU | Mechanical ventilation | Patient Ventilator | Consumables | Breathing circuit (tubes/balloon/valves/mask), NEONATE, single use | Per case |  | 0.05 |  |  |  |  | Per month. According to use per px. Consider to change to reusable. |  |  | 2.1429 | 3 |  |  | Breathing circuit (tubes/balloon/valves/mask), NEONATE, single use
ICU | Oxygen therapy | HFNC | Equipment | High Flow Nasal Cannula (HFNC) | Per bed | 0.5 |  |  |  |  | 1 | Per ICU bed. |  |  | 10 | 10 |  |  | High Flow Nasal Cannula (HFNC)
ICU | Oxygen therapy | HFNC | Consumables | Optiflow + Nasal Cannula, small, pack of 20 | Per case |  | 1 | [7 empty] | 42.8571 | 43 |  |  | Optiflow + Nasal Cannula, small, pack of 20
ICU | Oxygen therapy | HFNC | Consumables | Optiflow + Nasal Cannula, medium, pack of 20 | Per case |  | 1 | [7 empty] | 42.8571 | 43 |  |  | Optiflow + Nasal Cannula, medium, pack of 20
ICU | Oxygen therapy | HFNC | Consumables | Optiflow + Nasal Cannula, large, pack of 20 | Per case |  | 1 | [7 empty] | 42.8571 | 43 |  |  | Optiflow + Nasal Cannula, large, pack of 20
ICU | Oxygen therapy | HFNC | Consumables | Disinfection kit | Per case |  | 0.8 |  |  |  |  | Per month. According to use per px. |  |  | 34.2857 | 35 |  |  | Disinfection kit
ICU | Oxygen therapy | HFNC | Consumables | Cleaning Sponge Stick, pack of 20 | Per case |  | 0.15 |  |  |  |  | Per month. According to use per px. |  |  | 6.4286 | 7 |  |  | Cleaning Sponge Stick, pack of 20
ICU | Oxygen therapy | HFNC | Consumables | Disinfection filter, pack of 2 | Per case |  | 0.05 |  |  |  |  | Per month. According to use per px. |  |  | 2.1429 | 3 |  |  | Disinfection filter, pack of 2
ICU | Oxygen therapy | Suction Pump | Equipment | SUCTION PUMP, ELECTRICAL, 100-230V,50-60Hz | Per bed | 1 |  |  |  |  | 1 | Per ICU bed. |  |  | 20 | 20 |  |  | SUCTION PUMP, ELECTRICAL, 100-230V,50-60Hz
ICU | Oxygen therapy | Suction Pump | Consumables | BACTERIA FILTER, unit | Per case |  | 0.5 |  |  |  |  | Per month. According to packaging and use. Calculation is to change every 2 px. |  |  | 21.4286 | 22 |  |  | BACTERIA FILTER, unit
ICU | Oxygen therapy | Suction Pump | Consumables | TUBE, silicone, autoclavable, int. diam.8 mm, 10 m | Per case |  | 1 |  |  |  |  | Per month. According to use per px. Calculation is per meter. |  |  | 42.8571 | 43 |  |  | TUBE, silicone, autoclavable, int. diam.8 mm, 10 m
ICU | Drug administration | Infusion Pump | Equipment | INFUSION PUMP | Per bed | 1 |  |  |  |  | 1 | Per ICU bed. To consider the use of more than one channel (This one has a flow range of 0.1-1200 L. Other pumps could manage up 1500 L). |  |  | 20 | 20 | 1470.9942 |  | INFUSION PUMP
ICU | Drug administration | Infusion Pump | Consumables | (inf. pump) INFUSION LINE | Per case |  | 2 |  |  |  |  | Per month. According to use per px. Calculation is 2 lines per px. |  |  | 85.7143 | 86 | 0.85182 |  | (inf. pump) INFUSION LINE
ICU | Drug administration | Drill | Equipment | DRILL, FOR VASCULAR ACCESS, with transport bag | Per centre |  |  |  |  | 1 |  |  |  |  | 3 | 3 |  |  | DRILL, FOR VASCULAR ACCESS, with transport bag
ICU | Drug administration | Drill | Consumables | NEEDLE + STABILIZER KIT, 15 G, 15 mm, PAEDIATRIC | Per case |  | 0.1 |  |  |  |  | 50 % of the ICU infants. According to use, but equipment may have intermittent use. |  |  | 4.2857 | 5 |  |  | NEEDLE + STABILIZER KIT, 15 G, 15 mm, PAEDIATRIC
ICU | Drug administration | Drill | Consumables | NEEDLE + STABILIZER KIT, 15 G, 25 mm, ADULT | Per case |  | 0.08 |  |  |  |  | 10 % of the ICU adults. According to use, but equipment may have intermittent use. |  |  | 3.4286 | 4 |  |  | NEEDLE + STABILIZER KIT, 15 G, 25 mm, ADULT
ICU | Monitoring | Defibrillator | Equipment | DEFIBRILLATOR, mobile, semi-auto.,multi-paramet,AC/DC, w/acc+trolley | Per centre |  |  |  |  | 1 |  |  |  |  | 3 | 3 | 4040.9124 |  | DEFIBRILLATOR, mobile, semi-auto.,multi-paramet,AC/DC, w/acc+trolley
ICU | Monitoring | Defibrillator | Consumables | (defibrilator) ELECTRODE PADS, adult, adhesive, disp. | Per case |  | 0.08 |  |  |  |  | 5 % of CPR. According to use, but equipment may have intermittent use. |  |  | 3.4286 | 4 |  |  | (defibrilator) ELECTRODE PADS, adult, adhesive, disp.
ICU | Monitoring | Defibrillator | Consumables | (defibrilator) ELECTRODE PADS, paediat., adhesive, disp. | Per case |  | 0.02 |  |  |  |  | 5 % of CPR. According to use, but equipment may have intermittent use. |  |  | 0.85714 | 1 |  |  | (defibrilator) ELECTRODE PADS, paediat., adhesive, disp.
ICU | Monitoring | ECG | Equipment | ELECTROCARDIOGRAPH, portable, 3 ch+ACC | Per centre |  |  |  |  | 1 |  |  |  |  | 3 | 3 | 1384.8237 |  | ELECTROCARDIOGRAPH, portable, 3 ch+ACC
ICU | Monitoring | ECG | Consumables | (ECG) RECORDING PAPER, pack | Per Bed-Day |  |  | 0.01 |  |  |  | Per month. According to use. 5% of moderate patients. |  |  | 6 | 6 |  |  | (ECG) RECORDING PAPER, pack
ICU | Monitoring | ECG | Consumables | (ECG) ELECTRODE GEL, bottle | Per case |  | 0.01 |  |  |  |  | Per month. According to use. 5% of moderate patients. |  |  | 0.42857 | 1 | 8.9374 |  | (ECG) ELECTRODE GEL, bottle
ICU | Diagnostic imaging | Ultrasound | Equipment | ULTRASOUND, SYSTEM, MOBILE, transducer, trolley, 220V, w/ acc. | Per centre |  |  |  |  | 1 |  | Alternative ultrasound to be quoted and to evaluate ease of use. |  |  | 3 | 3 |  |  | ULTRASOUND, SYSTEM, MOBILE, transducer, trolley, 220V, w/ acc.
ICU | Diagnostic imaging | Ultrasound | Consumables | (ultrasound transducer) CONDUCTIVE GEL, 5L, container | Per case |  | 0.01 |  |  |  |  | Per month. According to use. 10 mL per patient, 2 US studies. Calculation is per litre. |  |  | 0.42857 | 1 | 12.8547 |  | (ultrasound transducer) CONDUCTIVE GEL, 5L, container
ICU | Blood Chemistry | Chemistry Analyser | Equipment | CLINICAL CHEMISTRY ANALYSER | Per centre |  |  |  |  | 1 |  | Per centre. One cartridge is aimed to perform a test panel for one patient. |  |  | 3 | 3 |  |  | CLINICAL CHEMISTRY ANALYSER
ICU | Blood Chemistry | Chemistry Analyser | Consumables | CARTRIDGE for Chemistry anaylser (lactic acidosis and hyperlactatemia) | Per case |  | 2 |  |  |  |  | Calculation of this cartridge to be higher than the second one. |  |  | 85.7143 | 86 |  |  | CARTRIDGE for Chemistry anaylser (lactic acidosis and hyperlactatemia)
ICU | Blood Chemistry | Chemistry Analyser | Consumables | CARTRIDGE ( metabolic status and renal function) | Per case |  | 1 |  |  |  |  | Calculation of this cartridge to be lower than the first one. |  |  | 42.8571 | 43 |  |  | CARTRIDGE ( metabolic status and renal function)
ICU | Blood Chemistry | Chemistry Analyser | Consumables | CONTROL,solution 1 | Per equipment |  |  |  | 2 |  |  | Two spare per equipment. |  |  | 6 | 6 |  |  | CONTROL,solution 1
ICU | Blood Chemistry | Chemistry Analyser | Consumables | CONTROL,solution 2 | Per equipment |  |  |  | 2 |  |  | Two spare per equipment. |  |  | 6 | 6 |  |  | CONTROL,solution 2
ICU | Blood Chemistry | Chemistry Analyser | Consumables | CONTROL,solution 3 | Per equipment |  |  |  | 2 |  |  | Two spare per equipment. |  |  | 6 | 6 |  |  | CONTROL,solution 3
ICU | Blood Chemistry | Chemistry Analyser | Consumables | TUBE, VACUUM, plastic, Li-HEPARIN, 2 mL green | Per case |  | 2 |  |  |  |  | Drug list? |  |  | 85.7143 | 86 |  |  | TUBE, VACUUM, plastic, Li-HEPARIN, 2 mL green
ICU | Blood Chemistry | Chemistry Analyser | Consumables | TUBE, VACUUM, plastic, Li-HEPARIN, 4 mL green | Per case |  | 2 |  |  |  |  | Drug list? |  |  | 85.7143 | 86 |  |  | TUBE, VACUUM, plastic, Li-HEPARIN, 4 mL green
ICU | Monitoring | Table, neonate | Equipment | Table, resuscitation, neonate | Per centre |  |  |  |  | 2 |  |  |  |  | 6 | 6 |  |  | Table, resuscitation, neonate
Ancillary | Drug administration | Scale, infant | Equipment | INFANT SCALE, electronic, 0-20 kg | Per centre |  |  |  |  | 2 |  | For drug calculation of paediatric and neonates. |  |  | 6 | 6 | 125 |  | INFANT SCALE, electronic, 0-20 kg
Ancillary | Drug administration | Scale, adult | Equipment | SCALE, electronic, 50g/0-200kg | Per centre |  |  |  |  | 2 |  | For drug calculation of paediatric and neonates. |  |  | 6 | 6 | 90 |  | SCALE, electronic, 50g/0-200kg
Ancillary | Sterilization | Autoclave, 39 L | Equipment | AUTOCLAVE, 39 L, with single burner | Per centre |  |  |  |  | 1 |  |  |  |  | 3 | 3 | 500 |  | AUTOCLAVE, 39 L, with single burner
Ancillary | Sterilization | Autoclave, 39 L | Consumables | Indicator, TST (Time, Steam, Temperature), Type 5 | Per Bed-Day |  |  | 0.5 | [6 empty] | 300 | 300 |  |  | Indicator, TST (Time, Steam, Temperature), Type 5
Ancillary | Sterilization | Autoclave, 39 L | Consumables | Paper sheet, crepe, for sterilization, 60 g/m2, 90 x 90 cm | Per Bed-Day |  |  | 0.5 | [6 empty] | 300 | 300 |  |  | Paper sheet, crepe, for sterilization, 60 g/m2, 90 x 90 cm
Ancillary | Sterilization | Autoclave, 39 L | Consumables | Paper sheet, non-woven, for sterilization, 120 x 120 cm | Per Bed-Day |  |  | 0.5 | [6 empty] | 300 | 300 |  |  | Paper sheet, non-woven, for sterilization, 120 x 120 cm
Ancillary | Sterilization | Autoclave, 39 L | Consumables | Tape, adhesive, indicator Type 1, for steam sterilizer, 18 mm x 48 m | Per Bed-Day |  |  | 0.5 | [6 empty] | 300 | 300 |  |  | Tape, adhesive, indicator Type 1, for steam sterilizer, 18 mm x 48 m
Ancillary | Sterilization | Autoclave, 39 L | Equipment | Kerosene stove, electric heating plate | Per centre |  |  |  |  | 1 |  |  |  |  | 3 | 3 |  |  | Kerosene stove, electric heating plate
Ancillary | Sterilization | Autoclave, 39 L | Equipment | Timer | Per centre |  |  |  |  | 1 |  |  |  |  | 3 | 3 |  |  | Timer

## Procurement List (USD)
Medical Purpose | Main equipment | Type | WHO code | Designation generic | Ratio remarks | Suspected cases | Severe cases | ICU cases | Total Quantity | Approx. Cost (USD) | Total Apprx. Cost (USD)
Monitoring | Patient Monitor | Equipment |  | MONITOR PATIENT, NIBP, w/o ECG , battery, trolley, +acc. | * Supplied with one of: -CUFF ADULT M, navy 23-33cm - CUFF ADULT L, wine 31-40cm - CUFF CHILD, green 12-19cm - CUFF NEON., orange 8-13cm - TUBING NIBP adult/child - SENSOR SPO2 adult - SENSOR SPO2, ped/inf. + adh.wrap OXI-P/I | 0 | 80 | 0 | 80 | 2740.8636 | 219269.0909
Monitoring | Patient Monitor | Equipment |  | MONITOR PATIENT, multiparamet, ECG/CAPNO/SpO2/NIBP/Temp,230V,+acc | * Supplied with one of: -CUFF ADULT L, wine 31-40cm 002207 - CUFF CHILD, green 12-19cm 002201 - TUBING NIBP adult/child 107363 -SENSOR SPO2 adult Nellcor DS100A - SENSOR SPO2, ped/inf. + adh.wrap OXI-P/I -SKIN TEMPERATURE PROBE ad/ped diam.10mm, reus., M1024254 -LEADWIRE ECG multilink 3 grabb.412682-003 -CABLE ECG, 3/5 leads 3.6m 2106305-003 | 0 | 0 | 20 | 5 | 9210.2386 | 46051.1932
Oxygen therapy | Pulse Oximeter | Equipment |  | PULSE OXYMETER - FINGERTIP |  | 0 | 64 | 16 | 16 | 25 | 400
Oxygen therapy | Concentrator O2 | Equipment |  | CONCENTRATOR O2 10L, 230V, 50 Hz + acc. | * Supplied with one of: OUTLET CONNECTOR, FITTING O2 OXYGEN OUTLET | 0 | 80 | 0 | 20 | 1486.4391 | 29728.7827
Oxygen therapy | Concentrator O2 | Accessories |  | (humidifier) HOSE CONNECTOR |  | 0 | 8 | 0 | 4 |  | 0
Oxygen therapy | Concentrator O2 | Accessories |  | HUMIDIFIER, autoclavable |  | 0 | 8 | 2 | 4 | 13 | 52
Oxygen therapy | Concentrator O2 | Accessories |  | (concentr.O2) FLOW SPLITTER pediat., 5 flowmeters 0-2L/min |  | 0 | 8 | 0 | 4 |  | 0
Oxygen therapy | Concentrator O2 | Consumables |  | NASAL OXYGEN CANNULA, 2.1m, 2 prongs + tube, adult |  | 0 | 91.4286 | 34.2857 | 27 | 0.29654 | 8.0067
Oxygen therapy | Concentrator O2 | Consumables |  | NASAL OXYGEN CANNULA, 2 prongs + tube, paediatric |  | 0 | 17.1429 | 6.4286 | 6 | 0.37748 | 2.2649
Oxygen therapy | Concentrator O2 | Consumables |  | NASAL OXYGEN CANNULA, 2 prongs + tube, neonate |  | 0 | 5.7143 | 2.1429 | 3 | 1.678 | 5.0341
Oxygen therapy | Concentrator O2 | Consumables |  | Mask, oxygen, with connection tube, reservoir bag and valve, high-concentration, adult, non-sterile, single use |  | 0 | 91.4286 | 42.8571 | 29 |  | 0
Oxygen therapy | Concentrator O2 | Consumables |  | Mask, oxygen, with connection tube, reservoir bag and valve, high-concentration, paediatric, non-sterile, single use |  | 0 | 17.1429 | 42.8571 | 15 |  | 0
Oxygen therapy | Concentrator O2 | Consumables |  | Venturi Mask, with percent O2 Lock + 2.1 m tubing , adult |  | 0 | 91.4286 | 42.8571 | 29 |  | 0
Oxygen therapy | Concentrator O2 | Consumables |  | Venturi Mask, with percent O2 Lock + 2.1 m tubing , paediatric |  | 0 | 17.1429 | 42.8571 | 15 |  | 0
Oxygen therapy | Concentrator O2 | Consumables |  | Catheter, nasal, 8 Fr, 40 cm, with lateral eyes, sterile, single use |  | 0 | 91.4286 | 42.8571 | 29 |  | 0
Oxygen therapy | Concentrator O2 | Consumables |  | CONNECTOR, biconical, symmetric, ext. diam.7-11 mm, autoclav. |  | 0 | 114.2857 | 42.8571 | 35 | 0.5 | 17.5
Oxygen therapy | Concentrator O2 | Consumables |  | TUBE, silicone, autoclavable, int. diam.5 mm, 25 m | * 3 Tubing of 25 metres (75 metres in total). | 0 | 228.5714 | 85.7143 | 3 | 28.1879 | 84.5638
Oxygen therapy | CPAP, infant | Equipment |  | CPAP 10 machine, w/twin flowmeters | * Supplied with two of: Nasal cannula kit w/prongs, baby bonnet, velcro fixing and chin strap; size 0-1. Nasal cannula kit w/prongs, baby bonnet, velcro fixing and chin strap; size 2-3. Nasal cannula kit w/prongs, baby bonnet, velcro fixing and chin strap; size 4-5. Nasal cannula kit w/prongs, baby bonnet, velcro fixing and chin strap; size 6-7. Humidifier Bubble bottle | 0 | 0.8 | 0 | 1 | 2000 | 2000
Oxygen therapy | CPAP, adult | Equipment |  | CPAP unit w/nasal tubing and mask for adult |  | 0 | 6.4 | 0 | 2 |  | 0
Oxygen therapy | Suction Pump | Equipment |  | SUCTION PUMP, MECHANICAL + collection bottles | * Supplied with two of: COLLECTION BOTTLE, 2 litres, glass | 0 | 12 | 12 | 8 | 400 | 3200
Drug administration | Drop Counter | Equipment |  | ELECTRONIC DROP COUNTER, IV fluids infu. gravity monitor, alarm, batt.AA |  | 0 | 80 | 0 | 15 |  | 0
Blood Chemistry | Chemistry Analyser | Equipment |  | CLINICAL CHEMISTRY ANALYSER | * Supplied with one spare: BATTERY, rechargeable | 0 | 3 | 3 | 2 |  | 0
Blood Chemistry | Chemistry Analyser | Consumables |  | CARTRIDGE for Chemistry anaylser (lactic acidosis and hyperlactatemia) |  | 0 | 228.5714 | 85.7143 | 65 |  | 0
Blood Chemistry | Chemistry Analyser | Consumables |  | CARTRIDGE ( metabolic status and renal function) |  | 0 | 114.2857 | 42.8571 | 33 |  | 0
Blood Chemistry | Chemistry Analyser | Consumables |  | CONTROL,solution 1 |  | 0 | 6 | 6 | 40 |  | 0
Blood Chemistry | Chemistry Analyser | Consumables |  | CONTROL,solution 2 |  | 0 | 6 | 6 | 40 |  | 0
Blood Chemistry | Chemistry Analyser | Consumables |  | CONTROL,solution 3 |  | 0 | 6 | 6 | 40 |  | 0
Blood Chemistry | Chemistry Analyser | Consumables |  | TUBE, VACUUM, plastic, Li-HEPARIN, 2 mL green |  | 0 | 228.5714 | 85.7143 | 65 |  | 0
Blood Chemistry | Chemistry Analyser | Consumables |  | TUBE, VACUUM, plastic, Li-HEPARIN, 4 mL green |  | 0 | 228.5714 | 85.7143 | 65 |  | 0
Oxygen therapy | O2 Wall | Equipment |  | Flowmeter, Thorpe tube, for oxygen 0-15L/min | For wall oxygen. | 0 | 160 | 40 | 10 |  | 0
Airway Management | Independent | Equipment |  | CRICOTHYROTOMY, SET, emergency, 6 mm, ster., s.u., w/acc. | Supplied with: 1 - 6.0 mm, cuffed cricothyroidotomy tube 1 - cricothyroidotomy, tube holder 1 - dilator 1 - scalpel blade, No. 15, for handle No. 3, sterile, single use 1 - suture, surgical, synthetic, non absorbable, monofilament, DEC 3.5 (0), 45 cm, with needle, 3/8 circle, 29.9 mm, cutting point 1 - syringe, 10 mL, two or three pieces, Luer type, sterile, single use 1 - Thermovent T (HME) | 0 | 0 | 12 | 4 |  | 0
Airway Management | Independent | Consumables |  | AIRWAY, NASOPHARYNGEAL, sterile, set, single use (from Fr20 to Fr36) | Airway, nasopharyngeal, sterile, single use. Range of sizes. Fr20, Fr22, Fr24, Fr26, Fr28, Fr30, Fr32, Fr34, Fr36 | 0 | 0 | 42.8571 | 11 |  | 0
Airway Management | Independent | Consumables |  | AIRWAY, OROPHARYNGEAL, Guedel, sterile, set, single use (00, 0, 1, 2, 3, 4, 5) | Airway, oropharyngeal, Guedel, sterile, single use. Range of sizes: 00, 0, 1, 2, 3, 4, 5. | 0 | 0 | 42.8571 | 11 |  | 0
Airway Management | Independent | Consumables |  | ENDOTRACHEAL TUBE INTRODUCER, Bougie, ster., s.u., Fr15, 70cm |  | 0 | 0 | 42.8571 | 11 |  | 0
Airway Management | Independent | Consumables |  | ENDOTRACHEAL TUBE INTRODUCER, Bougie, ster., s.u., Fr10, 60cm |  | 0 | 0 | 42.8571 | 11 |  | 0
Airway Management | Independent | Consumables |  | ENDOTRACHEAL TUBE INTRODUCER, Stylet, ster., s.u., Fr14, 30 to 45 cm |  | 0 | 0 | 42.8571 | 11 |  | 0
Airway Management | Independent | Consumables |  | ENDOTRACHEAL TUBE INTRODUCER, Stylet, ster., s.u., Fr10, 30 to 45 cm |  | 0 | 0 | 42.8571 | 11 |  | 0
Airway Management | Independent | Consumables |  | TUBE, ENDOTRACHEAL, No.2, w/o cuff, ster., s.u. |  | 0 | 0 | 42.8571 | 11 |  | 0
Airway Management | Independent | Consumables |  | TUBE, ENDOTRACHEAL, No.2.5, w/o cuff, ster., s.u. |  | 0 | 0 | 42.8571 | 11 |  | 0
Airway Management | Independent | Consumables |  | TUBE, ENDOTRACHEAL, No.3, w/o cuff, ster., s.u. |  | 0 | 0 | 42.8571 | 11 |  | 0
Airway Management | Independent | Consumables |  | TUBE, ENDOTRACHEAL, No.3.5, w/o cuff, ster., s.u. |  | 0 | 0 | 42.8571 | 11 |  | 0
Airway Management | Independent | Consumables |  | TUBE, ENDOTRACHEAL, No.4, w/o cuff, ster., s.u. |  | 0 | 0 | 42.8571 | 11 |  | 0
Airway Management | Independent | Consumables |  | TUBE, ENDOTRACHEAL, No.5, w/o cuff, ster., s.u. |  | 0 | 0 | 42.8571 | 11 |  | 0
Airway Management | Independent | Consumables |  | TUBE, ENDOTRACHEAL, No.4, w/ cuff, ster., s.u. |  | 0 | 0 | 42.8571 | 11 |  | 0
Airway Management | Independent | Consumables |  | TUBE, ENDOTRACHEAL, No.5, w/ cuff, ster., s.u. |  | 0 | 0 | 42.8571 | 11 |  | 0
Airway Management | Independent | Consumables |  | TUBE, ENDOTRACHEAL, No.6, w/ cuff, ster., s.u. |  | 0 | 0 | 42.8571 | 11 |  | 0
Airway Management | Independent | Consumables |  | TUBE, ENDOTRACHEAL, No.7, w/ cuff, ster., s.u. |  | 0 | 0 | 42.8571 | 11 |  | 0
Airway Management | Independent | Consumables |  | TUBE, ENDOTRACHEAL, No.8, w/ cuff, ster., s.u. |  | 0 | 0 | 42.8571 | 11 |  | 0
Airway Management | Independent | Consumables |  | TUBE, ENDOTRACHEAL, No.9, w/ cuff, ster., s.u. |  | 0 | 0 | 42.8571 | 11 |  | 0
Airway Management | Independent | Consumables |  | LARYNGEAL MASK AIRWAY (LMA), size 2, sterile, single use |  | 0 | 0 | 42.8571 | 11 |  | 0
Airway Management | Independent | Consumables |  | LARYNGEAL MASK AIRWAY (LMA), size 3, sterile, single use |  | 0 | 0 | 42.8571 | 11 |  | 0
Airway Management | Independent | Consumables |  | LARYNGEAL MASK AIRWAY (LMA), size 4, sterile, single use |  | 0 | 0 | 42.8571 | 11 |  | 0
Airway Management | Independent | Consumables |  | COLORIMETRIC END TIDAL CO2 DETECTOR, adult, s.u. |  | 0 | 0 | 42.8571 | 11 |  | 0
Airway Management | Independent | Consumables |  | COLORIMETRIC END TIDAL CO2 DETECTOR, paediatric, s.u. |  | 0 | 0 | 42.8571 | 11 |  | 0
Airway Management | Independent | Consumables |  | SYRINGE, Luer slip, 10mL, sterile, single use |  | 0 | 0 | 42.8571 | 11 | 0.8567 | 9.4237
Airway Management | Independent | Consumables |  | LUBRICATING Jelly, 5g, box of 150 |  | 0 | 0 | 42.8571 | 11 |  | 0
Airway Management | Independent | Consumables |  | FORCEPS, MAGILL, paediatric, 15 cm |  | 0 | 0 | 42.8571 | 11 |  | 0
Airway Management | Independent | Consumables |  | FORCEPS, MAGILL, child, 19 cm |  | 0 | 0 | 42.8571 | 11 |  | 0
Airway Management | Independent | Consumables |  | FORCEPS, MAGILL, adult, 24 cm |  | 0 | 0 | 42.8571 | 11 |  | 0
Airway Management | Laryngoscope, adult | Equipment |  | LARYNGOSCOPE,fib.opt, ad/ch, diam.28mm, blades(Macintosh2/3/4+Miller1), sp.bulbs | Supplied with: * Blades, Macintosh type (curved): No. 2, length 90 - 110 mm, for child. No. 3, length 110 - 135 mm, for small adult. No. 4, length 135 - 155 mm, for adult. * Blades, Miller type (straight): No. 1, length 100 mm. * Heavy-walled plastic or metal case. * Six compatible batteries in total. * Four extra halogen bulbs. | 0 | 0 | 24 | 8 | 675 | 5400
Airway Management | Laryngoscope, infant | Equipment |  | LARYNGOSCOPE,fib.opt, neonate, diam.19mm, blades(Macintosh0/1/2), sp.bulbs,batt. | Supplied with: * Blades, Macintosh type (curved): No. 0, length 55 mm, for newborn. No. 1, length 70 mm, for infant. No. 2, length 90 mm, for child. * Heavy-walled plastic or metal case. * Six compatible batteries in total. * Four extra halogen bulbs. | 0 | 0 | 12 | 4 | 600 | 2400
Mechanical ventilation | Resuscitator, adult | Equipment |  | SELF-INFLATING BAG, ad./child + masks RH5 / RH2 (Ambu type) | Compressible self-refilling ventilation bag, capacity: 1475-2000mL. Oxygen reservoir bag complete. Non-rebreathing patient valve with pressure limiting valve, patient connector outside/inside diameter: 22/15 mm. Inlet valve with nipple for 02 tubing. Masks, silicon, in 3 sizes (adult small, adult medium and adult large). | 0 | 0 | 12 | 4 | 80 | 320
Mechanical ventilation | Resuscitator, infant | Equipment |  | SELF-INFLATING BAG, child/neonate + masks RH2/S1 (Ambu type) | Compressible self-refilling ventilation bag, child, capacity: 500-700mL. Oxygen reservoir bag complete. Non-rebreathing patient valve with pressure limiting valve, patient connector outside/inside diameter: 22/15 mm. Inlet valve with nipple for 02 tubing. Masks, silicon, for infants. | 0 | 0 | 6 | 2 | 97 | 194
Mechanical ventilation | Suction device | Equipment |  | SUCTION BULB, for newborn, reusable, autoclavable |  | 0 | 0 | 6 | 2 |  | 0
Mechanical ventilation | Independent | Consumables |  | FILTER, HEAT AND MOISTURE EXCHANGER (HMEF), high efficiency, with connectors, adult, single use | Normally for long term ventilated patients. | 0 | 0 | 34.2857 | 9 |  | 0
Mechanical ventilation | Independent | Consumables |  | FILTER, HEAT AND MOISTURE EXCHANGER (HMEF), high efficiency, with connectors, paediatric, single use | Normally for long term ventilated patients. | 0 | 0 | 8.5714 | 3 |  | 0
Gastro-enteral Feeding | Independent | Consumables |  | TUBE, FEEDING, NASOGASTRIC, 10 Fr, 50 cm, ENFit tip, sterile, single use |  | 0 | 0 | 42.8571 | 11 |  | 0
Gastro-enteral Feeding | Independent | Consumables |  | TUBE, FEEDING, NASOGASTRIC, 12 Fr, 90 cm, ENFit tip, sterile, single use |  | 0 | 0 | 42.8571 | 11 |  | 0
Gastro-enteral Feeding | Independent | Consumables |  | TUBE, FEEDING, NASOGASTRIC, 14 Fr, 90 cm, ENFit tip, sterile, single use |  | 0 | 0 | 42.8571 | 11 |  | 0
Gastro-enteral Feeding | Independent | Consumables |  | TUBE, FEEDING, NASOGASTRIC, 6 Fr, 50 cm, ENFit tip, sterile, single use |  | 0 | 0 | 42.8571 | 11 |  | 0
Gastro-enteral Feeding | Independent | Consumables |  | TUBE, FEEDING, NASOGASTRIC, 8 Fr, 50 cm, ENFit tip, sterile, single use |  | 0 | 0 | 42.8571 | 11 |  | 0
Gastro-enteral Feeding | Independent | Consumables |  | SYRINGE, FEEDING, 1 mL, LDT, ENFit, sterile, single use |  | 0 | 0 | 42.8571 | 11 |  | 0
Gastro-enteral Feeding | Independent | Consumables |  | SYRINGE, FEEDING, 10 mL, ENFit, sterile, single use |  | 0 | 0 | 42.8571 | 11 |  | 0
Gastro-enteral Feeding | Independent | Consumables |  | SYRINGE, FEEDING, 2.5 mL, LDT, ENFit, sterile, single use |  | 0 | 0 | 42.8571 | 11 |  | 0
Gastro-enteral Feeding | Independent | Consumables |  | SYRINGE, FEEDING, 20 mL, ENFit, sterile, single use |  | 0 | 0 | 42.8571 | 11 |  | 0
Gastro-enteral Feeding | Independent | Consumables |  | SYRINGE, FEEDING, 5 mL, LDT, ENFit, sterile, single use |  | 0 | 0 | 42.8571 | 11 |  | 0
Gastro-enteral Feeding | Independent | Consumables |  | SYRINGE, FEEDING, 60 mL, ENFit, sterile, single use |  | 0 | 0 | 42.8571 | 11 |  | 0
Gastro-enteral Feeding | Independent | Consumables |  | LUBRICATING Jelly, 50g, tube |  | 0 | 0 | 20 | 5 |  | 0
Gastro-enteral Feeding | Independent | Consumables |  | PAD, absorbent |  | 0 | 0 | 42.8571 | 11 |  | 0
Gastro-enteral Feeding | Independent | Consumables |  | BASIN KIDNEY, stainless steel, 825 mL |  | 0 | 0 | 20 | 5 |  | 0
Gastro-enteral Feeding | Independent | Consumables |  | STETHOSCOPE, binaural, double cup, adult/child, single use |  | 0 | 0 | 20 | 5 |  | 0
Central line | Independent | Consumables |  | CENTRAL VENOUS CATHETERS KIT | Supplied with: Finder needle, syringe, wire, dilator, lidocaine, scapel, needle, thread (e.g., ARROW) | 0 | 0 | 42.8571 | 11 |  | 0
Central line | Independent | Consumables |  | Transparent adhesive plasters, washproof, 5x5 cm |  | 0 | 0 | 128.5714 | 33 |  | 0
Urine collection | Independent | Consumables |  | Bag, collecting, urine, with outlet tap, with non-return valve, 2000 mL, adult, non-sterile, single use |  | 0 | 0 | 42.8571 | 11 |  | 0
Urine collection | Independent | Consumables |  | CATHETER, URETHRAL, Foley, 2-way, 10 Fr, sterile, single use |  | 0 | 0 | 42.8571 | 11 |  | 0
Urine collection | Independent | Consumables |  | CATHETER, URETHRAL, Foley, 2-way, 12 Fr, sterile, single use |  | 0 | 0 | 42.8571 | 11 |  | 0
Urine collection | Independent | Consumables |  | CATHETER, URETHRAL, Foley, 2-way, 14 Fr, sterile, single use |  | 0 | 0 | 42.8571 | 11 |  | 0
Urine collection | Independent | Consumables |  | CATHETER, URETHRAL, Foley, 2-way, 16 Fr, sterile, single use |  | 0 | 0 | 42.8571 | 11 |  | 0
Urine collection | Independent | Consumables |  | CATHETER, URETHRAL, Foley, 2-way, 18 Fr, sterile, single use |  | 0 | 0 | 42.8571 | 11 |  | 0
Urine collection | Independent | Consumables |  | CATHETER, URETHRAL, Foley, 2-way, 20 Fr, sterile, single use |  | 0 | 0 | 42.8571 | 11 |  | 0
Urine collection | Independent | Consumables |  | CATHETER, URETHRAL, Foley, 2-way, 8 Fr, sterile, single use |  | 0 | 0 | 42.8571 | 11 |  | 0
General supplies | Independent | Consumables |  | Compress, gauze, 10 x 10 cm, 8 to 12 ply, sterile, single use |  | 0 | 0 | 42.8571 | 11 |  | 0
General supplies | Independent | Consumables |  | Tape, surgical, hypoallergenic, 5 x 2.5 cm |  | 0 | 0 | 42.8571 | 11 |  | 0
General supplies | Independent | Consumables |  | Drape, surgical, non woven, sterile, single use |  | 0 | 0 | 42.8571 | 11 |  | 0
General supplies | Independent | Consumables |  | Gloves, examination, nitrile, powder-free, pair-packed, large, sterile, single use | Sterile gloves. | 0 | 0 | 42.8571 | 11 |  | 0
General supplies | Independent | Consumables |  | Gloves, examination, nitrile, powder-free, pair-packed, medium, sterile, single use | Sterile gloves. | 0 | 0 | 42.8571 | 11 |  | 0
General supplies | Independent | Consumables |  | Gloves, examination, nitrile, powder-free, pair-packed, small, sterile, single use | Sterile gloves. | 0 | 0 | 42.8571 | 11 |  | 0
General supplies | Independent | Consumables |  | Antiseptic Wipe with Alcohol & Chlorhexidine |  | 0 | 0 | 42.8571 | 11 |  | 0
Blood Chemistry | Independent | Consumables |  | Arterial Blood Sample Kits | Arterial blood sampling syringes with dry lithium heparin [23.5 iu/mL] | 0 | 0 | 42.8571 | 11 |  | 0
Mechanical ventilation | Patient Ventilator | Equipment |  | VENTILATOR PATIENT, for adult, paediatric and neonate w/acc. | It includes all accessories for operation, including the CO2 detector. | 0 | 0 | 10 | 3 | 35000 | 105000
Mechanical ventilation | Patient Ventilator | Consumables |  | Breathing circuit (tubes/balloon/valves/mask), ADULT, single use |  | 0 | 0 | 34.2857 | 9 |  | 0
Mechanical ventilation | Patient Ventilator | Consumables |  | Breathing circuit (tubes/balloon/valves/mask), PAEDIATRIC, single use |  | 0 | 0 | 6.4286 | 2 |  | 0
Mechanical ventilation | Patient Ventilator | Consumables |  | Breathing circuit (tubes/balloon/valves/mask), NEONATE, single use |  | 0 | 0 | 2.1429 | 1 |  | 0
Oxygen therapy | HFNC | Equipment |  | High Flow Nasal Cannula (HFNC) |  | 0 | 0 | 10 | 3 |  | 0
Oxygen therapy | HFNC | Consumables |  | Optiflow + Nasal Cannula, small, pack of 20 |  | 0 | 0 | 42.8571 | 11 |  | 0
Oxygen therapy | HFNC | Consumables |  | Optiflow + Nasal Cannula, medium, pack of 20 |  | 0 | 0 | 42.8571 | 11 |  | 0
Oxygen therapy | HFNC | Consumables |  | Optiflow + Nasal Cannula, large, pack of 20 |  | 0 | 0 | 42.8571 | 11 |  | 0
Oxygen therapy | HFNC | Consumables |  | Disinfection kit |  | 0 | 0 | 34.2857 | 9 |  | 0
Oxygen therapy | HFNC | Consumables |  | Cleaning Sponge Stick, pack of 20 |  | 0 | 0 | 6.4286 | 2 |  | 0
Oxygen therapy | HFNC | Consumables |  | Disinfection filter, pack of 2 |  | 0 | 0 | 2.1429 | 1 |  | 0
Oxygen therapy | Suction Pump | Equipment |  | SUCTION PUMP, ELECTRICAL, 100-230V,50-60Hz | * Supplied with one of: COLLECTION BOTTLE, 1L, autoclavable LID W/CONNECTOR AND OVERFLOW DEVICE | 0 | 0 | 20 | 5 |  | 0
Oxygen therapy | Suction Pump | Consumables |  | BACTERIA FILTER, unit |  | 0 | 0 | 21.4286 | 6 |  | 0
Oxygen therapy | Suction Pump | Consumables |  | TUBE, silicone, autoclavable, int. diam.8 mm, 10 m | * 3 Tubing of 10 metres (30 metres in total). | 0 | 2.8571 | 43.9286 | 3 |  | 0
Drug administration | Infusion Pump | Equipment |  | INFUSION PUMP |  | 0 | 0 | 20 | 5 | 1470.9942 | 7354.9708
Drug administration | Infusion Pump | Consumables |  | (inf. pump) INFUSION LINE |  | 0 | 0 | 85.7143 | 22 | 0.85182 | 18.74
Drug administration | Drill | Equipment |  | DRILL, FOR VASCULAR ACCESS, with transport bag | * Supplied with one of: TRAINING KIT | 0 | 0 | 3 | 1 |  | 0
Drug administration | Drill | Consumables |  | NEEDLE + STABILIZER KIT, 15 G, 15 mm, PAEDIATRIC |  | 0 | 0 | 4.2857 | 2 |  | 0
Drug administration | Drill | Consumables |  | NEEDLE + STABILIZER KIT, 15 G, 25 mm, ADULT |  | 0 | 0 | 3.4286 | 1 |  | 0
Monitoring | Defibrillator | Equipment |  | DEFIBRILLATOR, mobile, semi-auto.,multi-paramet,AC/DC, w/acc+trolley |  | 0 | 0 | 3 | 1 | 4040.9124 | 4040.9124
Monitoring | Defibrillator | Consumables |  | (defibrilator) ELECTRODE PADS, adult, adhesive, disp. |  | 0 | 0 | 3.4286 | 1 |  | 0
Monitoring | Defibrillator | Consumables |  | (defibrilator) ELECTRODE PADS, paediat., adhesive, disp. |  | 0 | 0 | 0.85714 | 1 |  | 0
Monitoring | ECG | Equipment |  | ELECTROCARDIOGRAPH, portable, 3 ch+ACC | * Supplied with one of: PATIENT CABLE 10 leads, SET SUCTION ELECTRODES, reusable ELECTRODES CLIP, limb, set 4pcs/colors | 0 | 0 | 3 | 1 | 1384.8237 | 1384.8237
Monitoring | ECG | Consumables |  | (ECG) RECORDING PAPER, pack |  | 0 | 0 | 6 | 4 |  | 0
Monitoring | ECG | Consumables |  | (ECG) ELECTRODE GEL, bottle |  | 0 | 0 | 2.5714 | 3 | 8.9374 | 26.8121
Diagnostic imaging | Ultrasound | Equipment |  | ULTRASOUND, SYSTEM, MOBILE, transducer, trolley, 220V, w/ acc. | * Supplied with one of: LINEAR TRANSDUCER 5.0-7.5 MHz. PHASED ARRAY CARDIAC TRANSDUCER 5.0-7.5 MHz. | 0 | 0 | 3 | 1 |  | 0
Diagnostic imaging | Ultrasound | Consumables |  | (ultrasound transducer) CONDUCTIVE GEL, 5L, container |  | 0 | 0 | 0.42857 | 4 | 12.8547 | 51.4188
Monitoring | Table, neonate | Equipment |  | Table, resuscitation, neonate |  | 0 | 0 | 6 | 2 |  | 0
Drug administration | Scale, infant | Equipment |  | INFANT SCALE, electronic, 0-20 kg |  | 0 | 0 | 6 | 2 | 125 | 250
Drug administration | Scale, adult | Equipment |  | SCALE, electronic, 50g/0-200kg |  | 0 | 0 | 6 | 2 | 90 | 180
Sterilization | Autoclave, 39 L | Equipment |  | AUTOCLAVE, 39 L, with single burner |  | 0 | 0 | 3 | 1 | 500 | 500
Sterilization | Autoclave, 39 L | Consumables |  | Indicator, TST (Time, Steam, Temperature), Type 5 |  | 0 | 0 | 300 | 30 |  | 0
Sterilization | Autoclave, 39 L | Consumables |  | Paper sheet, crepe, for sterilization, 60 g/m2, 90 x 90 cm |  | 0 | 0 | 300 | 30 |  | 0
Sterilization | Autoclave, 39 L | Consumables |  | Paper sheet, non-woven, for sterilization, 120 x 120 cm |  | 0 | 0 | 300 | 30 |  | 0
Sterilization | Autoclave, 39 L | Consumables |  | Tape, adhesive, indicator Type 1, for steam sterilizer, 18 mm x 48 m |  | 0 | 0 | 300 | 30 |  | 0
Sterilization | Autoclave, 39 L | Equipment |  | Kerosene stove, electric heating plate |  | 0 | 0 | 3 | 1 |  | 0
Sterilization | Autoclave, 39 L | Equipment |  | Timer |  | 0 | 0 | 3 | 1 |  | 0

## WHO Specs_Generic Items (G)
Remarks about the item | Designation (EN) | Generic complete description | Technical specifications for procurement | Sizes | Material | Standards (Equivalent test methods could be accepted) | Packaging
- | (concentr.O2) FLOW SPLITTER pediat., 5 flowmeters 0-2L/min (Sureflow FM069-2) | Flow splitter, for oxygen supply | Flow splitter for distribution of the oxygen delivery. Each of five outlets with a flowmeter for independent control of the oxygen flow rates. Full scale is graduated in litres per minute. The device is connected to a single oxygen supply (e.g., concentrator). Input pressure: 50 to 350 kPa. Flowmeter capacity: 0.125 to 2 L/min. Accuracy: ± 5%. Suitable for table top or wall mount. |  |  |  | Each
- | NASAL OXYGEN CANNULA, 2.1m, 2 prongs + tube, adult | Cannula with nasal prongs, for oxygen, adult, non-sterile, single use | Cannula with nasal prongs designed for easy administration of medicinal oxygen through patient nostrils. An oxygen flow rate of 2 to 4 L/min delivers a FiO2 of 0.28 to 0.36, respectively. The twin prongs nasal tips are soft and smoothly finished to ensure equal oxygen flow to both nostrils. They are connected to a lip support and harness (one tube right / left side). The harness is fully adjustable (over the patient ear) with a double tubing (right and left side), interlinked through a moulded Y-connector to the oxygen supply line. All tubing is soft and flexible funnel, kink resistant, with star lumen, and with proximal end with an universal connector to oxygen source. Individually packed in a sealed plastic envelope. | Size: Adult. * Outer diameter of the prong: 4 mm. * Septum width: 10 mm or adjustable. * Maximum flow: 6 to 8 L/min. * Tube length: 1.6 to 2 m. | Polyvinyl chloride (PVC); transparent; medical grade. Hygienically clean for medical use. | ISO/DIS 23368: Anaesthetic and respiratory equipment — Low flow nasal cannula for oxygen therapy. ISO/DIS 17256: Anaesthetic and respiratory equipment — Respiratory therapy tubing and connectors. | Box of 100 units
- | NASAL OXYGEN CANNULA, 2 prongs + tube, paediatric | Cannula with nasal prongs, for oxygen, paediatric, non-sterile, single use | Idem. | Size: Paediatric. * Outer diameter of the prong: 3.7 mm. * Septum width: 8 mm. * Maximum flow: 6 L/min. * Tube length: 1.6 to 2 m. | Idem. | Idem. | Idem.
- | NASAL OXYGEN CANNULA, 2 prongs + tube, neonate | Cannula with nasal prongs, for oxygen, neonate, non-sterile, single use | Idem. | Size: Neonate. * Outer diameter of the prong: 2.4 mm. * Septum width: 3.5 mm. * Maximum flow: 2 L/min. * Tube length: 1.6 to 2 m. | Idem. | Idem. | Idem.
- | Mask, oxygen, with connection tube, reservoir bag and valve, high-concentration, adult, non-sterile, single use | Mask, oxygen, with connection tube, reservoir bag and one way valve, high-concentration, adult, non-sterile, single use | Facemask used to deliver medical oxygen directly to the upper airway of the patient. It allows the administration in a high oxygen concentration of 80 to 100 %. It includes two unidirectional valves, one that closes during inspiration to prevent room air mixing with oxygen in a reservoir bag, and one that closes during exhalation to prevent exhaled respiratory gases from entering the reservoir bag (Non-rebreathing oxygen face mask). Mask is soft, transparent, well-fitting moulded, with two side vents (either four to six small wholes, or two slits). The nose clip is soft, malleable and adjustable. The tubing (oxygen line) is non-kinking, well-fitted. Tube length: 2 m. A standard connection is recommended. Individually packed in a sealed plastic envelope. | Size: Adult. | Mask and tubing: Polyvinyl chloride (PVC); transparent; medical grade. Strap: Elastomer. Nose clip: Aluminium. Hygienically clean for medical use. | ISO 11712:2009: Anaesthetic and respiratory equipment — Supralaryngeal airways and connectors. | Box of 50 units
- | Mask, oxygen, with connection tube, reservoir bag and valve, high-concentration, paediatric, non-sterile, single use | Mask, oxygen, with connection tube, reservoir bag and one way valve, high-concentration, paediatric, non-sterile, single use | Idem. | Size: Paediatric. | Idem. | Idem. | Idem.
- | Venturi Mask, with percent O2 Lock + 2.1 m tubing , adult | Venturi Mask, with percent O2 Lock + tubing , adult | Also known as an air-entrainment mask. It delivers low to moderate flow oxygen, with a specific concentration from 24% to 50% with a flow rate of 3 or 6 L/min. The mask has an adjustable nose clip. The kit include the tubing, humidity cup and multiple jets, which are color-coded and indicating the percentage of oxygen. | Size: Adult. |  | ISO 11712:2009: Anaesthetic and respiratory equipment — Supralaryngeal airways and connectors.
- | Venturi Mask, with percent O2 Lock + 2.1 m tubing , paediatric | Venturi Mask, with percent O2 Lock + tubing , paediatric | Idem. | Size: Paediatric. |  | Idem.
- | Catheter, nasal, 8 Fr, 40 cm, with lateral eyes, sterile, single use | Catheter, oxygen, 40 cm, with lateral eyes, sterile, single use | Nasal catheter for the administration of medical oxygen. Open distal end with multiple lateral holes, or a central eye and distal cross perforation. Proximal end features a straight conical connector. | Length: 40 cm. ID: 8 or 10Fr. | PVC. Proximal end: Polyester piece of foam. | ISO/DIS 23368: Anaesthetic and respiratory equipment — Low flow nasal cannula for oxygen therapy. ISO/DIS 17256: Anaesthetic and respiratory equipment — Respiratory therapy tubing and connectors. | Each
- | CONNECTOR, biconical, symmetric, ext. diam.7-11 mm, autoclav. | Connector, biconical, symmetric, fits tubing with external diameter 7 to 11 mm, autoclavable | A rigid short tube used to connect catheters and tubes of different diameters. The external diameter gradually decreases in size, symmetrically, from the centre towards the ends. | Length: 5 cm. External diameter, gradually decrease from: 11 to 7 mm, approximately. | Polycarbonate; medical grade; transparent opaque; disinfectant resistant and autoclavable. |  | Box of 100 units
- | TUBE, silicone, autoclavable, int. diam.5 mm, 25 m | Tubing, medical gases, ID 5 mm, L 25 m, autoclavable | Medical grade tubing used for the circulation of medical gases. Translucent, flexible and soft plastic tubing of uniform diameter. Anti-kink tubing, non-permanent deformation if kinked or bent too tight. | Internal diameter: 5 mm Minimum length: 25 m. | Polycarbonate; medical grade; transparent opaque. It must be capable to resist cleaning and disinfection with chlorine solution (0.05% concentration). It must resist steam sterilization at 121 °C or 134 °C. | ISO/DIS 17256: Anaesthetic and respiratory equipment — Respiratory therapy tubing and connectors. ISO 5359:2014: Anaesthetic and respiratory equipment — Low-pressure hose assemblies for use with medical gases. | Each
It is needed the complete set for arterial blood sample collection. | TUBE, VACUUM, plastic, Li-HEPARIN, 2 mL green | Arterial blood sample collection kit with TUBE, VACUUM, plastic, Li-HEPARIN, 2 mL green | The kit includes: Vacuum sterile plastic tube, with Li-Heparin, intended to contain a blood specimen. Perforable plug colour: green. Volume 5 mL, vacuum 2 mL. Dimensions: 75 mm x 13 mm Ø. Paper label. Holder for vacuum tube with needle injector. Needle, 21 or 23 G, with wings, for vacuum tube. |  |  | ISO 6710:2012: Single-use containers for human venous blood specimen collection.
It is needed the complete set for arterial blood sample collection. | TUBE, VACUUM, plastic, Li-HEPARIN, 4 mL green | Arterial blood sample collection kit with TUBE, VACUUM, plastic, Li-HEPARIN, 4 mL green | The kit includes: Vacuum sterile plastic tube, with Li-Heparin, intended to contain a blood specimen. Perforable plug colour: green. Volume 5 mL, vacuum 4 mL. Dimensions: 75 mm x 13 mm Ø. Paper label. Holder for vacuum tube with needle injector. Needle, 21 or 23 G, with wings, for vacuum tube.
- | Flowmeter, Thorpe tube, for oxygen 0-15L/min | Flowmeter, Thorpe tube, for oxygen 0-15L/min | The Thorpe tube flowmeter is composed of inlet and outlet ports, a regulator, a valve and a clear tapered measuring tube. It is suitable for connection with various medical gas sources, such as centralized system, cylinders, or compressors. Standard (absolute, non-compensated) and pressure-compensated flowmeter versions, suitable for specific flow ranges.
Included in the TESK. | CRICOTHYROTOMY, SET, emergency, 6 mm, ster., s.u., w/acc. | CRICOTHYROTOMY, SET, emergency, 6 mm, ster., s.u., w/acc. | Supplied with: 1 - 6.0 mm, cuffed cricothyroidotomy tube 1 - cricothyroidotomy, tube holder 1 - dilator 1 - scalpel blade, No. 15, for handle No. 3, sterile, single use 1 - suture, surgical, synthetic, non absorbable, monofilament, DEC 3.5 (0), 45 cm, with needle, 3/8 circle, 29.9 mm, cutting point 1 - syringe, 10 mL, two or three pieces, Luer type, sterile, single use 1 - Thermovent T (HME)
| Airway, nasopharyngeal, sterile, set, single use | Airway, nasopharyngeal, sterile, set, single use | A Nasopharyngeal Airway is recommended for use as an airway adjunct in the semi-conscious or unconscious patient with an intact gag reflex. Individually packaged sterile with a conveniently attached surgical lubricant for quick access to facilitate ease of insertion. Flexible and soft material for maximum patient comfort. Rounded tip allows for gentle insertion. Trumpet design for secure placement. Diameter and size labelled according to standards. | Range of sizes. Fr20, Fr22, Fr24, Fr26, Fr28, Fr30, Fr32, Fr34, Fr36 |  |  | Each
- | Airway, oropharyngeal, Guedel, sterile, set, single use | Airway, oropharyngeal, Guedel, sterile, set, single use | One-piece, semi-rigid, curved plastic tube, used to be inserted through the oropharynges to facilitate airway management. Guedel type. Flange surface is permanently marked with tube size/length in mm, and the manufacturer or supplier’s name. It is bite resistant. Proximal (or buccal) end straight and reinforced. Distal end semi-rigid, curved, with atraumatic soft rounded edges. Colour coded. | Range of sizes: 00, 0, 1, 2, 3, 4, 5. | Polyethylene vinyl acetate (EVA) and Polyvinyl chloride (PVC); siliconized; transparent; medical grade. | EN12181 or ISO 5364: Oropharyngeal airways. ISO10993-1: Biological safety testing. | Each
- | ENDOTRACHEAL TUBE INTRODUCER, Bougie, ster., s.u., Fr15, 70cm | Endotracheal Tube Introducer (Bougie), elastic, Fr15, 70cm, sterile, single use | To assist with endotracheal intubations is used to guide the tube properly into the airway. Blue or yellow tube with graduated marking. Curved tip with distal rounded smooth tip. * Initial sterilisation method: Ethylene oxide gas or Gamma radiation. * Individually peel-packed in paper and/or plastic. |  | Flexible, medical grade, radiopaque, braided polyester base with a resin coating. | ISO 5361:2016: Anaesthetic and respiratory equipment — Tracheal tubes and connectors. ISO 10993-1:2018: Biological evaluation of medical devices — Part 1: Evaluation and testing within a risk management process. ISO 11135:2014: Sterilization of health-care products — Ethylene oxide — Requirements for the development, validation and routine control of a sterilization process for medical devices. | Each
- | ENDOTRACHEAL TUBE INTRODUCER, Bougie, ster., s.u., Fr10, 60cm | Endotracheal Tube Introducer (Bougie), elastic, Fr10, 60cm, sterile, single use | Idem. |  | Idem. | Idem. | Idem.
- | ENDOTRACHEAL TUBE INTRODUCER, Stylet, ster., s.u., Fr14, 30 to 45 cm | Guide (Stylet), for endotracheal tube, Fr14, 30 to 45 cm, sterile, single use | Flexible and malleable guide (stylet) to be inserted into the endotracheal tube to guide it properly during the intubation. It has a soft and round end-tip. It can be shaped as needed. It has graduated marking. Manufacturer name and tube size are indicated on the tube. |  | Stylet: malleable metal alloy covered by polyvinyl chloride (PVC) coated, latex free; white; disinfectant resistant; withstand steam sterilization. | ISO 5361:2016: Anaesthetic and respiratory equipment — Tracheal tubes and connectors. ISO 10993-1:2018: Biological evaluation of medical devices — Part 1: Evaluation and testing within a risk management process. | Box of 10 units
- | ENDOTRACHEAL TUBE INTRODUCER, Stylet, ster., s.u., Fr10, 30 to 45 cm | Guide (Stylet), for endotracheal tube, Fr10, 30 to 45 cm, sterile, single use | Idem. |  | Idem. | Idem. | Idem.
- | TUBE, ENDOTRACHEAL, No.2, w/o cuff, ster., s.u. | Tube, endotracheal, No. 2, without cuff, sterile, single use | Endotracheal tube designed to be introduced into the trachea via the mouth or nose, ensuring an unobstructed upper airway to convey gases and vapours to and from the lungs. It consists in a thin, flexible, transparent and single hollow cylinder, with an anatomical curvature Magill-type of 37.5°, black and legibly depth markings and graduation in centimetres, with radio-opaque continuous line mark, without cuff, with a standard connector in the proximal end. The proximal end of the tube safely fits the connector from which size is selected according to the tube size. The connector is straight and double-ended, with the proximal end being an outer, standard 15 mm internal diameter, conical tip that allows the tube to be connected to the ventilation system (breathing circuit or manual resuscitator). The distal end of the tube is open and bevelled (obliquely cut), atraumatic, with Murphy’s eye. The endotracheal tubes are standard in all aspects: dimension, markings and connectors. * Initial sterilisation method: Ethylene oxide gas or Gamma radiation. * Individually peel-packed in paper and/or plastic. | Tube: Size: 2. Internal diameter: 2.0 mm. External diameter: 3.0 mm. Length: 160 mm, minimum. Connector: Proximal end: Standard internal diameter: 15 mm. Distal end (fits the tube): Length: 9 mm, minimum. | Tube: Polyvinyl chloride (PVC) or polyurethane; medical grade; transparent. Connector: Polyvinyl chloride (PVC); medical grade; transparent or blue. | ISO 5361:2016: Anaesthetic and respiratory equipment — Tracheal tubes and connectors. ISO 10993-1:2018: Biological evaluation of medical devices — Part 1: Evaluation and testing within a risk management process. ISO 11135:2014: Sterilization of health-care products — Ethylene oxide — Requirements for the development, validation and routine control of a sterilization process for medical devices. | Box of 10 units
- | TUBE, ENDOTRACHEAL, No.2.5, w/o cuff, ster., s.u. | Tube, endotracheal, No. 2.5, without cuff, sterile, single use | Idem. | Tube: Size: 2.5. Internal diameter: 2.5 mm. External diameter: 3.5 mm. Length: 140 mm, minimum. Connector: Proximal end: Standard internal diameter: 15 mm. Distal end (fits the tube): Length: 9 mm, minimum. | Tube: Polyvinyl chloride (PVC) or polyurethane; medical grade; transparent. Connector: Polyvinyl chloride (PVC); medical grade; transparent or blue. | ISO 5361:2016: Anaesthetic and respiratory equipment — Tracheal tubes and connectors. ISO 10993-1:2018: Biological evaluation of medical devices — Part 1: Evaluation and testing within a risk management process. ISO 11135:2014: Sterilization of health-care products — Ethylene oxide — Requirements for the development, validation and routine control of a sterilization process for medical devices. | Box of 10 units
- | TUBE, ENDOTRACHEAL, No.3, w/o cuff, ster., s.u. | Tube, endotracheal, No. 3, without cuff, sterile, single use | Idem. | Tube: Size: 3. Internal diameter: 3 mm. External diameter: 4.2 mm. Length: 160 mm, minimum. Connector: Proximal end: Standard internal diameter: 15 mm. Distal end (fits the tube): Length: 9 mm, minimum. | Tube: Polyvinyl chloride (PVC) or polyurethane; medical grade; transparent. Connector: Polyvinyl chloride (PVC); medical grade; transparent or blue. | ISO 5361:2016: Anaesthetic and respiratory equipment — Tracheal tubes and connectors. ISO 10993-1:2018: Biological evaluation of medical devices — Part 1: Evaluation and testing within a risk management process. ISO 11135:2014: Sterilization of health-care products — Ethylene oxide — Requirements for the development, validation and routine control of a sterilization process for medical devices. | Box of 10 units
- | TUBE, ENDOTRACHEAL, No.3.5, w/o cuff, ster., s.u. | Tube, endotracheal, No. 3.5, without cuff, sterile, single use | Idem. | Tube: Size: 3.5. Internal diameter: 3.5 mm. External diameter: 4.8 mm. Length: 140 mm, minimum. Connector: Proximal end: Standard internal diameter: 15 mm. Distal end (fits the tube): Length: 9 mm, minimum. | Tube: Polyvinyl chloride (PVC) or polyurethane; medical grade; transparent. Connector: Polyvinyl chloride (PVC); medical grade; transparent or blue. | ISO 5361:2016: Anaesthetic and respiratory equipment — Tracheal tubes and connectors. ISO 10993-1:2018: Biological evaluation of medical devices — Part 1: Evaluation and testing within a risk management process. ISO 11135:2014: Sterilization of health-care products — Ethylene oxide — Requirements for the development, validation and routine control of a sterilization process for medical devices. | Box of 10 units
- | TUBE, ENDOTRACHEAL, No.4, w/o cuff, ster., s.u. | Tube, endotracheal, No. 4, without cuff, sterile, single use | Idem. | Tube: Size: 4. Internal diameter: 4 mm. External diameter: 5.4 mm. Length: 200 mm, minimum. Connector: Proximal end: Standard internal diameter: 15 mm. Distal end (fits the tube): Length: 11 mm, minimum. | Tube: Polyvinyl chloride (PVC) or polyurethane; medical grade; transparent. Connector: Polyvinyl chloride (PVC); medical grade; transparent or blue. | ISO 5361:2016: Anaesthetic and respiratory equipment — Tracheal tubes and connectors. ISO 10993-1:2018: Biological evaluation of medical devices — Part 1: Evaluation and testing within a risk management process. ISO 11135:2014: Sterilization of health-care products — Ethylene oxide — Requirements for the development, validation and routine control of a sterilization process for medical devices. | Box of 10 units
- | TUBE, ENDOTRACHEAL, No.5, w/o cuff, ster., s.u. | Tube, endotracheal, No. 5, without cuff, sterile, single use | Idem. | Tube: Size: 5. Internal diameter: 5.0 mm. External diameter: 6.9 mm. Length: 250 mm, minimum. Connector: Proximal end: Standard internal diameter: 15 mm. Distal end (fits the tube): Length: 12 mm, minimum. | Tube: Polyvinyl chloride (PVC) or polyurethane; medical grade; transparent. Connector: Polyvinyl chloride (PVC); medical grade; transparent or blue. | ISO 5361:2016: Anaesthetic and respiratory equipment — Tracheal tubes and connectors. ISO 10993-1:2018: Biological evaluation of medical devices — Part 1: Evaluation and testing within a risk management process. ISO 11135:2014: Sterilization of health-care products — Ethylene oxide — Requirements for the development, validation and routine control of a sterilization process for medical devices. | Box of 10 units
- | TUBE, ENDOTRACHEAL, No.4, w/ cuff, ster., s.u. | Tube, endotracheal, No. 4, with cuff, sterile, single use | Endotracheal tube designed to be introduced into the trachea via the mouth or nose, ensuring an unobstructed upper airway to convey gases and vapours to and from the lungs. It consists in a thin, flexible, transparent and single hollow cylinder, with an anatomical curvature Magill-type of 37.5°, black and legibly depth markings and graduation in centimetres, with radio-opaque continuous line mark, with cuff and pilot balloon, with a standard connector in the proximal end. The proximal end of the tube safely fits the connector from which size is selected according to the tube size. The connector is straight and double-ended, with the proximal end being an outer, standard 15 mm internal diameter, conical tip that allows the tube to be connected to the ventilation system (breathing circuit or manual resuscitator). The distal end of the tube is open and bevelled (obliquely cut), atraumatic, with Murphy’s eye. The cuff, situated at the distal end of the tube, provides an airtight seal between the tube and the tracheal wall. It seals the lungs against the liquid secretions sloshing around in the upper airway. And, it ensures that the environment below the cuff can be pressurised and ventilated with a carefully controlled gas mixture. The cuff has a low pressure in order to avoid inadequate pressure on the tracheal mucous membrane to prevent damage or even necrosis. It typically has a capacity of 10 mL. The cuff is inflated via a small-bore inflation tube welded to the outside of the tracheal tube or built into its wall. The pilot balloon indicates the cuff distension. One end is connected to the cuff through a thin inflation tube located close to the proximal end. The other end has a spring-loaded, one-way valve that maintains a pre-set pressure in the circuit, and has Luer tip connector for syringes. The endotracheal tubes are standard in all aspects: dimension, markings and connectors. * Initial sterilisation method: Ethylene oxide gas or Gamma radiation. * Individually peel-packed in paper and/or plastic | Tube: Size: 4. Internal diameter: 4.0 mm. External diameter: 6.7 mm. Length: 210 mm, minimum. Connector: Proximal end: Standard internal diameter: 15 mm. Distal end (fits the tube): Length: 11 mm, minimum. Cuff: Diameter: 10.5 mm. | Tube: Polyvinyl chloride (PVC) or polyurethane; medical grade; transparent. Connector: Polyvinyl chloride (PVC); medical grade; transparent or blue. Cuff: Siliconized rubber, or siliconized polyvinyl chloride (PVC). Pilot balloon: Siliconized latex, or siliconized polyvinyl chloride (PVC). | ISO 5361:2016: Anaesthetic and respiratory equipment — Tracheal tubes and connectors. ISO 10993-1:2018: Biological evaluation of medical devices — Part 1: Evaluation and testing within a risk management process. ISO 11135:2014: Sterilization of health-care products — Ethylene oxide — Requirements for the development, validation and routine control of a sterilization process for medical devices. | Box of 10 units
- | TUBE, ENDOTRACHEAL, No.5, w/ cuff, ster., s.u. | Tube, endotracheal, No. 5, with cuff, sterile, single use | Idem. | Tube: Size: 5. Internal diameter: 5.0 mm. External diameter: 6.7 mm. Length: 250 mm, minimum. Connector: Proximal end: Standard internal diameter: 15 mm. Distal end (fits the tube): Length: 12 mm, minimum. Cuff: Diameter: 13 mm. | Tube: Polyvinyl chloride (PVC) or polyurethane; medical grade; transparent. Connector: Polyvinyl chloride (PVC); medical grade; transparent or blue. Cuff: Siliconized rubber, or siliconized polyvinyl chloride (PVC). Pilot balloon: Siliconized latex, or siliconized polyvinyl chloride (PVC). | ISO 5361:2016: Anaesthetic and respiratory equipment — Tracheal tubes and connectors. ISO 10993-1:2018: Biological evaluation of medical devices — Part 1: Evaluation and testing within a risk management process. ISO 11135:2014: Sterilization of health-care products — Ethylene oxide — Requirements for the development, validation and routine control of a sterilization process for medical devices. | Box of 10 units
- | TUBE, ENDOTRACHEAL, No.6, w/ cuff, ster., s.u. | Tube, endotracheal, No. 6, with cuff, sterile, single use | Idem. | Tube: Size: 6. Internal diameter: 6.0 mm. External diameter: 8.0 mm. Length: 290 mm, minimum. Connector: Proximal end: Standard internal diameter: 15 mm. Distal end (fits the tube): Length: 13 mm, minimum. Cuff: Diameter: 18.5 mm. | Tube: Polyvinyl chloride (PVC) or polyurethane; medical grade; transparent. Connector: Polyvinyl chloride (PVC); medical grade; transparent or blue. Cuff: Siliconized rubber, or siliconized polyvinyl chloride (PVC). Pilot balloon: Siliconized latex, or siliconized polyvinyl chloride (PVC). | ISO 5361:2016: Anaesthetic and respiratory equipment — Tracheal tubes and connectors. ISO 10993-1:2018: Biological evaluation of medical devices — Part 1: Evaluation and testing within a risk management process. ISO 11135:2014: Sterilization of health-care products — Ethylene oxide — Requirements for the development, validation and routine control of a sterilization process for medical devices. | Box of 10 units
- | TUBE, ENDOTRACHEAL, No.7, w/ cuff, ster., s.u. | Tube, endotracheal, No. 7, with cuff, sterile, single use | Idem. | Tube: Size: 7. Internal diameter: 7.0 mm. External diameter: 9.3 mm. Length: 320 mm, minimum. Connector: Proximal end: Standard internal diameter: 15 mm. Distal end (fits the tube): Length: 16 mm, minimum. Cuff: Diameter: 24 mm. | Tube: Polyvinyl chloride (PVC), or polyurethane; medical grade; transparent. Connector: Polyvinyl chloride (PVC); medical grade; transparent or blue. Cuff: Siliconized rubber, or siliconized polyvinyl chloride (PVC). Pilot balloon: Siliconized latex, or siliconized polyvinyl chloride (PVC). | ISO 5361:2016: Anaesthetic and respiratory equipment — Tracheal tubes and connectors. ISO 10993-1:2018: Biological evaluation of medical devices — Part 1: Evaluation and testing within a risk management process. ISO 11135:2014: Sterilization of health-care products — Ethylene oxide — Requirements for the development, validation and routine control of a sterilization process for medical devices. | Box of 10 units
- | TUBE, ENDOTRACHEAL, No.8, w/ cuff, ster., s.u. | Tube, endotracheal, No. 8, with cuff, sterile, single use | Idem. | Tube: Size: 8. Internal diameter: 8.0 mm. External diameter: 10.7 mm. Length: 340 mm, minimum. Connector: Proximal end: Standard internal diameter: 15 mm. Distal end (fits the tube): Length: 16 mm, minimum. Cuff: Diameter: 26 mm. | Tube: Polyvinyl chloride (PVC), or polyurethane; medical grade; transparent. Connector: Polyvinyl chloride (PVC); medical grade; transparent or blue. Cuff: Siliconized rubber, or siliconized polyvinyl chloride (PVC). Pilot balloon: Siliconized latex, or siliconized polyvinyl chloride (PVC). | ISO 5361:2016: Anaesthetic and respiratory equipment — Tracheal tubes and connectors. ISO 10993-1:2018: Biological evaluation of medical devices — Part 1: Evaluation and testing within a risk management process. ISO 11135:2014: Sterilization of health-care products — Ethylene oxide — Requirements for the development, validation and routine control of a sterilization process for medical devices. | Box of 10 units
- | TUBE, ENDOTRACHEAL, No.9, w/ cuff, ster., s.u. | Tube, endotracheal, No. 9, with cuff, sterile, single use | Idem. | Tube: Size: 9. Internal diameter: 9.0 mm. External diameter: 12.0 mm. Length: 350 mm, minimum. Connector: Proximal end: Standard internal diameter: 15 mm. Distal end (fits the tube): Length: 16 mm, minimum. Cuff: Diameter: 28 mm. | Tube: Polyvinyl chloride (PVC) or polyurethane; medical grade; transparent. Connector: Polyvinyl chloride (PVC); medical grade; transparent or blue. Cuff: Siliconized rubber, or siliconized polyvinyl chloride (PVC). Pilot balloon: Siliconized latex, or siliconized polyvinyl chloride (PVC). | ISO 5361:2016: Anaesthetic and respiratory equipment — Tracheal tubes and connectors. ISO 10993-1:2018: Biological evaluation of medical devices — Part 1: Evaluation and testing within a risk management process. ISO 11135:2014: Sterilization of health-care products — Ethylene oxide — Requirements for the development, validation and routine control of a sterilization process for medical devices. | Box of 10 units
- | Laryngeal mask airway (LMA), size 2, sterile, single use | Laryngeal mask airway (LMA), size 2, sterile, single use | Standard laryngeal mask airway used for patients undergoing general anaesthesia or as a resuscitation device. Maximum cuff volume: 10 mL. |  | Flexible medical grade PE. Latex-free. | ISO 11712:2009: Anaesthetic and respiratory equipment — Supralaryngeal airways and connectors. ISO 10993-1:2018: Biological evaluation of medical devices — Part 1: Evaluation and testing within a risk management process. ISO 11135:2014: Sterilization of health-care products — Ethylene oxide — Requirements for the development, validation and routine control of a sterilization process for medical devices. | Box of 5 units
- | Laryngeal mask airway (LMA), size 3, sterile, single use | Laryngeal mask airway (LMA), size 3, sterile, single use | Standard laryngeal mask airway used for patients undergoing general anaesthesia or as a resuscitation device. Maximum cuff volume: 20 mL. |  | Flexible medical grade PE. Latex-free. | Idem. | Idem.
- | Laryngeal mask airway (LMA), size 4, sterile, single use | Laryngeal mask airway (LMA), size 4, sterile, single use | Standard laryngeal mask airway used for patients undergoing general anaesthesia or as a resuscitation device. Maximum cuff volume: 30 mL. |  | Flexible medical grade PE. Latex-free. | Idem. | Idem.
- | Colorimetric End Tidal CO2 detector, adult, s.u. | Colorimetric End Tidal CO2 detector, adult, single use | Used to confirm proper endotracheal tube (ETT) placement by assessing exhaled CO2. It is attached directly to the ETT and changes its colour as a response of the exhaled CO2. Internal volume: 25 cc. Resistance to flow: 4.4 cm H2O ± 1.0 cm at 60 L/min flow. Detector weight: Less than 20 g. Connector port, patient end: 22 mm O.D./15 mm I.D. Connector port, circuit end: 15 mm O.D./13 mm I.D. Color chart with three ranges of EtCO2. | Size: Patient weight more than 15 kg. |  |  | Box of 24 units
- | Colorimetric End Tidal CO2 detector, paediatric, s.u. | Colorimetric End Tidal CO2 detector, paediatric, single use | Used to confirm proper endotracheal tube (ETT) placement by assessing exhaled CO2. It is attached directly to the ETT and changes its colour as a response of the exhaled CO2. Internal volume: 3 cc. Resistance to flow: 2.5 cm H2O ± 0.5 cm at 10 L/min flow. Detector weight: Less than 5 g. Connector port, patient end: 18 mm O.D./15 mm I.D. Connector port, circuit end: 15 mm O.D./5 mm I.D. Color chart with three ranges of EtCO2. | Size: Patient weight less than 15 kg. |  |  | Idem.
- | SYRINGE, Luer slip, 10mL, sterile, single use | Syringe, 10 mL, two or three pieces, Luer slip, sterile, single use | Syringe options: a) Two pieces: barrel with Luer nozzle and piston. b) Or, three pieces: barrel with Luer nozzle, piston, and stopper. * Barrel: Permanent and legible graduated scale in mL with intervals of 0.20 or 0.50 mL. Increment of each mL to be numbered. Concentric Luer slip nozzle. Length with a maximum usable capacity of at least 10% more than the nominal capacity. * Plunger stopper with backstop. Double sealing ring on plunger. * Initial sterilisation method: Ethylene oxide gas. * Individually peel-packed in paper and/or plastic. | Barrel nominal capacity: 10 mL. | Barrel: Polyethylene (PE) or polypropylene (PP) or polystyrene (PS); sufficiently transparent. Piston: Polypropylene (PP). *Latex-free/PVC-free. | Barrel graduation: ISO 7886-1:2017: Sterile hypodermic syringe for single use: part 1: syringe for manual use. Needle, Luer type: ISO 80369-7:2016: Small-bore connectors for liquids and gases in healthcare applications -- Part 7: Connectors for intravascular or hypodermic applications. Sterilization method: ISO 11135:2014: Sterilization of health-care products -- Ethylene oxide -- Requirements for the development, validation and routine control of a sterilization process for medical devices. | Box of 100 units
Same description is for both jelly, presentation is different for practicality in the application. | LUBRICATING Jelly, 5g, box of 150 | LUBRICATING Jelly, 5 g, pack | Designed to assist with patient intubation and feeding tubes. |  |  |  | Box of 150
- | FORCEPS, MAGILL, paediatric, 15 cm | FORCEPS, MAGILL, 15 cm | Designed to assist with patient intubation. Made of stainless steel. |  |  |  | Each
- | FORCEPS, MAGILL, child, 19 cm | FORCEPS, MAGILL, 19 cm | Idem. |  |  |  | Idem.
- | FORCEPS, MAGILL, adult, 24 cm | FORCEPS, MAGILL, 24 cm | Idem. |  |  |  | Idem.
- | Suction bulb, for newborn, reusable, autoclavable | Suction bulb, for newborn, reusable, autoclavable | Portable, hand-held, manual suction device. Used to evacuate secretions and liquids from the nasal cavity or from high airways of infants. The top part can be opened and is attached with a strip to the body to permit easier cleaning. Can be subjected to boiling (HLD) and sterilization including autoclaving. |  | Latex-free. Translucent. It must resist steam sterilization at 121 °C or 134 °C. | ISO 10079-2:1999: Medical suction equipment - Part 2: Manually powered suction equipment. JIS T 7208-2:2005: Medical suction equipment - Part 2: Manually powered suction equipment (Japan). ISO 11135-1:2007: Sterilization of health care products - Ethylene oxide - Part 1: Requirements for development, validation and routine control of a sterilization process for medical devices. | Each
- | Filter, heat and moisture exchanger (HMEF), high efficiency, with connectors, adult, single use | Heat and moisture exchanger filter (HMEF), high efficiency, straight, adult, single use | Heat and moisture exchanging filters (HMEF) for respiratory breathing systems. 99.9% efficiency; humidification level >30 mg H2O/L; dead space <40 mL. Connector 22 Fr / 15 mm. Transparent cover. With luer connector. | Maximum tidal volume (breathing filter): 1500 mL. |  | ISO 9360-2:2001: Anaesthetic and respiratory equipment — Heat and moisture exchangers (HMEs) for humidifying respired gases in humans — Part 2: HMEs for use with tracheostomized patients having minimum tidal volumes of 250 ml. ISO 5367:2014: Anaesthetic and respiratory equipment — Breathing sets and connectors. | Each
- | Filter, heat and moisture exchanger (HMEF), high efficiency, with connectors, paediatric, single use | Heat and moisture exchanger filter (HMEF), high efficiency, straight, paediatric, single use | Idem. | Maximum tidal volume (breathing filter): 500 mL. |  | Idem. | Idem.
If few suppliers, it could be replaced by other option of gastroenteral feeding tubes (i.e. conical tip and double channel) with its respective feeding syringe conical tip. | Tube, feeding, nasogastric, 10 Fr, 50 cm, ENFit tip, sterile, single use | Tube, feeding, nasogastric, 10 Fr, 50 cm, ENFit tip, sterile, single use | Used for short term gastro-enteral feeding and drug administration when connected to feeding syringes; or, for ventricular lavage. The nasogastric tube is introduced via the nasopharynx into the gastrointestinal (GI) tract. It consists in a thin, flexible, transparent and single hollow cylinder with radio-opaque line marked from the distal end. The distal end is a soft and rounded closed-ended tip, with two laterals opposite alternated eyelets. The proximal end with a connector, ENFit tip, and a stopper, allows the tube to be connected to feeding syringes. Visible international-recognized colour code on cup connector. * Initial sterilisation method: Ethylene oxide gas. * Individually peel-packed in paper and/or plastic. | Size: 10 Fr. Length: 50 cm. | Polyvinyl chloride (PVC); medical grade. | ISO 80369-3:2016: Small-bore connectors for liquids and gases in healthcare applications — Part 3: Connectors for enteral applications. ISO 18250-3:2018: Medical devices — Connectors for reservoir delivery systems for healthcare applications — Part 3: Enteral applications. ISO 11135:2014: Sterilization of health-care products — Ethylene oxide — Requirements for the development, validation and routine control of a sterilization process for medical devices. | Box of 50 units
- | Tube, feeding, nasogastric, 12 Fr, 90 cm, ENFit tip, sterile, single use | Tube, feeding, nasogastric, 12 Fr, 90 cm, ENFit tip, sterile, single use | Idem. | Size: 12 Fr. Length: 90 cm. | Idem. | Idem. | Idem.
- | Tube, feeding, nasogastric, 14 Fr, 90 cm, ENFit tip, sterile, single use | Tube, feeding, nasogastric, 14 Fr, 90 cm, ENFit tip, sterile, single use | Idem. | Size: 14 Fr. Length: 90 cm. | Idem. | Idem. | Idem.
- | Tube, feeding, nasogastric, 6 Fr, 50 cm, ENFit tip, sterile, single use | Tube, feeding, nasogastric, 6 Fr, 50 cm, ENFit tip, sterile, single use | Idem. | Size: 6 Fr. Length: 50 cm. | Idem. | Idem. | Idem.
- | Tube, feeding, nasogastric, 8 Fr, 50 cm, ENFit tip, sterile, single use | Tube, feeding, nasogastric, 8 Fr, 50 cm, ENFit tip, sterile, single use | Idem. | Size: 8 Fr. Length: 50 cm. | Idem. | Idem. | Idem.
- | Syringe, feeding, 1 mL, LDT, ENFit, sterile, single use | Syringe, feeding, 1 mL, LDT, ENFit, sterile, single use | Low Dose Tip (LDT) syringe used for the administration of oral medication or delivery of enteral nutrition by connection to an enteral administration device. It consists of a calibrated hollow barrel (cylinder) made of plastic and a moveable plunger. The tip is designed to mate only with enteral administration devices and is specifically incompatible with Luer (6%) connectors and intravenous devices. * Initial sterilisation method: Ethylene oxide gas. * Individually peel-packed in paper and/or plastic. | * Barrel: 1 mL; graduations every 0.1 mL. | * Barrel: Transparent polypropylene, ending with a female ENFit tip, LDT to decrease the dead space. * Plunger: Polypropylene or polyethylene coloured to be distinguished from injection syringes. * Gasket: Synthetic elastomer (latex free). | ISO 80369-3:2016: Small-bore connectors for liquids and gases in healthcare applications — Part 3: Connectors for enteral applications. ISO 18250-3:2018: Medical devices — Connectors for reservoir delivery systems for healthcare applications — Part 3: Enteral applications. ISO 11135:2014: Sterilization of health-care products — Ethylene oxide — Requirements for the development, validation and routine control of a sterilization process for medical devices. | Each
- | Syringe, feeding, 10 mL, ENFit, sterile, single use | Syringe, feeding, 10 mL, ENFit, sterile, single use | Syringe used for the administration of oral medication or delivery of enteral nutrition by connection to an enteral administration device. It consists of a calibrated hollow barrel (cylinder) made of plastic and a moveable plunger. The tip is designed to mate only with enteral administration devices and is specifically incompatible with Luer (6%) connectors and intravenous devices. * Initial sterilisation method: Ethylene oxide gas. * Individually peel-packed in paper and/or plastic. | * Barrel: 10 mL; graduations every 0.5 mL. | * Barrel: Transparent polypropylene, ending with a female ENFit tip. * Plunger: Polypropylene or polyethylene coloured to be distinguished from injection syringes. * Gasket: Synthetic elastomer (latex free). |  | Each
- | Syringe, feeding, 2.5 mL, LDT, ENFit, sterile, single use | Syringe, feeding, 2.5 mL, LDT, ENFit, sterile, single use | Low Dose Tip (LDT) syringe used for the administration of oral medication or delivery of enteral nutrition by connection to an enteral administration device. It consists of a calibrated hollow barrel (cylinder) made of plastic and a moveable plunger. The tip is designed to mate only with enteral administration devices and is specifically incompatible with Luer (6%) connectors and intravenous devices. * Initial sterilisation method: Ethylene oxide gas. * Individually peel-packed in paper and/or plastic. | * Barrel: 2.5 to 3 mL; graduations every 0.1 mL. | * Barrel: Transparent polypropylene, ending with a female ENFit tip, LDT to decrease the dead space. * Plunger: Polypropylene or polyethylene coloured to be distinguished from injection syringes. * Gasket: Synthetic elastomer (latex free). | ISO 80369-3:2016: Small-bore connectors for liquids and gases in healthcare applications — Part 3: Connectors for enteral applications. ISO 18250-3:2018: Medical devices — Connectors for reservoir delivery systems for healthcare applications — Part 3: Enteral applications. ISO 11135:2014: Sterilization of health-care products — Ethylene oxide — Requirements for the development, validation and routine control of a sterilization process for medical devices. | Each
- | Syringe, feeding, 20 mL, ENFit, sterile, single use | Syringe, feeding, 20 mL, ENFit, sterile, single use | Syringe used for the administration of oral medication or delivery of enteral nutrition by connection to an enteral administration device. It consists of a calibrated hollow barrel (cylinder) made of plastic and a moveable plunger. The tip is designed to mate only with enteral administration devices and is specifically incompatible with Luer (6%) connectors and intravenous devices. * Initial sterilisation method: Ethylene oxide gas. * Individually peel-packed in paper and/or plastic. | * Barrel: 20 mL; graduations every 0.5 mL. | * Barrel: Transparent polypropylene, ending with a female ENFit tip. * Plunger: Polypropylene or polyethylene coloured to be distinguished from injection syringes. * Gasket: Synthetic elastomer (latex free). |  | Each
- | Syringe, feeding, 5 mL, LDT, ENFit, sterile, single use | Syringe, feeding, 5 mL, LDT, ENFit, sterile, single use | Low Dose Tip (LDT) syringe used for the administration of oral medication or delivery of enteral nutrition by connection to an enteral administration device. It consists of a calibrated hollow barrel (cylinder) made of plastic and a moveable plunger. The tip is designed to mate only with enteral administration devices and is specifically incompatible with Luer (6%) connectors and intravenous devices. * Initial sterilisation method: Ethylene oxide gas. * Individually peel-packed in paper and/or plastic. | * Barrel: 5 mL; graduations every 0.2 mL. | * Barrel: Transparent polypropylene, ending with a female ENFit tip, LDT to decrease the dead space. * Plunger: Polypropylene or polyethylene coloured to be distinguished from injection syringes. * Gasket: Synthetic elastomer (latex free). | ISO 80369-3:2016: Small-bore connectors for liquids and gases in healthcare applications — Part 3: Connectors for enteral applications. ISO 18250-3:2018: Medical devices — Connectors for reservoir delivery systems for healthcare applications — Part 3: Enteral applications. ISO 11135:2014: Sterilization of health-care products — Ethylene oxide — Requirements for the development, validation and routine control of a sterilization process for medical devices. | Each
- | Syringe, feeding, 60 mL, ENFit, sterile, single use | Syringe, feeding, 60 mL, ENFit, sterile, single use | Syringe used for the administration of oral medication or delivery of enteral nutrition by connection to an enteral administration device. It consists of a calibrated hollow barrel (cylinder) made of plastic and a moveable plunger. The tip is designed to mate only with enteral administration devices and is specifically incompatible with Luer (6%) connectors and intravenous devices. * Initial sterilisation method: Ethylene oxide gas. * Individually peel-packed in paper and/or plastic. | * Barrel: 60 mL; graduations every 1 mL. | * Barrel: Transparent polypropylene, ending with a female ENFit tip. * Plunger: Polypropylene or polyethylene coloured to be distinguished from injection syringes. * Gasket: Synthetic elastomer (latex free). |  | Each
Same description is for both jelly, presentation is different for practicality in the application. | LUBRICATING Jelly, 50g, tube | LUBRICATING Jelly, 50 g, tube | Designed to assist with patient intubation and feeding tubes.
- | Pad, absorbent | Pad, absorbent, single use | Absorbent pad used to contain body fluids and any water-based fluid. Absorbs roughly 300 mL, depending on the viscosity of the liquid. Latex-free and lint-free. Dimensions: minimum 30 x 25 cm.
- | Basin kidney, stainless steel, 825 mL | Basin kidney, stainless steel, 825 mL | Made of stainless steel.
- | Stethoscope, binaural, double cup, adult/child, single use | Stethoscope, binaural, double cup, adult/child, single use | Single use stethoscope.
For e.g. ARROW. | Central Venous Catheters kit | Central Venous Catheters kit | Supplied with: Finder needle, syringe, wire, dilator, lidocaine, scalpel, needle, thread.
For e.g. Tegaderm. | Transparent adhesive plasters, wash proof, 5x5 cm | Transparent adhesive plasters, wash proof, 5x5 cm | Transparent adhesive plaster.
- | Bag, collecting, urine, with outlet tap, with non-return valve, 2000 mL, adult, non-sterile, single use | Bag, collecting, urine, with outlet tap, with non-return valve, 2000 mL, adult, non-sterile, single use | Used to fit gastric tube (aspirating/feeding tube) or incontinence condom. Drainable bag for collecting urine with permanent and legible graduations every 100 mL. With reinforced eyelets for hanging. With a drainage valve (outlet tap) which permits the bag to be emptied without disconnecting, maintaining sterility. It is fitted to an outlet tube. With a non-return valve, located inside the urine collector bag at the upper part (urine entry point), which prevents urine backflow into the indwelling urinary catheter. It is fitted with a kink resistant, transparent plastic inlet tube, with universal connector and protective cap. | Capacity, bag: 2000 mL. Length, tube: 85 - 95 cm. Diameter, tube: 6.5 mm, approximately. | Bag: Polyvinyl chloride (PVC), polypropylene or ethylene vinyl acetate (EVA); medical grade. Tube, connector and protective cap: Polyvinyl chloride (PVC); medical grade. | ISO 8669-2:1996: Urine collection bags -- Part 2: Requirements and test methods. | Box of 10 units
- | Catheter, urethral, Foley, 2-way, 10 Fr, sterile, single use | Catheter, urethral, Foley, 2-way, 10 Fr, sterile, single use | A soft thin rubber tube with a balloon at the nelaton tip, designed for insertion in the bladder cavity, via the urethra, in order to drain off urine, instil a liquid or irrigate the bladder. A standard catheter consists of a hollow 2-way cylindrical tube with one central channel for urinary drainage; bladder side ending, Foley type, with a rounded atraumatic end tip (nelaton tip), with two opposing eyelets and one balloon; collector side ending that is a universal and hollow truncated cone (funnel) to connect the urine bag, spigot, syringe or irrigating device; and the balloon port, side channel ending with a non-return valve and a Luer tip connector. Catheter size is expressed in French gauge or Charrier (Fr or CH) and colour coded; balloon expansion capacity is expressed in mL. Dimensions and colour code must be legible and visible on the connectors. Preferable intended use for this size: children, medium or long term catheterization. * Initial sterilisation method: Ethylene oxide gas. * Individually peel-packed in paper and/or plastic. * Double-packaged: protected with an interior layer and an outer peel pack. | Length: 30 cm. External diameter: 10 Fr. Balloon expansion capacity: 3 to 5 mL. | Silicone-coated natural latex. | ISO 20696:2018: Sterile urethral catheters for single use. ISO 11135:2014: Sterilization of health-care products — Ethylene oxide — Requirements for the development, validation and routine control of a sterilization process for medical devices. | Box of 10 units
- | Catheter, urethral, Foley, 2-way, 12 Fr, sterile, single use | Catheter, urethral, Foley, 2-way, 12 Fr, sterile, single use | Idem. | Length: 40 cm. External diameter: 12 Fr. Balloon expansion capacity: 5 to 15 mL. | Idem. | Idem. | Idem.
- | Catheter, urethral, Foley, 2-way, 14 Fr, sterile, single use | Catheter, urethral, Foley, 2-way, 14 Fr, sterile, single use | Idem. | Length: 40 cm. External diameter: 14 Fr. Balloon expansion capacity: 5 to 15 mL. | Idem. | Idem. | Idem.
- | Catheter, urethral, Foley, 2-way, 16 Fr, sterile, single use | Catheter, urethral, Foley, 2-way, 16 Fr, sterile, single use | Idem. | Length: 40 cm. External diameter: 16 Fr. Balloon expansion capacity: 5 to 15 mL. | Idem. | Idem. | Idem.
- | Catheter, urethral, Foley, 2-way, 18 Fr, sterile, single use | Catheter, urethral, Foley, 2-way, 18 Fr, sterile, single use | Idem. | Length: 40 cm. External diameter: 18 Fr. Balloon expansion capacity: 5 to 15 mL. | Idem. | Idem. | Idem.
- | Catheter, urethral, Foley, 2-way, 20 Fr, sterile, single use | Catheter, urethral, Foley, 2-way, 20 Fr, sterile, single use | Idem. | Length: 40 cm. External diameter: 20 Fr. Balloon expansion capacity: 5 to 15 mL. | Idem. | Idem. | Idem.
- | Catheter, urethral, Foley, 2-way, 8 Fr, sterile, single use | Catheter, urethral, Foley, 2-way, 8 Fr, sterile, single use | Idem. | Length: 30 cm. External diameter: 8 Fr. Balloon expansion capacity: 3 to 5 mL. | Idem. | Idem. | Idem.
- | Compress, gauze, 10 x 10 cm, 8 to 12 ply, sterile, single use | Compress, gauze, 10 x 10 cm, 8 to 12 ply, sterile, single use | Used for general wound care, wound cleaning and wound covering; to treat wounds with minimal secretion. Absorbent gauze with minimum 8 to 12 folds to ensure better absorption. Edges folded over; with selvedge to prevent the gauze from fraying. No broken stitches, no loose threads. High level of tear-resistance. Air-permeable. Radiolucent. Hypoallergenic. Pre-washed. Width: 10 cm. Length: 10 cm. Material: 100% cotton fabric: hydrophobic, purified and bleached, plain weave. Type of gauze: V / 17 threads/cm2 Warp: 95 to 105 threads/dm. Weft: 66 to 74 threads/dm. Grammage: +/- 23g/m2. Initial sterilisation method: Ethylene oxide gas or ionising radiation. Packed per 2 or 5 in a sterile peel open pack. |  |  | EN 14079: Non-active medical devices - Performance requirements and test methods for absorbent cotton gauze and absorbent cotton and viscose gauze. ISO 10993-10:2010: Biological evaluation of medical devices — Part 10: Tests for irritation and skin sensitization. | Box of 50 units
- | Tape, surgical, hypoallergenic, 5 x 2.5 cm | Tape, surgical, hypoallergenic, 2.5 x 5 m | Latex free hypoallergenic. Porous and breathable. Radiolucent. Length: 5 to 10 m.
- | Drape, surgical, non woven, sterile, single use | Drape, surgical, non woven, sterile, single use | Non-woven surgical drape.
- | Gloves, examination, nitrile, powder-free, pair-packed, large, sterile, single use | Gloves, examination, nitrile, powder-free, pair-packed, large, sterile, single use | Sterile, pair-packed. Smooth surface, flanged cuff, moulded manufacture. Thickness, fingers and palm: 0.08 - 0.12 mm (3 - 4.7 mil). Nitrile (synthetic rubber), powder-free, blue or purple. Shelf life: 5 years. |  |  | ISO 11193-2:2006: Single-use medical examination gloves. BS EN 455-3:2015: Medical gloves for single use. Requirements and testing for biological evaluation.
- | Gloves, examination, nitrile, powder-free, pair-packed, medium, sterile, single use |  | Idem. |  |  | Idem.
- | Gloves, examination, nitrile, powder-free, pair-packed, small, sterile, single use |  | Idem. |  |  | Idem.
- | Antiseptic Wipe with Alcohol & Chlorhexidine |  | Box of square wipes packed individually for disinfection with 2% (or 4%) chlorhexidine gluconate and 70% isopropyl alcohol.
Move together with blood sampling kit for the I-Stat, but leave quantities as such for the time being. | Arterial Blood Sample Kits |  | Arterial blood sampling syringes with dry lithium heparin [23.5 iu/mL]. |  |  |  | Box of 100 units
- | TUBE, silicone, autoclavable, int. diam.8 mm, 10 m | Tube, suction, universal, ID 8 mm, L 10 m, autoclavable | Used as a suction tube for continuous or intermittent aspiration during surgery. Uniform diameter; non-permanent deformation if kinked or bent too tight. It is cut according to the length required between the surgery field and the suction machine. The material resist vacuum pressure of 559 mmHg. | Internal diameter: 8 mm [+/1.0 ] External diameter: 10 mm [+/1.0 ] Length: 10 m. | Polycarbonate; medical grade; transparent opaque. It must be capable to resist cleaning and disinfection with chlorine solution (0.05% concentration). It must resist steam sterilization at 121 °C or 134 °C. | ISO 80369-3:2016: Small-bore connectors for liquids and gases in healthcare applications — Part 3: Connectors for enteral applications. | Each
Same description is for both gels. | (ECG) ELECTRODE GEL, bottle | Electro-conductive gel, bottle | Water-based hypoallergenic gel that is capable to be used in applications such as electrocardiogram (ECG), electroencephalogram (EEG), Ultrasound, Transcutaneous Electrical Nerve Stimulation (TENS). The gel is specially formulated to act as a coupling agent and to reduce static.
Same description is for both gels. | (ultrasound transducer) CONDUCTIVE GEL, 5L, container | Electro-conductive gel, bottle | Water-based hypoallergenic gel that is capable to be used in applications such as electrocardiogram (ECG), electroencephalogram (EEG), Ultrasound, Transcutaneous Electrical Nerve Stimulation (TENS). The gel is specially formulated to act as a coupling agent and to reduce static.
- | Indicator, TST (Time, Steam, Temperature), Type 5 | Indicator, TST (Time, Steam, Temperature), Type 5 | TST (Time, Steam, Temperature) integrating chemical indicator (CI), type 5. Self-adhesive, coloured, to monitor steam sterilisation cycles. For use in steam sterilizers working at 121 °C during 15 minutes; or 134 °C during 7 minutes, with saturated steam. One temperature in between – that mimic the death curve, or biological response, of the microorganisms inside a BI (biological indicator) when exposed to ideal, saturated steam conditions. Labelling of the strip is conforming the standards. It is delivered as a pack with one record sheet. Shelf life: 3 years. |  |  | ISO 11140-1:2014: Sterilization of health care products — Chemical indicators — Part 1: General requirements. ISO 18472:2018: Sterilization of health care products — Biological and chemical indicators — Test equipment. | Box of 500 units
- | Paper sheet, crepe, for sterilization, 60 g/m2, 90 x 90 cm | Paper sheet, crepe, for sterilization, 60 g/m2, 90 x 90 cm | Crepe paper is used to wrap devices that will be subjected to sterilization. Made of cellulose fibres allows low flexibility and drapeability. Permeable to steam; efficient microbiological barrier; low permeability to air. Colour: white, green or blue. Non sterile, single use. Storage humidity: Between 30 - 60%. Storage temperature: Between 10 and 40 °C. | Size: 90 x 90 cm. | Grammage: 60 g/m2. Porosity: 50 µ, maximum. Water repellence, test time: 20 seconds, minimum. Tensile strength. | ISO 11607-1:2019: Packaging for terminally sterilized medical devices — Part 1: Requirements for materials, sterile barrier systems and packaging systems. | Box of 250 sheets
- | Paper sheet, non-woven, for sterilization, 120 x 120 cm | Paper sheet, non-woven, for sterilization, 120 x 120 cm | Non-woven paper is used to wrap devices that will be subjected to sterilization. Made of blend of cellulose fibres and synthetic fibres allows good flexibility and drapeability. Permeable to steam; efficient microbiological barrier; good permeability to air. Colour: blue. Non sterile, single use. Storage humidity: Between 30 - 60%. Storage temperature: Between 10 and 40 °C. | Size: 120 x 120 cm. |  | ISO 11607-1:2019: Packaging for terminally sterilized medical devices — Part 1: Requirements for materials, sterile barrier systems and packaging systems. | Box of 250 sheets
- | Tape, adhesive, indicator Type 1, for steam sterilizer, 18 mm x 48 m | Tape, adhesive, indicator Type 1, for steam sterilizer, 18 mm x 48 m | Process chemical indicator (CI) tape, type 1. Beige adhesive crepe paper with white diagonal stripes. When exposed to steam at 121°C for 3 minutes or 134°C for 30 seconds, the diagonal white ink stripes will irreversibly change colour to black. It only provides visual confirmation that a pack has been exposed to steam pressure. One roll with adequate protection and labelled "Autoclave" inside the spool. Shelf life: 3 years. | Length: 48 to 52 mm. Width: 18 to 20 mm. |  | ISO 11140-1:2014: Sterilization of health care products — Chemical indicators — Part 1: General requirements. | Each
Could be purchased locally. | Kerosene stove, electric heating plate | Stove with one heating plate | A heat source, preferably electric, with one heating plate. |  |  |  | Each
Could be purchased locally. | Timer | Timer | Stopwatch to be used in the sterilization room. |  |  |  | Each
This item is in the Drug list (latest version). |  | Spacer, for metered dose inhaler, medium size mask | External device attached to a metered dose inhaler to for better aerosol drug delivery by enhancing actuation and inhalation coordination. The mouthpiece has a one-way valve. The chamber is easy to clean. Possibility of adult and paediatric use. Supplied with one silicone mask, medium size. Single patient use. |  |  |  | Each