## Slide 1: Westgard Multirule System
a “multi-rule” system developed by Dr. James O. Westgard based on statistical concepts
a combination of decision criteria or rules to assess if a system is in control
use when at least 2 levels of control are run with the examination run
cannot use with only one control
Dr. Westgard
Notes: 8-5
<number>

## Slide 2: Westgard Multirule System Titles
12S rule
13S rule
22S rule
R4S rule
41S rule
10X rule
Used when 2 levels of control material are analyzed per run.
Notes: 8-5
<number>

## Slide 3: Westgard – 12S Rule
WARNING RULE – not cause for rejecting a run
One of two control results falls outside ±2SD
Alerts technologist to possible problems
Must then evaluate the 13S rule
Notes: 8-5
<number>

## Slide 4: Westgard – 22S Rule
Quantitative QC - Module 8
Two consecutive control values for the same level fall outside of ±2SD in the same direction, or
Both controls in the same run exceed ±2SD.
Patient results cannot be reported.
Requires corrective action
Notes: 8-5
<number>

## Slide 5: Westgard – R4S Rule
One control exceeds the mean by –2SD, and the other control exceeds the mean by +2SD.
The range between the two results will therefore exceed 4 SD.
Random error has occurred, test run must be rejected.
Notes: 8-5
<number>

## Slide 6: Westgard – 41S Rule
requires control data from previous runs
four consecutive QC results for one level of control are outside ±1SD, or
both levels of control have consecutive results that are outside ±1SD.
Notes: 8-5
<number>

## Slide 7: Westgard – 10X Rule
requires control data from previous runs
ten consecutive QC results for one level of control are on one side of the mean, or
both levels of control have five consecutive results that are on the same side of the mean
Notes: 8-5
<number>

## Slide 8: Westgard Multirule QC
Notes: 8-5
<number>

## Slide 9: Activity
Using the Westgard Multirule System, describe the control values for each day in the following slide.
Notes: Activity 3
<number>

## Slide 11: Answers for activity
Day 21, 22, 24, 26, 27, 30, 31, 33, 34, 36-44 – in control
Day 23, 28, 29 – 12s
Day 25 - 13s
Day 32 – 22s
Day 35 - R4s
Notes: Results of activity- verify
<number>