Activity 10-1: Scenario—Discuss Main Elements of EQA

| Purpose | The purpose of this activity is to discuss the main elements of external quality assessment based on the experience of the participants. |
| --- | --- |
| Suggested Time | 20 minutes |
| Instructional Method | Problem solving: use as an individual or group activity, or as a class discussion. |
| Instructions | Give the participants a copy of the following scenario and/or show it as a slide (Presentation 10, slide 3). Allow 2 minutes for them to read. “A newly appointed laboratory supervisor has begun reviewing past EQA results, and observes poor Gram stain performance had occurred several times.” Ask the question: How should the laboratory investigate this problem? |
| Expected Results | Write down the answers on the whiteboard grouping them as steps to investigate the problem. These include: - examine and evaluate condition of sample when received and how it is handled; - evaluate compliance with procedures; - repeat quality control check on stains; - check equipment to be sure it is properly calibrated and functioning well; - evaluate personnel competency; - carefully examine the reporting process, for example, checking for transcription errors. Examples of cause of errors follow: Pre-examination - sample may have been compromised during preparation, shipping, or after receipt by improper storage or handling; - improper sample management: identification problems, mislabelling. Examination - staff is not competent; - stains may be outdated, not well prepared, not producing the appropriate stain reaction; - microscope not well maintained or misused; - SOP not existing and/or not followed (test methods, quality control). Post-examination (frequent) - wrong report format; - clerical/transcription error. Refer to content sheet 10-1. Questions that might be raised: - What are the implications for patient test results performed and reported during the same time-period? Important implications could be raised for the patient test results. After having taken corrective actions and solved the problem, laboratory has to re-analyze if necessary the patient samples performed during the same time- period and inform if results for the concerned people are different. - What corrective actions might be taken? - provide training to the staff to assure consistent application; - prepare/improve SOPs; - prepare/purchase new reagents; - improve/implement quality control; - improve sample management process; - establish microscope maintenance program. |

Activity 10-2: Processing Proficiency Testing Samples

| Purpose | The purpose of this activity is to discuss how the laboratory processes proficiency testing samples. |
| --- | --- |
| Suggested Time | 15 minutes |
| Instructional Method | Group assignment (pairs). |
| Instructions | Provide instructions to the participants : “In groups of two, consider the following 2 questions related to proficiency testing samples : - How does your laboratory process proficiency testing samples? - Is it possible in your laboratory to process proficiency testing samples in the same way as other patient samples? Base your answer either on what is in place in your laboratory or on what you think should be in place. Provide as many elements as possible. You have 10 minutes.” After 10 minutes, ask several pairs to describe their answers. Then ask if the other pairs have anything different to add. |
| Expected Results | Points for consideration include: - Why is it difficult to handle PT sample in the same way as patient samples? - Can PT samples be blinded from the laboratory staff? - Is it appropriate to test PT sample several times? - Can you discuss PT results with another laboratory before sending the report to the PT provider? Refer to content sheet 10-2. |