## Slide 1: Microscope Maintenance

## Slide 2: Learning Objectives
At the end of this module, you will be able to
Store the microscope properly
Clean the microscope including the eye pieces and objective lenses
Replace the microscope bulb
Troubleshoot the common problems associated with use of microscope

## Slide 3: Content Overview
Microscope storage conditions
Materials required for Microscope Maintenance
Cleaning solutions and solvents required to clean optics, stage, and body
Materials required for cleaning

## Slide 4: Microscope Storage
Proper storage of the microscope will prevent or reduce problems!
Optics and mechanisms of the microscope must be protected from:
  Dust and dirt
  Fungus
Store the microscope
  Under a protective cover
  In a low humidity environment

## Slide 5: Cleaning Solutions and Solvents
Soap solution for cleaning of body and stage
Ethyl ether-alcohol, alcohol, or lens cleaner solution for cleaning of lenses
Refer to manufacturer’s guide for appropriate organic solvent

## Slide 6: Cleaning Solutions
Cleaning Solutions | Frequent Use | Infrequent Use
Manufacturer's recommendation | Yes | Yes
Ethyl ether/alcohol (80/20) | Yes | Yes
Alcohol | No | Yes
Benzene / petrol | No | Yes
Xylol | No | No

## Slide 7: Cleaning Materials
Lint-free cotton gauze pads
Lint-free cotton swabs
Lens paper
  Alternatives include:
    Fine quality tissue paper
    Muslin cloth
    Silk

## Slide 8: Microscope Cleaning Process
Eyepiece
Objectives
Microscope Stage
Microscope Body
Condenser

## Slide 9: Step 1: Cleaning the Eyepieces
Blow to remove dust before wiping lens
Clean the eyepieces with a cotton swab moistened with lens cleaning solution
Clean in a circular motion inside out

## Slide 10: Step 2: Cleaning the Eyepiece
Wipe the eyepieces dry with lens paper
Repeat cleaning and drying if required

## Slide 11: Step 3: Cleaning the Objectives
Objectives are cleaned while attached to microscope
  Moisten the lens paper with the cleaning solution
  Wipe gently the objective in circular motion from inside out
  Wipe with dry tissue or lens cleaning paper
Objectives should never be removed from the nosepiece.

## Slide 12: Step 4: Cleaning the Microscope Stage
Wipe the microscope stage using the cleaning solution on a soft cloth
Thoroughly dry the stage
Repeat above steps, if required

## Slide 13: Step 5: Cleaning the Microscope Body
Unplug the microscope from power source
Moisten the cotton pad with a mild cleaning agent
Wipe the microscope body to remove dust, dirt, and oil
Repeat steps1–3, if required

## Slide 14: Step 6: Cleaning the Condenser and Auxiliary Lens
Unplug the microscope from power source
Clean the condenser lens and auxiliary lens using lint-free cotton swabs moistened with lens cleaning solution
Wipe with dry swabs

## Slide 15: Replacing the Microscope Bulb
Unplug the microscope from power source
Find location of bulb
Follow manufacturer’s instructions to remove the bulb

## Slide 16: Replacing the Bulb (Cont.)
Use tissue paper or an appropriate device to remove the bulb from the microscope
Check the model number on the bulb to ensure the use of correct replacement bulb.
Replace the bulb by holding it with lens paper or an appropriate device.
Never touch the bulb with your fingers.

## Slide 17: Microscope Repair
Never disassemble the microscope
  Optics: eyepieces and objectives
  Mechanics: stage and focus adjustments
Repair of these items requires a service engineer

## Slide 18: Troubleshooting Problems
Many common microscope problems can be prevented or repaired by routine cleaning, adjustment, and maintenance
Seek professional service for problems that cannot be fixed

## Slide 19: Problem: Eye Strain and Headaches
Adjust interpupillary distance
Adjust eyepiece diopter setting
Use matched eyepieces
Notes: <number>
<footer>

## Slide 20: Problem: Poor Image Quality
Clean objective and lens
Clean scope thoroughly
Check oil, replace if any contamination or haze is visible
Ensure the slide is thoroughly dry before applying oil
Notes: <number>
<footer>

## Slide 21: Problem: Uneven Illumination
Ensure nosepiece is rotated to position where it clicks into place
Center condenser
Adjust Kohler illumination
Check to see bulb is correct model
Notes: <number>
<footer>

## Slide 22: Problem: Constant Refocusing
Be sure slide is flat on stage
Clean stage and underside of slide
Be sure microscope is on flat and leveled surface
Notes: <number>
<footer>

## Slide 23: Problem: Surging or Flickering Light
Loose plug or connection
Clean bulb contacts (Unplug from power source)
Assure halogen pins are in full contact
Bulb needs replacement
Check to see if cord is damaged - Replace, if any wear and tear is visible
If voltage supply is erratic, use a voltage stabilizer
Notes: <number>
<footer>
This is an electrical problem.
Check to see the bulb is fully in place and if it is a tungsten bulb that you have cleaned the contacts. Also, many scopes flicker because the operator has leaned on the cord where it enters the base of the scope and damaged the wiring.
If the wiring, rheostat or potentiometer switch are bad, this is a professional level repair.
Another reason for this problem could be power surges.

## Slide 24: Microscope Maintenance-1
Daily Maintenance
  Optimum illumination
  Cleanliness
  Light bulb
  Immersion oil
  Check cord
  Keep it covered
Professional Service
  Minimum of once a year if possible
  Problem that cannot be fixed by laboratory staff

## Slide 25: Microscope Maintenance-2
Maintain a Log Book for
  Problems
  Routine maintenance
  Repairs
  Spares

## Slide 26: Summary Questions
What is another term for “eyepiece”?
What are two reasons for problems with brightness of the viewing field?
Define the purpose of a cabinet box.