## Slide 1
Process Control: Quality Control for Quantitative Tests
<number>

## Slide 2: Learning Objectives
At the end of this module, participants will be able to:
Differentiate accuracy and precision.
Select control material for the laboratory.
Establish acceptable control limits for a method when only one level of control material is available.
Explain the use of a Levey-Jennings chart.
Describe how to correct “out of control” problems.
Notes: <number>

## Slide 3: The Quality Management System
Organization
Personnel
Equipment
Purchasing & Inventory
Process Control
Information Management
Documents & Records
Occurrence Management
Assessment
Process Improvement
Customer Service
Facilities & Safety

## Slide 4: Quantitative Tests
measure the quantity of a particular substance in a sample
quality control for quantitative tests is designed to assure that patient results are:
  accurate
  reliable
Notes: <number>

## Slide 5: Implementation steps
establish policies and procedures
assign responsibility, train staff
select high quality controls
establish control ranges
develop graphs to plot control values - Levey-Jennings charts
monitor control values
develop procedures for corrective action
record all actions taken
Notes: <number>

## Slide 6: What is a Control?
material that contains the substance being analyzed
  include with patient samples when performing a test
used to validate reliability of the test system
  run after calibrating the instrument
  run periodically during testing
Notes: <number>

## Slide 7: Calibrators vs. Controls
Notes: <number>

## Slide 8
Calibrators
A substance with a specific concentration. Calibrators are used to set (calibrate) the measuring points on a scale.
Controls
A substance similar to patients’ samples that has an established
concentration. Controls are used to ensure
the procedure is working
properly.
1
2
3
4
5
1
2
3
4
5
Notes: <number>

## Slide 9: Characteristics of Control Materials
appropriate for the diagnostic sample
values cover medical decision points
similar to test sample (matrix)
available in large quantity; ideally enough for one year
can store in small aliquots
Notes: <number>

## Slide 10: Types of Control Materials
may be frozen, freeze-dried, or chemically preserved
requires very accurate reconstitution if this step is necessary
Notes: <number>

## Slide 11: Sources of Controls Materials
commercially prepared
made “in house”
obtained from another laboratory, usually central or reference laboratory
Notes: <number>

## Slide 12: Control Materials
ASSAYED | Target value predetermined Verify and use
UNASSAYED | Target value not predetermined Full assay required before using
“IN-HOUSE” | In-house pooled sera Full assay, validation
Notes: <number>

## Slide 13: Choosing Control Materials
values cover medical decision points
similar to the test sample
controls are usually available in high, normal, and low ranges
Notes: <number>

## Slide 14: Preparation and Storage of Control Material
adhere to manufacturer’s instructions
keep adequate amount of same lot number
store correctly
Notes: <number>

## Slide 15: Steps in Implementing Quantitative QC
obtain control material
run each control 20 times over 30 days
calculate mean and +/-1,2,3 Standard Deviations
Mean
1SD
1SD
2SD
3SD
2SD
3SD
Notes: <number>

## Slide 16: Measurement of Variability
Variability is a normal occurrence when a control is tested repeatedly
Affected by:
Operator technique
Environmental conditions
Performance characteristics of the measurement
The goal is to differentiate between variability due to chance from that due to error
Notes: <number>

## Slide 17: Measures of Central Tendency
Although variable, sets of data are distributed around a central value
Measurement
F
r
e
q
u
e
n
c
y
Notes: <number>

## Slide 18: Measures of Central Tendency
Mode | the value which occurs with the greatest frequency
Median | the value at the center or midpoint of the observations
Mean | the calculated average of the values
Notes: <number>

## Slide 19: Not all central values are the same
Mode
Median
Mean
Measurement
F
r
e
q
u
e
n
c
y
Notes: <number>

## Slide 20: Symbols Used in Calculations
∑ is the sum of (add data points)
n = number of data points
x1 - xn = all of the measurements (1 through n)
__
X represents the mean
Notes: <number>

## Slide 21: Calculation of Mean
X = Mean
X1 = First measurement
X2 = Second measurement
Xn = Last measurement in series
n = Total number of measurements
Notes: <number>

## Slide 22: Example Calculation of Mean: ELISA Tests
Run controls 20 times in 30 days. Record both OD and cut off (CO) values for each measurement.
Divide the OD by the CO (OD/CO) for each data point or observation. This standardizes the data.
Add the ratios and divide by the number of measurements to get the mean.
Notes: <number>

## Slide 23: Data showing outlier
192 mg/dL
194 mg/dL
196 mg/dL
196 mg/dL
185 mg/dL
196 mg/dL
200 mg/dL
200 mg/dL
202 mg/dL
270 mg/dL
204 mg/dL
208 mg/dL
212 mg/dL
198 mg/dL
204 mg/dL
208 mg/dL
212 mg/dL
198 mg/dL
192 mg/dL
196 mg/dL
Notes: <number>

## Slide 24: Normal distribution
all values symmetrically distributed around the mean
characteristic “bell-shaped” curve
assumed for all quality control statistics
Frequency
mean
Notes: <number>

## Slide 25: What are accuracy and precision?
Quality Control is used to monitor
the accuracy and the precision of the assay.

## Slide 26: Definitions
Accuracy | The closeness of measurements to the true value
Precision | The amount of variation in the measurements
Bias | The difference between the expectation of a test result and an accepted reference value
Notes: <number>

## Slide 27: Accuracy and Precision
Accurate = Precise but not Biased
Accurate and Precise
Precise but Biased
Imprecise
Notes: <number>

## Slide 28: Standard Deviation and Probability
For a set of data with a normal distribution, a random measurement will fall within: + 1 SD 68.3% of the time + 2 SD 95.5% of the time + 3 SD 99.7% of the time
68.2%
95.5%
99.7%
Frequency
-3s - 2s -1s Mean +1s +2s +3s
Notes: <number>

## Slide 29: Standard Deviation (SD)
SD is the principle measure of variability used in the laboratory
Standard Deviation – Statistical Formula
Notes: <number>

## Slide 30: Coefficient of Variation
The coefficient of variation (CV) is the SD expressed as a percentage of the mean.
CV is used to monitor precision
CV is used to compare methods
CV ideally should be less than 5%
Notes: <number>

## Slide 31: Levey-Jennings Chart
Graphically Representing Control Ranges
Notes: <number>

## Slide 32: Statistics for Quantitative QC
  assay control material at least 20 data points over a 20-30 day period
  ensure procedural variation is represented
  calculate mean and + 1, 2 and 3 SD
Notes: <number>

## Slide 33: Draw lines for Mean and SDs (calculated from 20 controls)
MEAN
+1SD
+2SD
-1SD
-2SD
-3SD
+3SD
Days
190.5
192.5
194.5
196.5
188.5
186.5
184.6
Chart name:
Lot number:
Notes: <number>

## Slide 34: Levey-Jennings Chart
MEAN
+1SD
+2SD
-1SD
-2SD
-3SD
+3SD
Days
192.5
194.5
196.5
188.5
186.5
184.6
1
2
3
4
5
6
7
8
9
10
11
12
13
14
15
16
17
18
19
190.5
Plot daily control measurements
Notes: <number>

## Slide 35: Number of Controls
Interpretation depends on number of controls run with patients’ samples.
Good: If one control:
  accept results if control is within ± 2SD unless shift or trend
Better: If 2 levels of controls
  apply Westgard multirule system
Notes: <number>

## Slide 36: Detecting error
random error: variation in QC results with no pattern- only a cause for rejection if outside 2SDs.
systematic error: not acceptable, correct the source of error
  Examples:
  shift–control on one side of the mean 6 consecutive days
  trend–control moving in one direction– heading toward an “out of control” value
Notes: <number>

## Slide 37: Levey-Jennings Chart Shift
MEAN
+1SD
+2SD
-1SD
-2SD
-3SD
+3SD
Days
190.5
192.5
194.5
196.5
188.5
186.5
184.6
Notes: <number>

## Slide 38: Levey-Jennings Chart Trend
MEAN
+1SD
+2SD
-1SD
-2SD
-3SD
+3SD
Days
190.5
192.5
194.5
196.5
188.5
186.5
184.6
Notes: <number>

## Slide 39: Measurement Uncertainty
represents a range of values in which the true value is reasonably expected to lie
is estimated at “95% coverage”
the more precise the method, the smaller the range of values that will fall within 95%
for most instances, a range of + or - 2 SDs is accepted as measurement uncertainty that is explained by random variation
Notes: <number>

## Slide 40: If QC is out of control
STOP testing
identify and correct problem
repeat testing on patient samples and controls after correction
Do not report patient results until problem is solved and controls indicate proper performance
Notes: <number>

## Slide 41: Solving out-of-control problems
identify problem
refer to established policies and procedures for remedial action
Notes: <number>

## Slide 42: Possible Problems
degradation of reagents or kits
control material degradation
operator error
failure to follow manufacturer’s instructions
an outdated procedure manual
equipment failure
calibration error
Notes: <number>

## Slide 43: Summary
A quality control program for quantitative tests is essential. It should:
monitor all quantitative tests
have written policies and procedures, followed by laboratory staff
have a quality manager for monitoring and reviewing QC data
use statistical analysis, provide for good records
provide for troubleshooting and corrective action
Notes: <number>

## Slide 44: Key Messages
A QC program allows the laboratory to differentiate between normal variation and error.
The QC program monitors the accuracy and precision of laboratory assays.
The results of patient testing should never be released if the QC results for the test run do not meet the laboratory target values.

## Slide 45: Questions? Comments?
Organization
Personnel
Equipment
Purchasing & Inventory
Process Control
Information Management
Documents & Records
Occurrence Management
Assessment
Process Improvement
Customer Service
Facilities & Safety