Optional—

Quality Control of Antimicrobial Susceptibility Tests (AST)

| Overview | A number of methods are available to test organisms for their susceptibility to antimicrobial agents. The agar disk diffusion method provides qualitative interpretive category results of susceptible, intermediate, and resistant. The microdilution and agar gradient diffusion methods provide a quantitative result, a minimum inhibitory concentration, rather than a category result. This content sheet covers many of the quality control requirements for the agar disk diffusion method. Regardless of the method, quality control techniques are extremely important to insure that the results are accurate. |  |
| --- | --- | --- |
| Resources | The complexity of testing has become more complicated as new resistance patterns emerge. Laboratories are encouraged to obtain documents such as those from the Clinical Laboratory Standards Institute (CLSI), the French Society of Microbiology (SFM), or the British Society for Antimicrobial Chemotherapy (BSAC), to verify that all of their procedures follow recommended practices. Revisions do occur and laboratories should ensure the latest published version is used. |  |
| Agar disk diffusion method requirements | AST using the agar disk diffusion method is based on the principle that antimicrobial agents of specified concentration in discs will diffuse into media and inhibit the growth of susceptible organisms resulting in a zone of inhibition around the disk. Zones are measured and recorded. But, since each antimicrobial agent diffuses at a different rate, and organisms have unique growth patterns, interpretation of the zones of inhibition must be based on established criteria. Erroneous patient results can occur unless laboratory staff adheres to strict testing requirements. The following can influence the zone size and therefore the interpretation of the results, so each must be carefully controlled: - depth of the medium; - ratio of the medium to the plate size; - pH of the medium; - number of organisms in the inoculation broth; - purity and age of the organism; - storage of antimicrobial disks; - time and temperature of incubation of the inoculated plate; - sufficient light and calipers used to measure zone sizes. |  |
| QC organisms | Because all of the above factors influence the AST process, specific quality control organisms must be run along with patients’ samples. Interpretive criteria are published for specified QC organisms only, so the organisms must be purchased and run with patients’ samples. Control organisms allow the laboratory to verify that their media, processes, and procedures are working properly. The zone sizes of the control organisms for each antibiotic are stable; if they vary from established criteria, the testing system is not working properly and patient results should not be reported. In order to be sure the error is not due to the control organism, the controls must be handled and stored according to the manufacturer’s instructions. Specific storage criteria for disks containing the antimicrobial agents must also be followed. The recommended organisms below should be tested daily in laboratories that test a variety of organisms. The zone sizes for these organisms are published for each antimicrobial agent and are used to verify that test results are correct. - Escherichia coli ATCC 25922 - Staphylococcus aureus ATCC 25923 - Pseudomonas aeruginosa ATCC 27853 Other strains are recommended for fastidious organisms and those that have unique growth requirements. |  |
| McFarland turbidity standards | McFarland 0.5 turbidity standards are available from various manufacturers. Alternately, the 0.5 McFarland may be prepared by adding 0.5 mL of a 1.175% (wt/vol) barium chloride dihydrate (BaCl2•2H2O) solution to 99.5 mL of 1% (vol/vol) sulfuric acid. Note: a latex particle suspension may also be used. The turbidity standard is then aliquoted into test tubes identical to those used to prepare the inoculum suspension. Tightly seal the McFarland standard tubes to prevent evaporation and store in the dark at room temperature (22° to 25°C). The accuracy of the density of a prepared McFarland standard should be checked by using a spectrophotometer with a 1-cm light path; for the 0.5 McFarland standard, the absorbance at a wavelength of 625 nm should be 0.08 to 0.13. Alternately, the accuracy of the McFarland standard may be verified by adjusting a suspension of a control strain (e.g., E. coli ATCC 25922) to the same turbidity, preparing serial 10-fold dilutions, and then performing plate counts. The adjusted suspension should give a count of 108 colony forming units/mL. Check the densities of the McFarland standards monthly; replace as needed. Before each use, shake vigorously to mix the fine white precipitate of barium sulfate in the tube and inspect for a uniform appearance. Discard if large particles are present. Gently invert suspensions made with latex particles. |  |
| Preparation of inoculum | Each culture to be tested should be streaked onto a noninhibitory agar medium (blood agar, brain heart infusion agar, or tryptic soy agar) to obtain isolated colonies. After incubation at 35°C overnight, select 4 or 5 well-isolated colonies with an inoculating needle or loop, and transfer the growth to a tube of sterile saline or nonselective broth (Mueller-Hinton broth, heart infusion broth, or tryptic soy broth) and vortex thoroughly. The bacterial suspension should then be compared to the 0.5 McFarland standard. This comparison can be made more easily if the tubes are viewed against a sheet of white paper on which sharp black lines are drawn. The turbidity standard should be agitated on a vortex mixer immediately prior to use. If the bacterial suspension does not appear to be the same density as the McFarland 0.5, the turbidity can be reduced by adding sterile saline or broth or increased by adding more bacterial growth. Alternatively, the growth method may be used to prepare the inoculum. Four or five colonies are picked from overnight growth on agar and inoculated into broth (Mueller-Hinton broth, heart infusion broth, or tryptic soy broth). The broth is incubated at 35°C until turbid, and then the turbidity is adjusted, if necessary, to the proper density. |  |
| Medium | The recommended medium used for the agar disk diffusion method is Mueller Hinton agar. It must be prepared according to the instructions and poured in the plate at a depth of 3 to 4 mm. Each batch should be checked to confirm the acceptable pH range of 7.2 to 7.4. Before using a new batch of medium, incubate a randomly chosen uninoculated plate for 48-72 hours to check for sterility. |  |
| Antimicrobial disks storage | The working supply of antimicrobial disks should be stored in the refrigerator (4°C). Upon removal of the disks from the refrigerator, the package containing the cartridges should be left unopened at room temperature for approximately 1 hour to allow the temperature to equilibrate. This reduces the amount of condensation on the disks. If a disk-dispensing apparatus is used, it should have a tight fitting cover, be stored in the refrigerator, and be allowed to warm to room temperature before using. Disks containing the concentration of antibiotic that corresponds to the published interpretive criteria must be used, or zone sizes will be interpreted incorrectly. |  |
| Inoculation procedure | Within 15 minutes after adjusting the turbidity of the inoculum suspension, dip a sterile cotton swab into the suspension. Pressing firmly against the inside wall of the tube just above the fluid level, rotate the swab to remove excess liquid. Streak the swab over the entire surface of the medium three times, rotating the plate approximately 60 degrees after each application to ensure an even distribution of the inoculum. Finally, swab all around the edge of the agar surface. Disks may be applied with sterile forceps or with a disk dispenser. Larger laboratories will find it much more efficient to use a disk dispenser. Once applied, disks cannot be moved or their zone sizes may be affected. After the disks are placed on the plate, invert the plate and incubate at 35°C in ambient air for 16 to 18 hours. The time of incubation may need to be extended for certain organisms and antimicrobials. Zones are read from the back of the plate with calipers or rulers and recorded in millimeters. Adequate light is essential to avoid missing small colonies within the zone. |  |
| Antimicrobial gradient method | The Etest® (AB Biodisk, Sweden) combines the principles of disk diffusion and MIC tests. Different concentrations of the antimicrobial agent are imbedded in a paper strip. Organisms are inoculated onto a plate and the strip added. After incubation, the MIC can be determined by reading the strip at the point where the inhibition of the organism is seen in an elliptical shape. |  |
| Points to remember | When reporting results of AST, remember that many organisms have predictable susceptibility patterns. For example, most Staphylococci are susceptible to vancomycin, Streptococcus pyogenes are susceptible to penicillin, and Klebsiella pneumoniae are resistant to ampicillin. If these results are not obtained during testing, the results (identification and susceptibility testing) should be questioned and the test repeated. Monitoring antibiogram data can help generate needed community-specific estimates of resistance. Although all QC may be “in control,” patient results may be incorrect if: - the organism was misidentified; - a clerical error was made; - inappropriate choice of antimicrobials were tested and reported; - the wrong patient’s sample was examined; - the wrong test was ordered; - the sample was not preserved properly. |  |

Clinical Laboratory Standards Institute. Available at URL: http://www.clsi.org/

Société Française de Microbiologie. Available at URL: http://www.sfm.asso.fr/publi/general.php?pa=1

British Society for Antimicrobial Chemotherapy. Available at URL: http://www.bsac.org.uk/