## Slide 1: 10
Assuring quality results

## Slide 2: Disclaimer
WHO Health Security Learning Platform – Training Materials
These WHO Training Materials are © World Health Organization (WHO) 2022. All rights reserved.
Your use of these materials is subject to the “WHO Health Security Learning Platform, Training Materials – Terms of Use”, which you accepted when downloading them and which are available on the Health Security Learning Platform at: https://extranet.who.int/hslp.
Should you adapt, modify, translate, or in any other way revise the contents of these materials, you shall not imply that WHO is in any way affiliated with such modifications and shall not use the WHO name or emblem in such modified materials.
Furthermore, please inform WHO of any modifications of these materials that you use publicly, for record-keeping purposes and continued development, by emailing ihrhrt@who.int.

## Slide 3: Learning objectives
At the end of this module, you will be able to:
define quality assurance (QA) and its main components;
describe the origin of quality control (QC) materials;
define External Quality Assessment (EQA) and its main components;
describe the process of supervisory visits.

## Slide 4: Quality assurance (QA)
QA refers to the processes that are put in place to ensure that a testing site produces quality results.
Quality results are accurate, reliable, relevant and timely.
The main components of QA are:
  process control (quality control [QC])
  assessment (External Quality Assessment [EQA]).

## Slide 5: QA roll-out
Given the need to quickly roll out SARS-CoV-2 Antigen RDTs, a full QA system may not be in place when testing is started.
Consequently, to maximize quality, it is essential that a good-quality RDT product is purchased, and then transported and stored according to the manufacturer’s instructions.
Users must be trained and supervised, with a mechanism for users to report concerns/complaints.
The essential components of QA should be introduced in the following order:
  QC testing
  Supervisory visits
  New lot testing (or new lot verification) of incoming kits
  Proficiency testing

## Slide 6: Quality control (QC) (1)
QC involves the continuous and ongoing evaluation or monitoring of the performance characteristics of the SARS-CoV-2 Antigen RDT to ensure that it is performing to the specifications required for accurate diagnosis.
QC consists of samples with known POSITIVE or NEGATIVE results that are tested with batches of patient samples.
QC results are analysed and compared to the expected results.
If the QC results do not match the expected results, QC has failed. Investigation and corrective actions are required.
Patient test results cannot be released if QC fails.

## Slide 7: Quality control (QC) (2)
QC materials may be supplied with the SARS-CoV-2 Antigen RDT kit.
If QC materials are not supplied with the kit:
  QC materials can be purchased from the SARS-CoV-2 Antigen RDT kit supplier or a third-party provider
  QC materials can be obtained from central or reference laboratories (e.g., standardized control materials or made in-house by pooling specimens from different patients). However, the best options for preparing, characterizing and preserving QC materials for SARS-CoV-2 Antigen RDTs have not yet been determined.
Currently, it is recommended that the QC materials provided by the manufacturer be used to verify test performance prior to initial roll-out of RDTs, intermittently (e.g., for new lot testing and during supervisory visits), and to investigate any complaints/concerns.

## Slide 8: Quality control (QC) (3)
QC must be performed regularly in accordance with the manufacturer’s instructions and national policy.
It is advisable to perform QC at least once a week or more frequently if the testing site has conducted a high number of tests.
Adapt according to the guidelines in your country

## Slide 9: New lot testing
At the time of purchase, SARS-CoV-2 Antigen RDTs should be checked by lot testing, i.e., by evaluating the performance of production lots (batches) before they are deployed in the field.
New lot testing helps to ensure that the SARS-CoV-2 Antigen RDT delivered to the field performs according to the manufacturer's specifications, and if not this should be reported immediately to the manufacturer (Consider using WHO’s user feedback form here )
New lot testing is usually carried out in the laboratory, following standardized procedures and using QC materials (positive and negative controls).
At least five controls (three positive and two negative) should be tested for each new lot.
Adapt according to the guidelines in your country

## Slide 10: Demonstration & proficiency (competency) assessments
QC materials should be used to demonstrate how to perform the SARS-CoV-2 Antigen RDT testing.
After training, all testers are required to perform a competency assessment to certify them as proficient in the testing procedures.
See Module 11: Training users for more information on conducting a competency assessment.
Competency assessments are conducted using QC materials.
If possible, the testers should be blinded to the results from the QC materials for the competency assessments.

## Slide 11: Generic instructions for preparing QC materials for demonstration and proficiency (competency) assessments
QC test materials must be reconstituted and/or tested in accordance with the manufacturer’s instructions:
  Reconstituted controls
  Add the positive and negative control tablets to the extraction buffer tubes (one tube for each control) and close using the cap provided. Mix the controls by hand or vortex. Perform the SARS-CoV-2 Antigen RDT as instructed.
  Swab controls
  Place the positive and negative control swabs in the extraction buffer tubes (one tube for each control) and close using the cap provided. Perform the SARS-CoV-2 Antigen RDT as instructed.
This slide may need to be adapted based on the SARS-CoV-2 Antigen RDT in use

## Slide 12: External Quality Assessment

## Slide 13: External Quality Assessment (EQA)
EQA is a process of assessing the quality of diagnostic testing.
EQA consists of:
  supervisory visits
  proficiency testing
  re-testing

## Slide 14: Supervisory visits (1)
Following SARS-CoV-2 Antigen RDT training and deployment, immediate and sustained follow-up is important to facilitate and support laboratory and healthcare workers in integrating SARS-CoV-2 Antigen RDT testing into routine patient management and record keeping.
Before testing can begin, the readiness of the testing site should be assessed.*
After testing has started, the testing site should be assessed.*
* Testing Site Readiness Checklist and Testing Site Supervisory Checklist templates are provided as part of the SARS-CoV-2 Antigen Rapid Diagnostic Test training package
https://extranet.who.int/hslp/content/sars-cov-2-antigen-rapid-diagnostic-test-training-package
Adapt according to the guidelines in your country

## Slide 15: Supervisory visits (2)
Supervisory visits provide an opportunity to support testing at the testing site and to provide feedback for improving testing.
Supervisory visits should be performed periodically (recommended four times per year).
More frequent visits should be arranged, particularly when a new test (e.g., SARS-CoV-2 Antigen RDT) is introduced, or when frequent, continuous or protracted problems are identified at the testing site.
Adapt according to the guidelines in your country

## Slide 16: Proficiency testing (1)
Proficiency testing (PT) is an important tool for communicating with and motivating testers. Used correctly, a PT programme can be used to identify and resolve problems in diagnostic testing.
The PT process includes the following components:
  An external or internal agency (called a PT provider) regularly provides well-characterized samples to the testing site.
  The person undertaking RDT testing tests the well-characterized samples and reports the results back to the external or internal agency.
  The PT provider analyses the results against the known results, compares the results to other testers from the same site (or from different testing sites) and reports the outcome back to the tester.

## Slide 17: Proficiency testing (2)
External agencies (e.g., Royal College of Pathologists of Australia or United Kingdom NEQAS) provide well-characterized samples for SARS-CoV-2 testing (for NAAT).
Internal agencies (e.g., Influenza EQA programme) may also produce well-characterized samples for testing.
Adapt according to the guidelines in your country

## Slide 18: Proficiency testing (3)
As a tester, conduct the SARS-CoV-2 Antigen RDT PT as directed by the PT programme.
Report the results of the PT tests as required.
If the PT programme reports that the PT results are incorrect, investigate and take corrective action.
Adapt according to the guidelines in your country

## Slide 19: Re-testing
Re-testing is a process by which some tests performed at one testing site are sent to another testing site for re-testing.
Re-testing SARS-CoV-2 Antigen RDTs is not advisable due to the practicality of collecting multiple samples and the safety associated with their transport.
Adapt according to the guidelines in your country

## Slide 20: Taking action
Failed PT tests and identification of areas of improvement during supervisory visits require investigation to determine the reason for the error or deficiency in the quality of testing (often called root cause analysis).
When the root cause has been identified, it must be addressed by taking corrective action.
Preventive actions prevent reoccurrence of the error or deficiency in the quality of testing.
Root cause analysis of issues, as well as the corrective and preventive actions taken must be documented by the testing site.
Adapt according to the guidelines in your country

## Slide 21: Post market surveillance

## Slide 22: Post-market surveillance *
Post-market surveillance - is the process conducted by the manufacturer to collect and analyze experiences with a product on the market.
Examples of issues that may be reported:
  Event-stability of reagent degraded
  Hazards-false negative result, too low result
  Harms delayed diagnosis, delayed treatment, onwards transmission
  Action destruction, shortening shelf-life
* For more information on post market surveillance: Guidance for post-market surveillance and market surveillance of medical devices, including in vitro diagnostics. June 2021.

## Slide 23: Advice to users of SARS-CoV-2 Antigen RDT
All users (lay providers, health care professionals, laboratory personnel, self-testers) should contribute to post-market surveillance by reporting any RDT problems to manufacturers as soon as they notice them
WHO to provide user feedback to?
  Notify the manufacturer (their contact details are in the IFU) or their local authorized representative in your country.
  Notify WHO, if the product has WHO or EUL rapidalert@who.int
WHEN to provide user feedback?
  Immediately (or as soon as possible)
  Document rather than investigate
  You may not be the only user with this problem – you can avert harm to others
For more information on post market surveillance: Guidance for post-market surveillance and market surveillance of medical devices, including in vitro diagnostics. June 2021.

## Slide 24: IVD user feedback form
Product - lot number, expiry date
Event - specimen type, no. tests, no. of patients involved
Health impact - misdiagnosis, delayed diagnosis, delayed treatment, etc.
Consider using WHO’s user feedback form here

## Slide 25
Key points (1)
Quality assurance (QA) aims to guarantee the production of accurate, reliable, relevant and timely test results.
Quality control (QC) involves the monitoring of the performance characteristics of the SARS-CoV-2 Ag-RDT using materials with known POSITIVE or NEGATIVE results.
Patient test results cannot be released if QC fails.
New lot testing should be performed centrally with QC materials before the kits are used.

## Slide 26
Key points (2)
External Quality Assessment (EQA) consists of:
  supervisory visits
  proficiency testing
  re-testing.
PT process involves the testing of well-characterized samples, provided by an external or internal agency to the testing site, reporting of results, and receiving feedback from the provider on the testing site performance with the PT.
TAKE ACTION: Investigate, identify root causes if insufficient quality of testing and implement corrective and preventive actions to resolve.

## Slide 27
Key points (3)
Play your part in post market surveillance - You should report any kit problems to manufacturers immediately.
Provide user feedback to the manufacturer and, if the product has WHO EUL, to WHO.

## Slide 28: Questions?