Biological Safety Cabinet Class II Purchasing Requirements Example

When purchasing a Biological Safety Cabinet Class II the requirements to provide to bidders are outlined as follows:

| Description of requirements | -complying with DIN 12950, part 10 and EN ISO 12429; -internal working area approx. width 1700 x depth 650 x height 650 mm; inside finish, stainless steel; -external housing made of stainless steel or equivalent resistant coating; -vertically adjustable sliding windows (safety glass); -single piece working surface with V – shaped air slits or equivalent system to establish a proper inward flow as well as laminar flow; -luminosity level > 800 lux; -noise level < 58 dBA; -internal fittings: two gas taps with solenoid valve; -two 230± 10V AC main power sockets, European plug; -UVC light 30 W for sterilization with hour counter, to switch on only when white light source is switched off; -electrical control / indicators; -electronic fan control; -flow meter for influx; -flow meter for laminar flow; -filter saturation indicator; -operating hours indicator; -UV timer. |
| --- | --- |
| Filters | -pre-filter construction: easy accessible, filter change without tools preferred; -two HEPA filters (1 main and 1 exhaust), at least Type H14 conform EN 1822; -laminar downflow >= 0.38 m/s; -total volume 1400-1700 m³/h; -ducted air volume 300-450 m³/h. |
| Electricity | -connections: AC 230 ± 10 V, 50Hz, European plug; -power consumption approx. 700 W. (Or if required, connections: AC 115 ± 10 V, 60Hz, American plug, connections: 3-phase current, 360 ± 10 V) |
| Accessories needed | -table or support frame; -working height 80 ± 2 cm, adjustable at least at three feet (points) to level; -all standard accessories/consumables/parts required for the proper operation of the above item shall be included in the offer by the supplier and has to quantified and specified by the supplier; -all standard maintenance tools and cleaning /lubrication materials where applicable shall be included by the supplier and has to quantified and specified by the supplier; -operation manuals and service manuals, at least 1 set each in the purchaser’s language and English; -manufacturing certificate; -spare parts list; -standard accessories; -two additional sets of HEPA filters as specified above; -a hood (air duct) built as an exhaust air valve set up for the biosafety cabinet with motorized flap and switch for external ventilation. |
| Notes | Bidder may propose any product/system which is equivalent or better than the requirements specified above. Bidder shall furnish a clause-by-clause commentary on the requirements above. Bidder who commented with texts such as "Yes", "Refer to catalogue" or any text of this nature shall be considered as NON-RESPONSIVE. |