Activity 7-1: Calculation of Mean and Standard Deviation

| Purpose | To practice calculating the mean and standard deviation (SD) of a set of data to use for establishing control ranges. |
| --- | --- |
| Suggested Time | 15 minutes |
| Instructional Method | Individual assignment |
| Instructions | Ask participants: Calculate the mean and SD using Annex 7-A and the two Standard Deviation Worksheets provided. |

| Standard Deviation Worksheet |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Method: Glucose |  |  | Sample: Pooled Serum, Level 1 Control |  |  |  |
| Day | Observation (x) |  | Difference (x-mean) |  |  | Difference (x-mean) Squared |
| 1 | 98 |  |  |  |  |  |
| 2 | 95 |  |  |  |  |  |
| 3 | 103 |  |  |  |  |  |
| 4 | 102 |  |  |  |  |  |
| 5 | 100 |  |  |  |  |  |
| 6 | 105 |  |  |  |  |  |
| 7 | 103 |  |  |  |  |  |
| 8 | 95 |  |  |  |  |  |
| 9 | 92 |  |  |  |  |  |
| 10 | 93 |  |  |  |  |  |
| 11 | 108 |  |  |  |  |  |
| 12 | 98 |  |  |  |  |  |
| 13 | 98 |  |  |  |  |  |
| 14 | 101 |  |  |  |  |  |
| 15 | 103 |  |  |  |  |  |
| 16 | 104 |  |  |  |  |  |
| 17 | 100 |  |  |  |  |  |
| 18 | 100 |  |  |  |  |  |
| 19 | 92 |  |  |  |  |  |
| 20 | 91 |  |  |  |  |  |
| Sum: |  |  |  |  |  |  |
| Mean: |  |  |  |  |  |  |
| Variance: |  |  |  |  |  |  |
| SD: |  |  |  |  |  |  |
|  | Lower Limit |  | Upper Limit |  |  |  |
| 1 SD |  |  |  |  |  |  |
| 2 SD |  |  |  |  |  |  |
| 3 SD |  |  |  |  |  |  |
| Standard Deviation Worksheet |  |  |  |  |  |  |
| Method: Glucose |  |  |  | Sample: Pooled Serum, Level 2 Control |  |  |
| Day |  | Observation (x) |  | Difference (x-mean) | Difference (x-mean) Squared |  |
| 1 |  | 293 |  |  |  |  |
| 2 |  | 279 |  |  |  |  |
| 3 |  | 277 |  |  |  |  |
| 4 |  | 282 |  |  |  |  |
| 5 |  | 279 |  |  |  |  |
| 6 |  | 279 |  |  |  |  |
| 7 |  | 281 |  |  |  |  |
| 8 |  | 286 |  |  |  |  |
| 9 |  | 276 |  |  |  |  |
| 10 |  | 282 |  |  |  |  |
| 11 |  | 265 |  |  |  |  |
| 12 |  | 267 |  |  |  |  |
| 13 |  | 273 |  |  |  |  |
| 14 |  | 289 |  |  |  |  |
| 15 |  | 270 |  |  |  |  |
| 16 |  | 276 |  |  |  |  |
| 17 |  | 269 |  |  |  |  |
| 18 |  | 280 |  |  |  |  |
| 19 |  | 285 |  |  |  |  |
| 20 |  | 279 |  |  |  |  |
| Sum: |  |  |  |  |  |  |
| Mean: |  |  |  |  |  |  |
| Variance: |  |  |  |  |  |  |
| SD: |  |  |  |  |  |  |
|  |  | Lower Limit |  | Upper Limit |  |  |
| 1 SD |  |  |  |  |  |  |
| 2 SD |  |  |  |  |  |  |
| 3 SD |  |  |  |  |  |  |

Answer sheets for Activity 7-1 follow on pages 3 (Level 1 Control) and 4 (Level 2 Control).

| Activity 7-1 Answers to Level 1 Control Mean and SD Activity |  |  |  |
| --- | --- | --- | --- |
| Standard Deviation Worksheet |  |  |  |
| Method: Glucose |  | Sample: Pooled Serum, Level 1 |  |
| Day | Observation (x) | Difference (x-mean) | Difference (x-mean) Squared |
| 1 | 98 | -1.050 | 1.102 |
| 2 | 95 | -4.050 | 16.403 |
| 3 | 103 | 3.950 | 15.603 |
| 4 | 102 | 2.950 | 8.703 |
| 5 | 100 | 0.950 | 0.903 |
| 6 | 105 | 5.950 | 35.403 |
| 7 | 103 | 3.950 | 15.603 |
| 8 | 95 | -4.050 | 16.403 |
| 9 | 92 | -7.050 | 49.703 |
| 10 | 93 | -6.050 | 36.603 |
| 11 | 108 | 8.950 | 80.103 |
| 12 | 98 | -1.050 | 1.102 |
| 13 | 98 | -1.050 | 1.102 |
| 14 | 101 | 1.950 | 3.803 |
| 15 | 103 | 3.950 | 15.603 |
| 16 | 104 | 4.950 | 24.503 |
| 17 | 100 | 0.950 | 0.903 |
| 18 | 100 | 0.950 | 0.903 |
| 19 | 92 | -7.050 | 49.703 |
| 20 | 91 | -8.050 | 64.803 |
| Sum: | 1981 |  | 438.950 |
| Mean: | 99.05 |  | 23.103 (Sum/19) |
| Variance: | 23.103 |  |  |
| Standard Deviation: | 4.807 |  |  |
|  | Lower Limit | Upper Limit |  |
| 1 SD | 94.243 | 103.857 |  |
| 2 SD | 89.437 | 108.663 |  |
| 3 SD | 84.630 | 113.470 |  |

| Activity 7-1 Answers to Level 2 Control Mean and SD Results |  |  |  |
| --- | --- | --- | --- |
| Standard Deviation Worksheet |  |  |  |
| Method: Glucose |  | Sample: Pooled Serum, Level 2 |  |
| Day | Observation (x) | Difference (x-mean) | Difference (x-mean) Squared |
| 1 | 293 | 14.65 | 214.6225 |
| 2 | 279 | 0.65 | 0.4225 |
| 3 | 277 | -1.35 | 1.8225 |
| 4 | 282 | 3.65 | 13.3225 |
| 5 | 279 | 0.65 | 0.4225 |
| 6 | 279 | 0.65 | 0.4225 |
| 7 | 281 | 2.65 | 7.0225 |
| 8 | 286 | 7.65 | 58.5225 |
| 9 | 276 | -2.35 | 5.5225 |
| 10 | 282 | 3.65 | 13.3225 |
| 11 | 265 | -13.35 | 178.2225 |
| 12 | 267 | -11.35 | 128.8225 |
| 13 | 273 | -5.35 | 28.6225 |
| 14 | 289 | 10.65 | 113.4225 |
| 15 | 270 | -8.35 | 69.7225 |
| 16 | 276 | -2.35 | 5.5225 |
| 17 | 269 | -9.35 | 87.4225 |
| 18 | 280 | 1.65 | 2.7225 |
| 19 | 285 | 6.65 | 44.2225 |
| 20 | 279 | 0.65 | 0.4225 |
| Sum: | 5567 |  | 974.55 |
| Mean: | 278.350 |  | 51.292 (Sum/19) |
| Variance: | 51.292 |  | 7.162 |
| Standard Deviation: | 7.162 |  |  |
|  | Lower Limit | Upper Limit |  |
| 1 SD | 271.188 | 285.512 |  |
| 2 SD | 264.026 | 292.674 |  |
| 3 SD | 256.864 | 299.836 |  |

Activity 7-2: Creation of Levey-Jennings Charts

| Purpose | To create Levey-Jennings Charts to monitor examination results by visualizing daily control data. |
| --- | --- |
| Suggested Time | 15 minutes |
| Instructional Method | Individual assignment |
| Instructions | Give graph paper to participants. Use the mean and SDs calculated in Activity 7-1 for Level 1 and Level 2 Controls to make two Levey-Jennings Charts. - Draw a line in the middle of one section of the graph paper to represent the mean of the Level 1 control. Write the value of the mean on the left of the line. - Mark off plus (+) and minus (-) 3 SDs at appropriate intervals on the graph, then draw a horizontal line through each, and write the values beside each line. - Label the x-axis with days, runs, or other interval used to run QC. Divide the x-axis into equal spaces and label with time intervals, for example, days. - Label the chart with the name of the test and the control’s lot number. - Repeat for the Level 2 control. - Plot the data provided in Table 1 on the respective Levey-Jennings charts. Table 1. Activity 7-2. Data to plot on Levey-Jennings charts.. Day; Level 1 Observation (x); Level 2 Observation (x). 1; 98; 280. 2; 101; 283. 3; 110; 285. 4; 105; 282. 5; 114; 285. 6; 97; 277. 7; 98; 278. 8; 101; 262. 9; 102; 264. 10; 96; 277. 11; 99; 281. 12; 110; 294. 13; 98; 279. 14; 95; 283. 15; 86; 295. 16; 96; 276. 17; 91; 277. 18; 90; 283. |