Content

- Scope

- Definitions and abbreviations

- Personnel qualifications

3.1 Medical fitness

3.2 Education and training

- Procedure

4.1 Principle

4.2 Samples

4.3 Equipment and materials

4.4 Reagents and solutions

4.5 Procedure for defrosting and cleaning

4.6 Reading and reporting

4.7 Maintenance

4.8 Waste management

- Related documents

Annex 1. Temperature record form

Annex 2. Maintenance logbook

|  | Compiled by | Examined by | Approved by | Replaced | New version |
| --- | --- | --- | --- | --- | --- |
| Name |  |  |  | Code: | Code: |
| Date |  |  |  |  |  |
| Signature |  |  |  |  |  |
| Laboratory area: |  | No of copies: | Reason for change: |  |  |

- Scope

This SOP describes the optimal operation of the freezer (or deep freezer) – achieved through proper installation and maintenance (regular defrosting and cleaning) – and relates to equipment in an infectious area of the laboratory.

- Definitions and abbreviations

°C:	degrees Celsius

NA:	not applicable

UPS:	uninterrupted power supply.

- Personnel qualifications

- Medical fitness

In accordance with national laws and practices, arrangements should be made for  appropriate health surveillance of TB laboratory workers:

- before enrolment in the TB laboratory;

- at regular intervals thereafter, annually or bi-annually;

- after any biohazard incident;

- at the onset of TB symptoms.

Ideally, individual medical records shall be kept for up to 10 years following the end of occupational exposure.

Laboratory workers should be educated about the symptoms of TB and provided with ready access to free medical care if symptoms arise.

Confidential HIV counselling and testing should be offered to laboratory workers. Options for reassignment of HIV-positive or immuno-suppressed individuals away from the high-risk areas of the TB laboratory should be considered.

All cases of disease or death identified in accordance with national laws and/or practice as resulting from occupational exposure to biological agents shall be notified to the competent authority.

- Education and training

Education and training must be given on the following topics:

- potential risks to health (symptoms of TB disease and transmission);

- hygiene requirements;

- wearing and use of protective equipment and clothing;

- handling of potentially infectious materials;

- prevention of incidents and steps to be taken by workers in the case of incidents (biohazard incidents, chemical, electrical and fire hazards);

- good laboratory practice;

- organization of work flow;

- waste handling;

- use of equipment (operation, identification of malfunctions, maintenance).

The training shall be:

- given before a staff member takes his/her post;

- strictly supervised;

- adapted to take account of new or changed conditions; and

- repeated periodically, preferably every year.

- Procedure

- Principle

The freezer is intended to ensure a suitable environment for material that requires preservation at –18 °C.

- Samples

Typically, a freezer is used to preserve:

- mycobacterial cultures and specimens stored for specific purposes;

- specific reagents and solutions used for drug-susceptibility testing.

Never store flammable solutions in a freezer that is not approved and certified for this purpose.

Use separate freezers for clean and infectious materials.

- Equipment and materials

The freezer should be installed and operated in accordance with the specific manufacturer’s manual. Specifically:

- After transportation, leave the freezer in the upright position for some hours before connecting it to the electricity supply.

- Do not install the freezer close to heat sources.

- Locate the freezer in a dry, well-ventilated area.

- Ensure that the freezer is placed on a level surface.

- To prevent loss of valuable cultures, chemicals, etc. in case of a power failure, the freezer should ideally be connected to a suitable UPS system.

- Reagents and solutions

NA

- Procedure

4.5.1 Defrosting

Identify an adequate volume of available space in another freezer to store materials during the defrosting procedure. Ensure that materials kept in the freezer are clearly marked.

Switch the freezer off and disconnect it from the electrical supply. Open the freezer door and leave it open.

Position a container to catch the melted ice.

Never use sharp tools to chip off the ice.

Sponge up any melted ice.

4.5.2 Cleaning

Clean the interior of the freezer with a disinfectant solution (see SOP on disinfectants).

Clean the outside of the freezer with a cleaning solution, and dry with a soft cloth.

Reconnect the freezer to the main power supply and switch it on.

Replace the original freezer contents once the temperature has reached –18 °C. 	Do not overcrowd.

Notify staff of the replacement of the freezer contents.

Note the replacement of freezer contents in the freezer logbook.

- Reading and reporting

Use a record form to keep a daily record of the freezer temperature (Annex 1).

Record the regular maintenance of the freezer on a maintenance card (Annex 2).

- Maintenance

4.7.1  Daily

Check the compressor for any unusual sound and for overheating.

4.7.2  Monthly

Clean the filters and screens of the ventilator system with a brush or vacuum cleaner.

4.7.3  Every six months

Defrost and clean the freezer as described in section 4.5. This can be done more frequently if necessary, particularly in the event of leakage of biological materials onto the internal surface of the freezer or when the ice build-up on the internal wall reaches a thickness of 5–6 mm.

Clean the condenser coils and fan with a brush or vacuum cleaner.

Repairs should be performed only by a qualified service technician.

- Waste management

Freezer contents may be disposed of only with the prior consent of the laboratory manager.

Details of disposal must be noted in the freezer logbook.

Infectious cultures or contaminated materials must be autoclaved before disposal.

- Related documents

Manufacturer’s manual, specific to each freezer.

Maintenance and repair of laboratory, diagnostic imaging and hospital equipment. Geneva, World Health Organization, 1994.

Maintenance manual for laboratory equipment, 2nd ed. Geneva, World Health Organization, 2008 (available at www.who.int/entity/diagnostics_laboratory/documents/guidance/guidance2/en/)

Annex 1. Temperature record form

| Equipment: FREEZER | Reference: |
| --- | --- |
| Location: | Installation date: |
| Temperature required: –18 °C ± 2 °C (acceptable variation) |  |

| Trimester: |  |  |  | #### Year: |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| ## Month | Temp. °C | Operator (initials) | ## Month | Temp. °C | Operator (initials) | ## Month | Temp. °C | Operator (initials) |
| 1 |  |  | 1 |  |  | 1 |  |  |
| 2 |  |  | 2 |  |  | 2 |  |  |
| 3 |  |  | 3 |  |  | 3 |  |  |
| 4 |  |  | 4 |  |  | 4 |  |  |
| 5 |  |  | 5 |  |  | 5 |  |  |
| 6 |  |  | 6 |  |  | 6 |  |  |
| 7 |  |  | 7 |  |  | 7 |  |  |
| 8 |  |  | 8 |  |  | 8 |  |  |
| 9 |  |  | 9 |  |  | 9 |  |  |
| 10 |  |  | 10 |  |  | 10 |  |  |
| 11 |  |  | 11 |  |  | 11 |  |  |
| 12 |  |  | 12 |  |  | 12 |  |  |
| 13 |  |  | 13 |  |  | 13 |  |  |
| 14 |  |  | 14 |  |  | 14 |  |  |
| 15 |  |  | 15 |  |  | 15 |  |  |
| 16 |  |  | 16 |  |  | 16 |  |  |
| 17 |  |  | 17 |  |  | 17 |  |  |
| 18 |  |  | 18 |  |  | 18 |  |  |
| 19 |  |  | 19 |  |  | 19 |  |  |
| 20 |  |  | 20 |  |  | 20 |  |  |
| 21 |  |  | 21 |  |  | 21 |  |  |
| 22 |  |  | 22 |  |  | 22 |  |  |
| 23 |  |  | 23 |  |  | 23 |  |  |
| 24 |  |  | 24 |  |  | 24 |  |  |
| 25 |  |  | 25 |  |  | 25 |  |  |
| 26 |  |  | 26 |  |  | 26 |  |  |
| 27 |  |  | 27 |  |  | 27 |  |  |
| 28 |  |  | 28 |  |  | 28 |  |  |
| 29 |  |  | 29 |  |  | 29 |  |  |
| 30 |  |  | 30 |  |  | 30 |  |  |
| 31 |  |  | 31 |  |  | 31 |  |  |

Annex 2. Maintenance logbook

| ## ITEM IDENTIFICATION |  |  |
| --- | --- | --- |
| Equipment: FREEZER | Brand name: |  |
| Purchase date: | Model/type: |  |
| Location within laboratory: | Serial no. |  |
| Warranty expiry date: |  |  |
| Manufacturer: |  | Tel: |
| Address:. |  |  |
| Contact person: |  |  |
| Technical service representative: |  | Tel: |

Maintenance card

| PERIODICITY: Every six months or when needed |  |  |
| --- | --- | --- |
| Date | Maintenance operation | Operator |
|  | Defrosting: Relocation: What/Where: What/Where: What/Where: Refilling: |  |
| Remarks: |  |  |
| Date | Maintenance operation | Operator |
|  | Defrosting: Relocation: What/Where: What/Where: What/Where: Refilling: |  |
| Remarks: |  |  |
| Date | Maintenance operation | Operator |
|  | Defrosting: Relocation: What/Where: What/Where: What/Where: Refilling: |  |
| Remarks: |  |  |

| FAILURE EVENTS |  |  |  |
| --- | --- | --- | --- |
| Date | Event | Corrective action taken | Operator |