## Sheet1
Additional EMS Testing Requirements for DL01 devices
For a new datalogger, a full application must be submitted with a test report to E006/DL01-VP.3 from a PQS approved lab.
Note: these are tests that need to be repeated if the appliance differs from the appliance with which the datalogger was approved.
All appliances should verify specification E006/DL01 Clause 4.2.6, to confirm the backup time of the datalogger complies with the holdover time of the appliance.
"-" denotes no further testing required.
Appliance difference | Test 2 | Test 3 | Test 4 | Test 5 | Test 6 | Test 7 | Test 8 | Comments
Appliance is in a new category, e.g. RF05, RF03, TS01, compared to previously prequalified appliances. | 3rd party lab | - | - | - | 3rd party lab | In-house | - | Verification the logger still meets the requirements of E006/DL01.3 Clause 4.2.6. Additionally submit all data files in their raw format from Test 7.
New appliance has different electrical architecture or components than previously prequalified appliances. For example, different compressor, controller, or AC/DC power supply. | 3rd party lab | - | - | - | 3rd party lab | In-house | - | Verification the logger still meets the requirements of E006/DL01.3 Clause 4.2.6. Additionally submit all data files in their raw format from Test 7.
New appliance has identical electrical architecture and components as a prequalified appliance, except that door sensors are different models or different geometry is used to trigger them. | - | - | - | In-house | - | - | - | Submit all data files in their raw format from Test 5.
New appliance uses a different temperature sensor (e.g. vaccine compartment, ambient, etc.) for the datalogger. | - | - | - | - | - | In-house | - | Additional submission information per E006/DL01.3 Clause 4.2.15, 4.2.16, and 4.2.17 are required. Additionally submit all data files in their raw format from Test 7.
New appliance has identical electrical architecture and components as a prequalified appliance. | [7 empty] | Verification the logger still meets the requirements of E006/DL01.3 Clause 4.2.6. | bar: "-" denotes no further testing required. Test 2 to "-" denotes no further testing required. Test 8
Test 2: Logger power wakeup and recharging
Test 3: Logger data access
Test 4:Endurance of timekeeping and recording functions, including timestamp drift
Test 5: Door opening
Test 6: M2M load safety
Test 7: Logger data review
Test 8: Disruption to timekeeping
Additional EMS Testing Requirements for EM01 devices
For a new Level 2 or 3 EMD, a full application must be submitted with a test report to E006/EM01-VP.2 from a PQS approved lab.
For E-EMDs no additional testing is required.
For I-EMDs E006/EM01.2 Clause 4.5.2 and 4.5.4 should be verified.