## Review
E003 Vaccine freezer or combined vaccine and water-pack freezer submission template |  |  | Assessment of appliance: | [Brand and model here]
| Tested in accordance with PQS verification protocol: | E003/FZ01-VP.2 (06 JUL 2010) | Submitted by: | [Applicant's company here]
| To verify compliance with PQS specification: | E003/FZ01.2 (08 SEP 2020)
Instructions in [square brackets] are for your guidance. Please delete these instructions and replace with evidence e.g. a precise reference (including page and section numbers) showing compliance with that particular PQS specification requirement. In general a "pass" or "complies" statement alone is not acceptable - evidence for compliance must be provided.
One document may provide evidence for several clause requirements. Please carefully reference the sub-sections of any such document so that it is clear that the clause requirement has been met. This will help speed the pre-qualification process.
Greyed-out cells are for information and do not require a response.
NOTE: If it is not clear or doubtful whether a requirement has been met, then this is a non-compliance and pre-qualification might be delayed.
|  |  | REVIEWER’S SUMMARY: |  | REVIEWER’S SUMMARY: |  | REVIEWER’S SUMMARY:
Section 7: | Product dossier | Applicant's submission | Reviewer's remarks [Name & Date] | Applicants response [Name & Date] | Reviewer's remarks [Name & Date] | Applicants response [Name & Date] | Reviewer's remarks [Name & Date]
The legal manufacturer or reseller must provide WHO with a prequalification dossier containing the following:
| Dossier examination fee in US dollars. | [Please contact PQS.]
| General information about the legal manufacturer, including name and address. | [Please state or refer to a document.]
| Unique identification reference for the product type. | [Please state.]
| Brand name of the product. | [Please state.]
| Full specifications of the product being offered, covering all the requirements set out in this document, including details of product marking and traceability. | [Please list the documents submitted in the dossier. (See adjacent excel spreadsheet.)]
| A comprehensive set of photographs showing all external surfaces of the unit, the interior layout, the compressor and a close-up of the thermometer and the control panel. | [Please refer to a document.]
| Certified photocopies of all type-approvals obtained for the product, including CE marking and the like. | [Please refer to a document.]
| Certified photocopies of the legal manufacturer’s ISO 9001 quality system certification. | [Please refer to a document & section.]
| Where relevant, certified photocopies of the legal manufacturer’s ISO14001 certification, EMAS registration or registration with an equivalent environmental audit scheme. Conformity with an environmental audit scheme is not mandatory however, preference will be given to manufacturers who are able to demonstrate compliance with good environmental practice. | [Please refer to a document.]
| Laboratory test report(s) proving conformity with PQS specifications. | [Please refer to a document.]
| Indicative cost of the product per unit, per 10 units and per 100 units, EXW (Incoterms 2010). | [Please state or refer to a document.]
Section 4. | Requirements
4.1 | General
| Compression-cycle vaccine freezers or combined vaccine and water-pack freezers, powered by mains electricity, are used primarily in areas with a reliable electricity supply (i.e. 20 or more hours of continuous electricity per typical day). Manufacturers may offer products suitable for one or more temperature zones. | [Please state the minimum hours of electricity per day required by the appliance.]
4.2 | Performance
4.2.1 | Operating temperature range
| As indicated on the temperature zone rating sticker attached to the product (see Annex 1 of E003/FZ01.2). | [Please state the temperature zone and show a photo of this label on the appliance.]
4.2.2 | Refrigeration system
| Compression-cycle unit operating on alternating current electricity. | [Please state.]
4.2.3 | Voltage and frequency
| 220-240-volt 50/60 Hz and 100-127-volt 50/60 Hz options are to be offered. Performance is to be identical for all options, regardless of the nominal voltage and frequency rating of the appliance. | [Please state which voltage options are available.]
4.2.4 | Water-pack freezing
| In combined freezers, not less than 7.2 kg of water-packs must be frozen per 24 hours while maintaining the temperature control specified in clause 4.2.6. | [Please state the water-pack freezing capacity and the precise test reference that proves this.]
4.2.5 | Areas not suitable for vaccine storage
| Areas of an otherwise acceptable appliance which are too warm must be excluded from use by design. | [Please state how the vaccine storage is physically prohibited from areas warmer than –15°C.]
4.2.6 | Temperature control
| All freezers: The vaccine load must remain below -15°C during any continuous ambient temperature test(s) or day/night cycling temperature test(s). Combined units must achieve this performance with no water-packs in the water-pack compartment. | [Compliance to this clause must be demonstrated in the laboratory testing. Please make precise reference to these tests.]
| Combined freezers only: While freezing a quantity of water-packs equal to its water-pack freezing capacity, the temperature of the full load of vaccines must remain below -5°C and return to below -15°C within the 24- hour freezing cycle under the maximum continuous ambient temperature test conditions of its rated temperature zone. | [Compliance to this clause must be demonstrated in the laboratory testing. Please make precise reference to these tests.]
4.2.7 | Thermometer
| Option A: Externally readable cabinet-mounted gas or vapour pressure dial thermometer complying with PQS specification E006/TH02. | [Please state which thermometer is built into your appliance and refer to a photo.]
| Option B: Externally readable cabinet-mounted electronic thermometer conforming to PQS specification E006/TH06. | [Please state which thermometer is built into your appliance and refer to a photo.]
4.2.8 | Holdover time
| No standard set; however, performance data will be published. | [Please state the holdover time and the laboratory test reference that confirms this.]
4.2.9 | Compressor starting voltage
| At 22% below manufacturers stated voltage, 10 out of 10 cold starts and 10 out of 10 hot starts must all be successful. | [Please state the precise laboratory test reference that confirms this.]
4.2.10 | Power consumption
| No standard set however results will be reported. | [Please state power or energy consumption as confirmed by various laboratory tests.]
4.2.11 | Evaporator configuration
| If the evaporator is mounted in shelves the minimum clearance between shelves must be 130 mm. | [Please state.]
4.2.12 | Lock
| The door or lid must be fitted with a lock. Two keys are to be supplied with every unit. | [Please provide a photo of the lock.]
4.2.13 | Corrosion
| Internal and external cabinet, lid and frame protected against corrosion to DIN 8985. | [Please state the level of corrosion protection for all internal and external components and refer to the testing that proves this. (NOTE: The level of corrosion protection must be as good or better than DIN 8985.)]
4.2.14 | Electrical safety rating
| Manufacturer to certify compliance with IEC 60335-1 and IEC60335-2-24. | [Please provide test evidence for electrical safety compliance.]
4.2.15 | Markings
| Compressors must be marked with the blue identifying symbol shown in Annex 2. In addition, the freezer cabinet must be permanently marked, near the compressor position, with the chemical name of the refrigerant, or with the refrigerant number, formula or proportion (for blended refrigerants). All appliances require an asset identification label (bar code) as specified in WHO/PQS/E003/BC01.1: Global asset identification. Effective date June 30, 2020. | [Please refer to a photo of the blue label on the compressor. Please refer to a photo of the bar code.]
4.2.16 | Vaccine storage advice
| All units must carry a factory-fitted non-removable label, designed to last the lifetime of the appliance, carrying the following information: • Vaccine freezers: Vaccine storage instructions and the appropriate temperature zone symbol as Annex 1. • Combined freezers: Vaccine storage instructions, water-pack freezing instructions and the appropriate temperature zone symbol as Annex 1. The instructions should be fixed to the lid of chest freezers and near the top of the door on upright freezers. Instructions should be in one of the languages specified in clause 4.11, as indicated by the purchaser at the time of ordering. | [Please provide a clear copy (not a photo) of these instructions in accordance with this clause and provide a photo of the position of the instructions on the appliance.]
4.3 | Environmental requirements
4.3.1 | Ambient temperature range during transport and storage
| Ambient temperature range during transport and storage is -30°C to +55°C when the product is inactivated. | [Please provide evidence of compliance.]
4.3.2 | Ambient humidity range during transport, storage and use
| 5% to 95% RH, non-condensing. | [Please provide evidence of compliance.]
4.4 | Physical characteristics
4.4.1 | Overall dimensions
| To allow for manoeuvring around corners and through corridors and doorways, the minimum dimension of the product (either length, width or height) should not exceed 710 mm; exceptionally a minimum dimension up to 830 mm can be accepted, but this may restrict the number of sites where the appliance can be installed. The maximum dimension should not exceed 1700 mm and the maximum diagonal (corner to corner) dimension must not exceed 1850 mm. All dimensions must be clearly stated on the PQS information sheet. | [Please refer to a precise section in the test report which details the dimensions.]
4.4.2 | Weight
| Mechanical lifting equipment will often not be available at the installation sites. It is recommended that the refrigerator and any associated components should be designed for lifting in such a way that no single worker is required to carry more than 25 kg while working on their own, or in a group. Mechanical lifting equipment at the installation sites is necessary when component weight exceed practical weight limits of the workforce. | [Please state the weight of the appliance and explain how the appliance may be easily lifted or manoeuvred into position on site. e.g. does the appliance have handles or are "carry straps" provided?]
4.5 | Interface requirements
4.5.1 | Voltage stabilizer compatibility
| All electrical components must be compatible with voltage stabilizers that use tap-changing technology. If the product contains components that are incompatible with this type of voltage stabilizer it must be supplied with an appropriate device of equivalent performance to those voltage stabilizers that are currently prequalified in PQS section E007. A warning must be affixed to the unit stating the type(s) of voltage stabilizer that may be used and the user manual and spare parts list must clearly record this warning. | [Please state which PQS approved voltage stabilizer (VS) will be supplied with the appliance and refer to the test report proving PQS compliance for this VS. Also show the warning on the appliance as to which VS must be used. (NOTE: A power adaptor must be qualified as an approved PQS voltage stabilizer or the power adaptor must be plugged into a VS.)]
4.5.2 | Power lead
| The appliance must be supplied with a power lead with a sealed-on plug compatible with the electrical socket standard in the country where the equipment must be installed. The power lead must be at least 1.5 metre and not more than 2.0 metre in length. | [Please state the measured length of the power cord.]
4.6 | Human factors
4.6.1 | Generally
| The product must be usable by the widest practicable range of active health workers, regardless of age, gender, size or minor disability, including colour blind users and long-sighted people without glasses, in accordance with the general principles laid out in ISO 20282-1:2006. | [Please explain how your appliance complies with the general principles in ISO 20282-1:2006.]
4.6.2 | Control panel and thermometer
| Controls, thermometer and other visual displays may be positioned on the front of the unit preferably as close to eye level as possible. Alternatively, they may be mounted on top of the unit at a height not exceeding 1.3 metres. If a low-level position is essential, the display must be aligned so that it can easily be read without the user having to squat or kneel down. The on-off and/or defrost switch, if present, should be recessed or otherwise protected so that it is not possible inadvertently to activate it. | [Please provide a close-up photo of the display with annotated information of the LEDs or other indicators and a separate photo of the position of the display on your appliance.]
4.7 | Materials and construction
4.7.1 | Refrigerant
| HFC (hydro fluorocarbon) or HC (hydrocarbon) refrigerant. CFC (chlorofluorocarbon) and HCFC (hydrochlorofluorocarbon) gases are not acceptable. The suitability of alternative refrigerant gases will continue to be assessed and preference will be given to products that use gases with low global warming potential (GWP). | [Please state type of refrigerant(s) and their GWP.]
4.7.2 | Thermal insulation foaming agents
| Any gas complying with the limitations and deadlines set by the Montreal Protocol on the elimination of ozone-depleting chemicals. | [Please state type of foaming agent, it’s compliance with the Montreal Protocol and its GWP.]
4.7.3 | Other restricted materials
| The product and its constituent components, including batteries, must not contain lead, mercury, cadmium, hexavalent chromium, polybrominated biphenyls (PBB) or polybrominated biphenyl ethers (PBDE). | [Please provide a RoHS report or equivalent.]
4.7.4 | PCM
| Integrated thermal buffer materials may be used to prevent freezing temperatures from propagating to the vaccine storage compartment or for other thermal purposes. The buffer material may be PCM-based but if so, must comply with WHO/PQS/ E005/PCMC0.1 – PCM specification for phase-change material containers. | [Please state any thermal buffer material and provide evidence for compliance with WHO/PQS/ E005/PCMC0.1.]
4.8 | Warranty
| The appliance to be covered by a three-year replacement warranty from receipt of appliance in the event of any component failure arising from defective design, materials or workmanship. | [Please refer to a statement of the warranty with terms and conditions. (This is often provided in the user manual.)]
4.9 | Servicing provision
| The product is to be designed to achieve a maintenance-free life of not less than 10 years apart from routine de-frosting and cleaning and replacement of batteries (if any). | [Please explain how the appliance is designed to last at least 10 years. Please state if the appliance has any batteries. (It may be helpful to refer to photographic examples of similar products which already have been in the field for at least 10 years.]
4.10 | Disposal and recycling
| The manufacturer must provide information to the buyer on the hazardous materials contained within the system and suggestions for resource recovery/recycling and/or environmentally safe disposal. For the European Union WEEE compliance in accordance with European Union Directive 2002/96/EC is mandatory. | [Please list hazardous material(s) in the user manual with the provision for disposal/recycling. (NOTE: This should comply with local regulations for disposal.)]
4.11 | Instructions
User and maintenance instructions in Arabic, English, French, Mandarin Chinese, Russian and Spanish. The instructions are to be written for users and repair technicians and are to cover the following topics:
| • installation procedures | [Please refer to a precise section in the user manual.]
| • compatible types of voltage stabilizer | [Please refer to a precise section in the user manual.]
| • temperature adjustments (if applicable) | [Please refer to a precise section in the user manual.]
| • simple daily, weekly and monthly maintenance tasks | [Please refer to a precise section in the user manual.]
| • periodic preventative maintenance checks | [Please refer to a precise section in the user manual.]
| • diagnostic and repair procedures | [Please refer to a precise section in the user manual.]
| • itemized list of spare parts including part numbers | [Please refer to a precise section in the user manual.]
| • end-of-life resource recovery and recycling procedures | [Please refer to a precise section in the user manual.]
4.12 | Training
| Not required. | [Training is not required but could training be available if requested? e.g. this could be online or via a website.]
4.13 | Verification
| In accordance with PQS Verification Protocol E003/FZ01-VP.2. | [Please refer to the laboratory test report.]
5. | Packaging
| Materials used for packaging the finished product are to be free of ozone-depleting compounds as defined in the Montreal Protocol. The general specification of shipping containers will be subject to agreement with the individual procurement agencies. | [Please state the packaging materials and how they comply with the requirements of this clause.]

## Documents submitted
Document reference | Document description