## Sheet1
PQS submission template for a Vaccine freezer or combined vaccine and water-pack freezer: Solar-direct drive |  |  | Brand & Model:
| Tested in accordance with PQS verification protocol: | E003/FZ04-VP.1 | Company name:
| To verify compliance with PQS specification: | E003/FZ04.1
Main contact:
Instructions in [square brackets] are for your guidance. Please delete these instructions and replace with evidence e.g. a precise reference (including page and section numbers) showing compliance with that particular PQS specification requirement. In general a "pass" or "complies" statement alone is not acceptable - evidence for compliance must be provided.
One document may provide evidence for several clause requirements. Please carefully reference the sub-sections of any such document so that it is clear that the clause requirement has been met. This will help speed the pre-qualification process.
Greyed-out cells are for information and do not require a response.
NOTE: If it is not clear or doubtful whether a requirement has been met, then this is a non-compliance and pre-qualification might be delayed.
|  |  | REVIEWER’S SUMMARY: |  | REVIEWER’S SUMMARY: |  | REVIEWER’S SUMMARY:
Section 7: | Product dossier | Applicant's submission [Name & Date] | Reviewer's remarks [Name & Date] | Applicant's response [Name & Date] | Reviewer's remarks [Name & Date] | Applicant's response [Name & Date] | Reviewer's remarks [Name & Date]
The legal manufacturer or reseller must provide WHO with a prequalification dossier containing the following:
| Dossier examination fee in US dollars. | [Please arrange payment with PQS.]
| General information about the legal manufacturer, including name and address. | [Please state.]
| Unique identification reference for the product type. | [Please state.]
| Brand name of the product. | [Please state.]
| Full specifications of the product being offered, covering all the requirements set out in this document, including details of product marking and traceability. | [Please list all documents included with the dossier submission including product markings and traceability.]
| Full details of the recommended compatible Type 2 solar power system (see specification WHO/PQS/E003/PV01). | [Please refer to a document with full details.]
| A comprehensive set of photographs including a views of all external sides of the appliance with the door open. Take additional photographs showing the interior layout including optional configurations (e.g. with baskets or shelves) the compressor or cooling system and a close-up of the thermometer, indicator light(s), the control(s), control panel and any special features. | [Please refer to a document showing the photos.]
| Certified copies of all type-approvals obtained for the appliance, including UL marking, CE marking and the like. | [Please refer to certified copies of all type approvals.]
| Certified photocopies of the legal manufacturer’s ISO 9001 quality system certification. | [Please refer to certified copies of the ISO 9001 certificate.]
| Certified photocopies of the legal manufacturer’s ISO14001 certification, EMAS registration or registration with an equivalent environmental audit scheme. Conformity with an environmental audit scheme is not mandatory however, preference will be given to manufacturers who are able to demonstrate compliance with good environmental practice. | [Please refer to certified copies of the ISO 14001 certificate or some other environmental conformity scheme.]
| Laboratory test report(s) proving conformity with the appliance specifications as tested by a PQS accredited laboratory. | [Please refer to laboratory test report(s).]
| Indicative cost of the appliance per unit, per 10 units and per 100 units, EXW (Incoterms 2015) including appliance and solar power system. | [Please state or refer to a document.]
Section 4. | Requirements
4.1 | General
| Solar powered vaccine storage freezers will be used primarily where there is no electricity or where the electricity is unreliable. These freezers are not to be used for simultaneous vaccine storage and water-pack freezing unless provided with separate dedicated compartments with separate cooling and control systems. These freezers are not for ultra-low temperature storage (i.e. -40°C or colder). The associated power system must be designed to match both the freezer energy consumption and local climate conditions (i.e. ambient temperatures and solar radiation resource). All appliances must fully meet the stated performance requirements with a minimum solar radiation reference period of 3.5 kWh/m²/day and hot zone (+43°C) test temperatures. In addition, manufacturers may offer appliances suitable for hot zone and a lower temperature zone and/or lower solar radiation reference period. | [Please state product type and refer to a specification sheet.]
4.1.1 | Field evaluation
| At the discretion of the WHO PQS Secretariat and prior to PQS prequalification a WHO PQS pre-approved field evaluation may be required in accordance with the requirements of the Generic Guide for the Field Evaluation of New Technologies for WHO PQS Prequalification. | [Please contact the WHO PQS Secretariat (if not already done so) to discuss whether a field evaluation will be necessary.]
4.2 | Performance
4.2.1 | Operating temperature range
| The operating temperature range is indicated on the temperature zone symbol attached to the appliance (see Annex 1). All appliances must be suitable for hot zone (+43°C). Solar direct drive (SDD) vaccine freezers are not required to indicate a minimum rated ambient temperature (note: some other products in the WHO PQS E003 appliance category are required to indicate this rated temperature). | [Please provide a photo of the temperature zone symbol as attached to the appliance.]
4.2.2 | Refrigeration cycle
| Electrically powered compression-cycle, Stirling engine, thermoelectric freezers operating with input of direct current (DC) electricity. Other technologies with equal or better performance may be considered. | [Please state the exact product type.]
4.2.3 | Voltage
| Solar power system input voltage up to 45 Voc is acceptable. | [Please state working voltage or working voltage range.]
4.2.4 | Vaccine storage capacity
| Areas of an otherwise acceptable appliance which are too warm for vaccine storage must be excluded from use by design. No minimum vaccine storage capacity volume is specified. Net vaccine storage capacity volume will be established and reported per WHO/PQS/E003/FZ04/VP.1. | [Please show (photos or diagram) how areas that are too warm for vaccine storage are excluded from use. Please refer to the test report section that shows measured vaccine storage capacity volume.]
4.2.5 | Water-pack freezing and storage compartment capacity (optional).
| If included, the water-pack freezing compartment shall be capable of freezing a minimum mass of water based on the volume of the compartment over a 24-hour period. For gross water-pack freezing compartment volume of less than 50 litres, the appliance shall be capable of freezing a minimum of 1.6 kg of water. For compartments greater than 50 litres the appliance shall be capable of freezing a minimum of 2.4 kg per each 50 litres of gross storage volume. The water-pack freezing compartment shall hold a minimum of 3.2 kg of water-packs and at least twice the daily water-pack freezing capacity determined by WHO/PQS/E003/FZ04//VP.1. The water-packs must comply with WHO/PQS/E005/IP01. | [If applicable, please refer to the test report section that shows compliance with this requirement.]
4.2.6 | Simultaneous vaccine storage and water-pack freezing warning
| These appliances are not suitable for simultaneously storing frozen vaccine and water-pack freezing unless separate compartments (i.e. at least one compartment being dedicated to vaccine storage and other compartments dedicated to water pack freezing), with separate temperature control systems, are provided. All appliances must carry a factory-fitted non-removable label, designed to last the lifetime of the appliance, carrying the words: ‘When storing frozen vaccine do not freeze water-packs or store any other items in the vaccine storage compartment’ in letters no smaller than 20 mm high in the languages specified in clause 4.11. The label should be fixed to the lid of chest freezers and near the top of the door on upright freezers. | [Please state whether your appliance can store frozen vaccine and freeze waterpacks. Refer to photos of the a factory-fitted non-removable label showing the wording in detail and also the position on the appliance.]
4.2.7 | Temperature control
| All freezers: The vaccine load must remain within the acceptable temperature range during any continuous ambient temperature test(s) or day/night cycling temperature test(s). Combined freezers must achieve this performance with no water packs in the water-pack compartment. | [Please refer to the test report section that shows compliance with this requirement.]
| Combined freezers only: While simultaneously freezing a quantity of water-packs equal to its water-pack freezing capacity, the temperature of the full load of vaccines must remain at -15°C or colder under the maximum continuous ambient temperature test conditions of its rated temperature zone. Dedicated water-pack freezing compartments must remain at -5°C or colder during freezing operations. | [If applicable, please refer to the test report section that shows compliance with this requirement.]
4.2.8 | Thermostat
| A thermostat must be provided that is effective throughout the ambient operating temperature range. A vaccine compartment thermostat must be designed so that it cannot be adjusted by the user. A means for adjustment by a technician is acceptable provided the device is protected from user interference (e.g. by location within the appliance cabinet). Alternatively, programmable thermostats may be password-protected. Bulb and capillary tube thermostats are not acceptable. | [Please specify all thermostat controls, their siting and how they may be adjusted.]
4.2.9 | Thermometer
| Option A: Electronic thermometer powered by a photovoltaic cell which forms part of the device. This type draws no power from the appliance in which it is installed. | [Please state which thermometer is built into your appliance and refer to a photo.]
| Option B: Externally readable cabinet-mounted electronic thermometer. | [Please state which thermometer is built into your appliance and refer to a photo.]
4.2.10 | Indicator light
| A minimum of one green LED indicator light per vaccine storage compartment is required to be located on the front or top of the appliance to alert users that the cooling system is actively operating. A constant green LED light is required to indicate that the compressor or cooling system is active, and the light must go off when the compressor or cooling system is off. Optionally, additional indicator lights may be added to indicate other operating conditions including water pack freezing operations, temperature and faults. | [Please specify all indicator lights and where they are sited on the appliance (photos or diagram). Refer to a particular section in the user manual.]
4.2.11 | Autonomy
All solar direct drive (SDD) vaccine freezers or combined vaccine freezer and water-pack freezer must have a minimum autonomy of three days at the minimum solar radiation reference period and when tested for the hot zone unless an alternative autonomy (of greater days) for a specific location(s) can be calculated using one of the two methods defined in the normative reference: Toma, H. and Markvart T. Solar Autonomy Calculation Tool, University of Southampton, UK, 2009. If an autonomy of greater than three days can be calculated for a specific location(s) then the installation(s) must then be sized to provide the required autonomy at the specified location(s). In no case can the solar powered frozen vaccine storage freezer have an autonomy of less than three days. The two autonomy calculation methods are:
| 1. selection of autonomy for the specific locations listed in the supporting document to the Solar Autonomy Calculation Tool² or | [Please show your calculation if applicable.]
| 2. calculation of autonomy using the formulas and the required long term daily solar radiation data as described in the Solar Autonomy Calculation Tool². | [Please show your calculation if applicable.]
| Autonomy of appliance is determined by testing in accordance with WHO/PQS/E003/FZ04/VP.1. | [Please refer to the test report section that shows compliance with this requirement.]
NOTE 2 Instructions for Autonomy Calculation, Table 1 includes a subset of sites for which the data are considered to be sufficiently accurate for solar vaccine storage applications.
4.2.12 | Power system requirements
| PQS specification WHO/PQS/E003/PV01 specifies a compatible Type 2 solar power system to directly power the appliance. A battery powered/assisted cooling system for freezer operation will not be accepted. Ancillary power storage systems (e.g. capacitor) for nonessential cooling functions may be included provided these have a minimum guaranteed design life of 10 years under the environmental conditions for which the freezer is prequalified. Solar module voltage up to 45 volts open circuit (Voc at Standard Test Condition of solar radiation 1000 W/m², cell temperature +25°C, air mass 1.5) is acceptable provided all electrically powered system components are integrated in such a way that performance and component life is not reduced by voltage input from the solar array. | [Please state or refer to a document showing all specifications of the solar module configuration necessary to power your appliance.]
4.2.13 | Evaporator configuration, humidity control and defrosting
| If an evaporator is mounted in shelves the minimum clearance between shelves must be 130 mm. Freezer may be defrosted manually or have automatic defrosting. Defrost drainage must be supplied by design or by including a drain with plug where condensate can accumulate. If active defrost used, the defrost switch(es) and heaters must be accessible to the user without tools but must be protected from accidental changes in position. If a separate compartment for water pack freezing is provided legal manufacturers are to propose water-pack placement instructions, racks and/or structure to prevent frozen water-packs from adhering to other water-packs or adhering to freezer walls and/or other surfaces. Spare parts will be required for all drain plugs, switches and heaters. | [Please state the evaporator configuration. The method of defrosting. The method of draining the condensate. The operation of active defrost. Ease of removal of frozen waterpacks.]
4.2.14 | Lock
| The door or lid must be fitted with a lock or a means of locking (e.g. with a padlock). Two keys are to be supplied with every unit. | [Please refer to the appropriate section and page in the user manual.]
4.2.15 | Corrosion resistance
| The legal manufacturer must declare compliance that internal and external components are protected against corrosion as appropriate to EN ISO 6270-1 / ASTM D2247 / EN 13523-26, EN ISO 6270-2 / EN 13523-25, ISO 6272 / EN 13523-5 and ISO 2409. Evidence for compliance shall be demonstrated. | [Please state the corrosion protection for internal and external components and provide evidence for this level of protection.]
4.2.16 | Electrical safety rating
| The legal manufacturer must declare compliance with IEC 60335-1, IEC 60335-2-24 and IEC 60364-1. Evidence for compliance shall be demonstrated. | [Please refer to the test report that proves electrical safety compliance.]
4.2.17 | Markings
| If used, compressors must be marked with the blue identifying symbol shown in Annex 2. In addition, the cabinet must be permanently marked, near the compressor position, with the chemical name of the refrigerant, or with the refrigerant number, formula or proportion (for blended refrigerants). Appliances not utilizing compressor-based cooling methods must be marked identifying any refrigerant and/or heat transfer fluids used including, but not limited to, the chemical name, formula or proportion (for blended refrigerants). All appliances to label hazardous materials and include a Safety Data Sheet. Label and Safety Data Sheet must comply with the Globally Harmonized System for the Classification and Labelling of Chemicals GHS Rev.5. All appliances require an asset identification label (bar code) as specified in WHO/PQS/E003/BC01.1: Global asset identification. | [Please provide photos showing all markings and labels on the appliance in accordance with Clause 4.2.17.]
4.2.18 | Labelling
| All units must carry a factory-fitted non-removable label, designed to last the lifetime of the appliance, carrying the following information: • All freezers: Warning against simultaneous frozen vaccine storage and water-pack freezing in same compartment per clause 4.2.6. Warning against storing ultra-low temperature vaccines. • Vaccine freezers: Vaccine storage instructions and the appropriate temperature zone symbol as Annex 1. • Combined freezers: Vaccine storage instructions, water-pack freezing instructions, simultaneous vaccine storage and water-pack freezing warning and the appropriate temperature zone symbol as Annex 1. All key information must be pictorial and summarized on a single sheet fixed onto the appliance cabinet. The sheet should be sufficiently durable to last the life of the appliance. The instructions should be fixed to the lid of chest freezers and near the top of the door on upright freezers. Labelling must be available in the languages as specified in clause 4.11. The appliance must carry the following additional information fixed to the front of the cabinet: manufacturer and model number (unless already located on the front of the unit), serial number, date of manufacture, PQS identification number, applicable service phone number and website URL. | [Please provide photos showing all labels on the appliance in accordance with Clause 4.2.18.]
4.2.19 | Electromagnetic compatibility
| The legal manufacturer must declare compliance with the requirements of the latest edition of IEC 61000-6-1 and IEC 61000-6-3. Evidence for compliance shall be demonstrated. | [Please refer to the test report that proves EMC compliance.]
4.2.20 | Equipment monitoring system
| Future revisions of this specification will require that the legal manufacturer must certify compliance with the requirements of the latest edition of WHO/PQS/E006/EMS 01. | Please contact PQS for up-to-date information on compliance requirements.
4.3 | Environmental requirements
4.3.1 | Ambient temperature range during transport and storage
| The ambient temperature range during transport and storage is -30°C to +70°C when the appliance is inactivated. | [Please provide evidence of compliance.]
4.3.2 | Ambient humidity range during transport, storage and use:
| The ambient humidity range during transport, storage and use is 5% to 95% RH, non-condensing. | [Please provide evidence of compliance.]
4.4 | Physical characteristics
4.4.1 | Overall dimensions
| To allow for manoeuvring through corners, corridors and doorways, the minimum dimension of the appliance (either length, width or height) should not exceed 710 mm; exceptionally a minimum dimension up to 830 mm can be accepted, but this will restrict the number of sites where the appliance can be installed. The maximum dimension must not exceed 1700 mm and the maximum diagonal (corner to corner) dimension must not exceed 1850 mm. | [Please refer to the test report section that shows measured dimensions.]
4.4.2 | Weight
| Mechanical lifting equipment will typically not be available at the installation sites. It is recommended that the freezer and any associated components should be designed for lifting in such a way that no single worker is required to carry more than 25 kg whilst working on their own, or in a group. | [Please refer to the test report section that shows measured weight.]
4.4.3 | Spacers
| If air clearances around or on particular sides are critical for appliance performance (e.g. to protect the condenser), spacers or similar must be provided (e.g. as screw-in rods). For clearances greater than 50 mm (e.g. for side venting) prominent warning signs in letters no smaller than 20 mm high must be provided. If no spacers or warning signs are provided, then it can be assumed that air clearances are not critical. | [Please state the required air clearances and/or refer to the user manual section/page number. Please provide photos of the spacers provided in various positions on the appliance.]
4.5 | Interface requirements
4.5.1 | Electrical components
All electrical components must be compatible with a Type 2 solar power system as specified in specification WHO/PQS/E003/PV01.
4.5.2 | Power switch
| Each appliance to provide an on and off power switch that is readily accessible to the user either on the outside of the appliance cabinet or in a wall-mounted switch within one meter of the appliance. Switches must be rated for DC voltage and currents required by the SDD appliance. Plug and play solar array cable connectors (i.e. locking female and male coupler system) are not acceptable for on and off switching of the solar power system or appliance. | [Please provide a photo of the power switch or refer to the appropriate section/page number in the user manual and specify the rating of this switch.]
4.6 | Human factors
4.6.1 | General design
| The appliance must be useable by the widest practicable range of active health workers, regardless of age, gender, size or minor disability, including colour blind users and long-sighted (hyperopia) people without glasses, in accordance with the general principles laid out in ISO 20282-1: 2006. | [Please state how the design of your appliance conforms to these requirements.]
4.6.2 | Control panel and thermometer
| Controls panel, indicator light, thermometer and other visual displays may be positioned on the front of the unit preferably as close to eye level as possible. Alternatively, they may be mounted on top of the unit at a height not exceeding 1.3 metres. If a low-level position is essential, the display should be aligned so that it can easily be read without the user having to squat or kneel down. The on and off switch and/or defrost switch, if present, should be recessed or otherwise protected from accidental change in position so that it is not possible to inadvertently activate it. | [Please provide a photo of the control panel, lights, thermometer and visual display indicating where these are sited on your appliance.]
4.7 | Materials
4.7.1 | Refrigerant
| Refrigerants with global warming potential (GWP) ≤ 11 and zero ozone depletion potential (ODP) must be used. The suitability of alternative refrigerant gases will continue to be assessed and preference will be given to appliances that use gases with low global warming potential (GWP). | [Please state the refrigerant used with its GWP.]
4.7.2 | Thermal insulation foaming agents
| Any gas complying with the limitations and deadlines set by the Montreal Protocol on the elimination of ozone-depleting chemicals with global warming potential (GWP) ≤ 11 and zero ozone depletion potential (ODP) must be used and clearly specified. | [Please state the foaming agent used with its GWP.]
4.7.3 | Other restricted materials
| The appliance and its constituent components must not contain lead, mercury, cadmium, hexavalent chromium, polybrominated biphenyls (PBB) or polybrominated biphenyl ethers (PBDE). | [Please state how you can be sure that these materials are not in the components of your appliance.]
4.7.4 | PCM
| Integrated phase change materials (PCM) may be used for thermal purposes and must comply with WHO/PQS/ E005/PCMC0.1 specification for phase-change material containers. | [Please state any PCM used and refer to the test report that shows compliance with WHO/PQS/ E005/PCMC0.1.]
4.8 | Warranty
| The appliance must be covered by a three-year replacement warranty in the event of any component failure arising from defective design, materials or workmanship. Ancillary power storage (e.g. capacitors) to be covered by a 10-year warranty. Solar power system warranties as required per WHO/PQS/E003/PV01. | [Please refer to a statement of the warranty with all terms and conditions.]
4.9 | Servicing provision
| The appliance and solar power system must be designed to achieve a low-maintenance life of not less than 10 years (with the exception of routine defrosting, cleaning and solar array cleaning and shading prevention). | [Please state how the appliance and solar power system are designed to last at least 10 years.]
4.9.1 | Essential spare parts
| Based on design and requirements of appliance and solar power system the type and quantity of spare parts, basic installation tools/supplies, user and technician maintenance manuals (see clause 4.11), must be determined and agreed upon in advance of order placement. As a minimum, each appliance and solar power system is to be supplied with 10 spare fuses of all fuse size and type used in the appliance. The spares fuses are to be attached within or on the appliance. Each appliance with drain plugs and active defrost components will be supplied with a minimum of one spare drain plug and one set of active defrost components (e.g. switches and defrost heater). Legal manufacturers are to publish a list of spare parts recommended for purchases of 10 and 50 appliance and solar power systems. The legal manufacturer must ensure supply of spare parts for a minimum of five years from the time of cessation of the last production of equipment. Spare parts are to be provided in kit form for storage in appropriate quantities at National or Sub-national level in the purchasing country, as agreed with the purchasing agency. | [Please show the position of the 10 spare fuses of all fuse size and type within or on the appliance. Also, a drain plug and active defrost components. Please refer to a document listing the spare parts.]
4.10 | Disposal and recycling
| The legal manufacturer must provide information to the buyer on proper disposal, including child-proofing the cabinet, the hazardous materials contained within the system and suggestions for resource recovery/recycling and/or environmentally safe disposal. For legal manufacturers from the European Union WEEE compliance in accordance with European Union Directive 2002/96/EC is mandatory. | [Please state all hazardous materials used in your appliance and refer to the user manual section/page number describing the provision for disposal/recycling in the user manual.]
4.11 | Instructions
Each freezer to include a separate user manual and technician installation manual in Arabic, English, French, Mandarin Chinese, Russian and Spanish. An English version of all instructions and manuals are required to be supplied at time of laboratory testing. Instructions to include easy to understand visuals whenever possible to avoid reliance on text.
The user manual must include the following information:
| Health and safety guidance; | [Please refer to the appropriate section and page in the user manual.]
| Basic operations description; | [Please refer to the appropriate section and page in the user manual.]
| Warning against simultaneous frozen vaccine storage and water-pack freezing; | [Please refer to the appropriate section and page in the user manual.]
| Warning against storing refrigerated vaccines, refreezing vaccines and storing ultra-low temperature vaccines; | [Please refer to the appropriate section and page in the user manual.]
| If PCM included then safe PCM handling in case of leakage; and | [Please refer to the appropriate section and page in the user manual.]
| Preventive maintenance tasks (e.g. daily, weekly, and monthly). | [Please refer to the appropriate section and page in the user manual.]
The technician installation manual must include the following information:
| Health and safety guidance; | [Please refer to the appropriate section and page in the user manual.]
| Detailed operations description; | [Please refer to the appropriate section and page in the user manual.]
| Correct handling to avoid appliance damage and for the safety of handling persons; | [Please refer to the appropriate section and page in the user manual.]
| If PCM included then safe PCM handling in case of leakage; | [Please refer to the appropriate section and page in the user manual.]
| Warning against simultaneous frozen vaccine storage and water-pack freezing; | [Please refer to the appropriate section and page in the user manual.]
| Warning against storing ultra-low temperature vaccines; | [Please refer to the appropriate section and page in the user manual.]
| Installation procedures; | [Please refer to the appropriate section and page in the user manual.]
| Compatible solar power system voltage; | [Please refer to the appropriate section and page in the user manual.]
| Technical maintenance tasks (e.g. daily, weekly and monthly); | [Please refer to the appropriate section and page in the user manual.]
| Periodic preventive maintenance procedures; | [Please refer to the appropriate section and page in the user manual.]
| Corrective maintenance, diagnostic and repair procedures; | [Please refer to the appropriate section and page in the user manual.]
| Itemized list of spare parts including part numbers; | [Please refer to the appropriate section and page in the user manual.]
| End-of-life resource recovery and recycling procedures; | [Please refer to the appropriate section and page in the user manual.]
| User training guidance. | [Please refer to the appropriate section and page in the user manual.]
| Printed user operations and routine maintenance instructions specifically directed at the health centre or store staff must be pictorial. All key information must be summarized on a single sheet pictogram fixed onto the appliance cabinet per clause 4.2.18 - Labelling; the sheet should be sufficiently durable to last the life of the appliance. | See above.
| Installation, repair and servicing instructions must be supplied in printed format, and optionally on DVD and/or on-line to instruct the installation teams in installation standards and practices specific to the appliance and its power system. In addition, supporting video material supplied on DVD and/or on-line can be supplied to assist the instructor when delivering on-site user training. | [Please provide a copy of the Installation, repair and servicing instructions and state where these can be found if available online.]
4.12 | Training
| If requested, all legal manufacturers are required to have the capability of providing in-person training in the countries where their product is deployed. Training may be delivered by the legal manufacturer, manufacturer’s representative or reseller. The legal manufacturer must provide user training guidance to enable installation technicians to present user training. In addition, supporting video material supplied on DVD and/or online can be supplied to assist the instructor when delivering on-site user training. | [Please state how training could be provided if requested, e.g. in-person, on video.]
4.13 | Verification
| In accordance with PQS Verification Protocol WHO/PQS/E003/FZ04-VP.1. | [Please refer to the test report.]
5. | Packaging
| Materials used for packaging the finished appliance are to be free of ozone-depleting compounds as defined in the Montreal Protocol and Kigali Amendment 2016. The packaging must be a sturdy export quality and of a commercial standard that will provide adequate protection of the goods for carriage by air, sea and/or road to final destinations worldwide, including remote locations under adverse climatic and storage conditions and high humidity. The packaging must be not less than 17kN edge crush resistance with minimum 60% remaining with 90% humidity at a temperature of +70°C (tropical conditions). To avoid destructive unpacking prior to installation, legal manufacturers are encouraged to add a re-sealable observation opening in their packaging to aid inspectors in finding labelling and/or placing additional markings prior to installation. Instructions on the packaging alerting inspectors to use of the opening and what information will be revealed are also advised. The general specification of shipping containers will be subject to agreement with the individual procurement agencies. | [Please describe the packaging and refer to the user manual section/page number on how this packaging may be disposed / recycled.]