## Sheet1
E003 Submission to PQS for approval: Solar power system for vaccine refrigerator or combined vaccine refrigerator and water-pack freezer submission template (V1.0 RFR 05 MAY 2025)
| Assessment of appliance: | [Brand & Model here] | For dossier review on [TBC]
| Tested in accordance with PQS verification protocol: | E003/PV01-VP1.2 (2 August 2007)
| To verify compliance with PQS specification: | E003/PV01.4 (2 February 2018)
ASSESSOR’S SUMMARY:
Instructions in [square brackets] are for your guidance. Please delete these instructions and replace with evidence or a precise reference (including page and section numbers) showing compliance with that particular PQS specification requirement. In general a "pass" or "complies" statement alone is not acceptable.
One document may provide evidence for several clause requirements. Please carefully reference the sub-sections of any such document so that it is clear that the clause requirement has been met. This will help speed the pre-qualification process.
Greyed-out cells do not require a response.
NOTE: If it is not clear or doubtful whether a requirement has been met, then this is a non-compliance and pre-qualification might be delayed.
|  |  | REVIEWER’S SUMMARY: |  | REVIEWER’S SUMMARY: |  | REVIEWER’S SUMMARY:
Section 7. | Product dossier | Applicant's submission | Reviewer's remarks [Name & Date] | Applicants response [Name & Date] | Reviewer's remarks [Name & Date] | Applicants response [Name & Date] | Reviewer's remarks [Name & Date]
The legal manufacturer or reseller is to provide WHO with a pre-qualification dossier containing the following:
a. | Dossier examination fee in US dollars. | [Please contact PQS.]
b. | General information about the legal manufacturer, including name and address. | [Please state or refer to a document.]
c. | General information about the reseller, including name and address (where applicable). | [If a reseller is used, please state or refer to a document.]
d. | A list of at least 50 solar photovoltaic systems installations in one or more less-developed countries completed and supported over a period of at least two years prior to the date of dossier submission. Provide full details of each installation, including details of installed equipment, country and location, procurement agency and named contacts. | [Please state or refer to a document.]
e. | A list of the countries or regions in which the legal manufacturer or reseller is able to offer the optional support services described in Clauses 6 (On-site installation) and 8 (On-site maintenance), together with details of these services. | [If available, please state or refer to a document.]
f. | Unique identification reference for the system type. | [Please state.]
g. | System type offered (Type 1 [battery], Type 2 [SDD] or both). | [Please state.]
h. | Full specifications of the components being offered, covering all the requirements set out in this document, including details of product marking and traceability. | [Please refer to a document.]
i. | Certified photocopies of all type-approvals obtained for the product, including CE marking and the like. | [Please refer to a document & section.]
j. | Certified photocopies of the legal manufacturer’s ISO 9001 quality system certification. | [Please refer to a document.]
k. | Where relevant, certified photocopies of the legal manufacturer’s ISO 14001 certification, EMAS registration or registration with an equivalent environmental audit scheme. Conformity with an environmental audit scheme is not mandatory; however, preference will be given to manufacturers who are able to demonstrate compliance with good environmental practice. | [Please refer to a document.]
l. | Where available, laboratory test report(s) proving conformity with the product specifications. | [Please refer to a document.]
m. | Indicative cost of the product per unit, per 10 units and per 100 units, EXW (Incoterms 2010) including appliance and solar power system. | [Please state or refer to a document.]
Section 4. | Design criteria
4.1 | General
4.1.1 | System characteristics
The solar power system will be connected to appliances that comply with PQS specifications E003/FZ03, E003/RF04, E003/RF05. This equipment is used primarily in areas without any electricity or where there is less than four to eight hours of reliable electricity over a typical day. Reliability, durability and effective maintenance are essential for a successful installation.
Each power system must be purpose-designed to match the power consumption of the connected appliance and the site-specific climate conditions, taking account of ambient temperatures, available solar radiation resource and autonomy requirements. In all cases the power system must provide the power capacity (Watts) equal to or greater than the power system capacity used for PQS prequalification testing. Each power system must be complete to allow the entire installation to occur without the need for additional components.
No additional loads, such as lighting or pumping, are to be connected to the solar power system.
4.1.2 | Design responsibility
The procurement organization will specify the capacity and characteristics of the appliance that is to be connected to the solar power system. The installation must then be sized to ensure that the specified appliance will operate continuously within specified temperature limits, both when solar radiation is powering the system and throughout the autonomy period.
| All solar systems must provide a minimum of five days of autonomy unless an alternative autonomy (of greater or lesser days) can be calculated using one of the two methods defined in the normative reference: Toma, H. and Markvart T. Solar Autonomy Calculation Tool, University of Southampton, UK, 2009. | [Please state system autonomy and provide calculations of minimum autonomy requirement.]
| The two methods for calculating autonomy are: 1. Selection of autonomy for the specific locations which are listed in the supporting document to the Solar Autonomy Calculation Tool. 2. Calculation of autonomy using the formulas and the required long term daily solar radiation data as described in the Solar Autonomy Calculation Tool. | [Please state which method was used to calculate system autonomy and provide calculations.]
Two design cases apply: when the installation site is known, and when it is unknown. • Known sites: The procurement organization will specify, as accurately as possible, the site location for each installation. This will include including the country name, site name, latitude and longitude and, where possible, the elevation in meters above sea level. • Unknown sites: The procurement agency will specify the country name and, where possible, the region within the country.
| In both cases the qualified supplier must then design the installation to provide the required autonomy anywhere within the specified geographical area. In both cases, the qualified supplier must provide the procurement agency with the data used to determine the installation design, including reference periods for temperatures, solar radiation reference period, autonomy determination and assumptions made regarding site conditions. The data must include the assumed monthly minimum solar radiation reference period, assumed average minimum and maximum ambient temperatures and the month requiring the most power. Reference period climate conditions are to be in closest accordance with available meteorological data. | [Please state system autonomy and provide calculations of minimum autonomy requirement. Please provide sizing calculations for the solar array.]
4.2 | Performance
4.2.1 | Photovoltaic array and array output cable
| Solar modules must comply with IEC 61215 (crystalline solar modules) or IEC 61646 (thin film solar modules). Individual solar module open circuit voltage of up to 45 Voc is acceptable provided all electrically-powered system components are integrated in such a way that performance and component life is not reduced by voltage input from the solar array. The power warranty is to be a minimum of 25 years to 80% of the initial power rating. The minimum solar array capacity is to be no smaller than an array capable of sustaining the maximum load, including all power system losses multiplied by a factor of 1.25. | [Please state type of photovoltaic module used and refer to document showing IEC certification, maximum output voltage, and power warranty. Refer to sizing calculations showing power losses as well as minimum 1.25 design margin.]
| Except for equipment grounding conductors all solar module cables, module to module interconnect cables and solar array output cables must provide a locking female and male coupler system with minimum IP rating of 65 (“plug and play”) for cable to cable interconnection. They must also be equipped with strain-relief fittings unless otherwise permanently fixed or in conduit. All cable and connectors not in conduit must display a rating confirming its suitability for exterior use in direct sunlight, and for direct burial if used underground. Exterior-rated conduit with weatherproofing or equal protection must be provided at the point where the cable enters the building. A minimum solar array output cable length of 20 meters must be supplied. A separate equipment grounding conductor of 20 meters must also be supplied. Manufacturers are to provide instruction for estimating cable diameter and for ordering non-standard lengths of cable. | [Please refer to document or provide details as to how cabling is compliant to the requirements of clause 4.2.1.]
4.2.2 | Solar array support structure
| The photovoltaic array support structure forms part of the installation and must be constructed of anodized aluminium, stainless steel or steel, hot dip galvanised to ISO 1461 after cutting and drilling. Generalized roof/ground mounts with adjustable tilt angle of up to 45 degrees and optional pole mount structures (excluding pole and concrete) must be offered by all manufacturers. Generalised roof/ground mounts with adjustable tilt angle of up to 45 degrees are to be supplied unless one of the following location-specific mounts is specified at time of procurement: | [Please state type of solar array support structures offered. Please state the fabrication material for the support structure. Confirm degrees of tilt angle adjustment.]
| • Pitched-roof mounting: The support frame is provided with adjustable tilt angle of up to 45 degrees for connection to a pitched roof structure. Any type of pitched roof finish may be encountered, including corrugated sheeting, tiles and slates. | [Please state if pitched-roof mount is offered.]
| • Flat-roof mounting: The support frame should hold the solar array at the optimum angle and orientation for the site. Depending upon the roof construction, the frame may either be mechanically fixed to the roof or held down with weights. | [Please state if flat-roof mount is offered. Please state if the structure is mechanically attached or weighted]
| • Wall mounting: The support frame should hold the solar array at the optimum angle for the site. Orientation will follow the alignment of the wall. Any type of wall construction may be encountered, including brick, stone, concrete, steel or timber framing. | [Please state if wall mount is offered.]
| • Ground mounting: The support frame should hold the solar array at the optimum angle and orientation for the site. The frame will either be bolted to concrete pads or to a concrete slab. | [Please state if ground mount is offered.]
| • Pole mounting: The support frame should hold the solar array at the optimum angle and orientation for the site and may be top mounted or side mounted. It is not required to supply pole, concrete and maintenance tools. However, these materials and any required tools are to be clearly specified to prospective buyers in advance of purchase. | [Please state if pole mount is offered.]
| The combined array structure and photovoltaic panel assembly must be designed to withstand wind loads at site of installation. Sites that experience Category One hurricane and/or F1 tornado must have a solar array certified or engineered to be capable of withstanding local wind conditions. | [Please state maximum wind load design of structure.]
| A solar array must be able to withstand the full force of the wind (wind-load force) expected at the installation site whether it is fixed to a roof, ground, building or pole structure. This force should be calculated from knowledge of local conditions including wind speed and local features. Particular attention should be given to force caused by wind gusts and possible resonance conditions. Guidance for calculating the wind-load force can be found in internationally-accepted building codes. When site wind conditions are unknown, but it is known that the installation site does not experience Category One hurricane or F1 tornado, then the structure with solar array must provide wind load resistance equal to the installed solar module complying with IEC 61215. | [Please provide calculations for wind load design of structure.]
| All solar array structure fasteners accessible to unauthorised persons must be theft deterrent, including fasteners required to fix the structure to building, foundation or pole. All solar array structures packages are to be supplied with a complete set of all unique tools required for fixing the theft deterrent fasteners supplied. It must be possible for an authorised person to adjust or remove the array assembly if required. Breakaway nuts and similar devices are not acceptable. | [Please describe type of theft deterrent hardware offered.]
| The generalised roof/ground mount must include all fasteners that are needed to attach to a pitched roof, flat concrete roof or ground foundation, including theft deterrent fasteners that are required to attach to a pitched roof with wood structure and flat concrete on either an existing roof or ground foundation cast at time of installation. The array structure package is also to be supplied with sealant for weather-proofing building penetrations and all other necessary subsidiary components, including foundation fasteners with necessary drill bits and the components needed to bond the solar array and the structure for lightning protection purposes, in accordance with the solar module manufacturer’s instructions. | [Please describe type of attachment hardware offered.]
| Where site-specific works are required to install the support structure, these will be designed by the qualified supplier and constructed by a local contractor. | [If applicable, please describe any site-specific works required to install the support structure.]
4.2.3 | Battery set sizing
| Batteries must be sized to meet the autonomy period determined as described in clause 4.1.2, based on the energy required for the appliance(s) to which the solar system components are connected. In no case should the battery have a capacity of less than three days. Required battery capacity is estimated over the discharge rate and temperature most closely matching the autonomy period and the temperature at the site (e.g. at C/120 hours at +25°C for a five-day autonomy period to a final voltage that ensures the minimum required battery set life of 1800 cycles to 80% discharge at +20 °C). Battery capacity must be sufficient to ensure that the designed autonomy period is maintained throughout the minimum required battery set life of 1800 cycles to 80% discharge | [If applicable, please describe how the battery has been sized to meet the requirements of clause 4.1.2.]
4.2.4 | Battery type
| Both sealed and flooded lead acid batteries are acceptable for Type 1 systems, although sealed batteries are preferred for installations in remote areas where maintenance will be difficult. Only sealed lead acid batteries are acceptable for Type 2 ancillary batteries. Lithium batteries are not permitted. | [If applicable, please describe the type of battery offered.]
| The maintenance cycle for flooded batteries must be six months or greater. They must have clear casings to allow the user to inspect electrolyte levels and they must be supplied dry/charged with the electrolyte packed in separate hermetically-sealed containers. | [If applicable, please describe how the battery complies with these requirements.]
4.2.5 | Battery set housing (Type 1 systems only)
| The battery set must be housed within a lockable ventilated cabinet, which has a clear viewing window for maintenance inspection by the user. The housing for flooded batteries must include protection against spilled electrolyte. Circuit breakers or fuses must be installed in the positive line, near the battery, and the fuse holder must be of non-corroding material. If fuses are used there must be a set of 10 spare fuses for each fuse size and type used, and they must be supplied in a polyethylene bag fixed near to the fuse box. | [If applicable, please describe how the system complies with these requirements.]
4.2.6 | Battery charge regulator (Type 1 systems only)
| Battery charge regulator type may be pulse width modulated, series switching, shunt type or maximum power point control. All regulator types must limit the number of field connections to no more than six conductor connections (e.g. solar + and - battery + and -, appliance + and -). Additional battery sensing connections are acceptable. | [If applicable, please state the type of battery charge regulator and how the regulator complies with these requirements.]
| Maximum power point controls are acceptable when used to connect a solar array of higher voltage to a battery of lower voltage. The battery charge regulator must be sized for a DC current rating of at least 150% of the solar array short circuit rating. The load disconnect current rating must be 200% of the maximum load current and must operate the appliance compressor under all starting-load surge conditions. Maximum power point controls are not permitted to increase field wiring complexity compared to other regulator types. | [If applicable, please state how the regulator complies with these requirements.]
| The regulator must use either varistor, spark-gap or surge protection. The regulator must be pre-set to suit the installed battery type. Voltage regulation set points must not be user-adjustable. The set points must be set precisely to meet the charge and temperature requirements of the installed battery set and they must disconnect the load when the battery has reached a state of charge which can be repeated a minimum of 1800 cycles to 50% depth of discharge. If battery voltage drops below the load disconnect set point, the load must be disconnected and the solar array must remain connected. The load must be automatically reconnected when the system voltage recovers. | [If applicable, please state how the regulator complies with these requirements.]
| The regulator must automatically compensate for temperature changes. If temperature compensation becomes disabled, the regulator must regulate at voltage regulation set points at +25°C. | [If applicable, please state how the regulator complies with these requirements.]
| The regulator unit must have a clearly marked color-coded voltmeter, LED’s or other indicators to show: • Array charging (green); • Low battery (orange or yellow), with a clearly labelled warning: Do not freeze water-packs in the appropriate language; • Load disconnected (red). | [If applicable, please state how the regulator complies with these requirements.]
| An acoustic alarm may be included as a supplementary high/low battery and/or load disconnect warning device. | [If applicable, please state how the regulator complies with these requirements.]
4.2.7 | Lead acid battery safety kit (Type 1 systems only)
| Comprehensive safety instructions, hazardous materials data sheets and a lead acid battery safety equipment kit must be provided. The safety equipment for lead acid flooded batteries is to include the following: • Temperature-compensated hydrometer, • One-litre plastic decanting jug, • Plastic filling funnel, • Eye, hand and clothing protection, • Sodium bicarbonate for cleaning electrolyte spills, and • A hand pump for electrolyte is to be provided if the electrolyte is supplied in containers of more than five litres. | [If applicable, please refer to document showing what safety items are included with the battery.]
4.2.8 | Power switch
| At least one method to isolate (disconnect) the appliance from the solar power system must be supplied. This switch can be built into the appliance or wall mounted. Disconnecting cables is not an acceptable means of disconnect, although this method can be used in case of emergency. | [Please refer to photo or other document showing or describing disconnect between the appliance and the solar system.]
4.2.9 | Electrical safety rating
| The manufacturer must certify compliance with IEC 60335-1, with particular reference to Section 8: Protection against access to live parts, and Annex B: Appliances powered by rechargeable batteries. All on-site electrical installation work must comply with IEC 60364-1. | [Please reference the test information that proves compliance to this clause.]
4.2.10 | Electrical protection
| The system design must prevent damage to all components in the event of short circuits, electrical storms and reversed polarity connections. Circuit breakers or fuses must be installed in the positive line, near the battery, and the fuse-holder must be of non-corroding material. The fuse type and rating must be marked close to the fuse holder. Ten spare fuses must be provided in a polyethylene bag fixed near to the fuse box. | [Please describe how the system design prevents damage in the event of short circuits, electrical storms, and reverse polarity. For type 1 systems please describe or show photos of overcurrent protection.]
4.2.11 | Lightning surge protection
| Lightning surge protection must be provided per manufacturers requirements. At minimum, a complete wire and hardware kit must be supplied to provide an equipment-grounding conductor to bond the photovoltaic modules and support structure to earth per the solar module manufacturer’s requirements. Any other system component requiring lightning surge protection must also be included per manufacturer(s) requirements. The equipment grounding conductor must be a minimum of 20 meters with a minimum of 1.5-metre grounding rod with connectors. | [Please refer to a document or describe the grounding system for the solar system.]
4.2.12 | User maintenance and technician installation tool kits
| All sites are to be supplied with one complete user maintenance kit consisting of all necessary operations and maintenance tools as proposed by the legal manufacturer. Specialised tools and materials required for installations are to be clearly identified to prospective buyers and offered as an option by legal manufacturer. The type and quantity of specialised tools must be determined and agreed upon in advance of order placement. | [Please state if any specialised tools are required.]
| Solar direct drive installation sites must be sufficiently shade free, as determined by a solar site analysis that accounts for the daily average site shading on no less than a monthly basis. Specialised tools for shading analysis must be offered to prospective buyers. | [Please state which specialised tools for shading analysis are offered.]
Note: sites with multiple solar-powered appliances can be supplied with a single maintenance kit suitable for the type of solar power system provided.
4.2.13 | Electromagnetic compatibility
| The legal manufacturer must certify compliance with the requirements of the latest edition of IEC 61000-6-1 and IEC 61000-6-3. | [Please refer to a test report proving EMC compliance.]
4.3 | Environmental requirements
4.3.1 | Ambient temperature range during transport and storage
| Components must be able to withstand ambient temperatures of -30°C to +70°C when in transit. If individual components (e.g. batteries) are not able to withstand transport temperature conditions then special transport arrangements may have to be made. In advance of purchase, the legal manufacturer must notify procurement agency of the need and requirements for special transport arrangements. | [Please refer to documents confirming ambient temperature range during transport and storage.]
4.3.2 | Ambient temperature range during use
| Photovoltaic panels may be exposed to temperatures from -40°C to +90°C when mounted in their final positions. Battery sets and battery charge regulators may be exposed to temperatures from -10°C to +43°C after installation and commissioning. | [Please refer to documents confirming ambient temperature range during use.]
4.3.3 | Ambient humidity range during transport and storage
| The ambient humidity range during transport and storage is 5% to 95%, non-condensing | [Please refer to documents confirming ambient humidity range during transport and storage.]
4.4 | Physical characteristics
4.4.1 | Overall dimensions
| System components must be able to fit through an 830 mm-wide door opening (with the door leaf removed if necessary). | [Please state the external dimensions.]
4.4.2 | Weight
| System components must be capable of being safely manhandled into their final positions. Mechanical lifting equipment will typically not be available at the installation sites. It is recommended that individual component packages should be designed so that they can be lifted in such a way that no single worker is required to carry more than 25 kg whilst working on their own, or in a group. | [Please state the weight of the system components and how they may safely lifted or maneuvered into position on site.]
4.5 | Interface requirements
4.5.1 | Appliances
| In order to complete the installation, the qualified supplier must supply all the solar system components needed to connect to the specified appliance. Appliances for Type 1 systems must comply with PQS specification E003/RF04. Appliances for Type 2 systems must comply with specification E003/FZ03, E003/RF05 or E003/RF06. | [Please state which appliances are compatible with the solar system.]
4.6 | Human factors
4.6.1 | Generally
| The solar system components must be easy to assemble and commission by installation technicians who have received basic training, and using normally available hand tools. The completed installation must be useable by the widest practicable range of health workers, regardless of age, gender, size or minor disability, including colour blind users and long-sighted people without glasses, in accordance with the general principles laid out in ISO 20282-1: 2006. | [Please state how the solar system complies with this clause.]
4.6.2 | Safe access
| Photovoltaic arrays must be mounted in positions that allow safe access for installation, maintenance and cleaning, using appropriate fixed access equipment as necessary. Avoid array mounting locations which require access across fragile roof structures and finishes. In locations where safe access is not possible provide specialized cleaning tools (e.g. extended handle glass cleaning tools) to allow cleaning from a safe and secure position. | [Please describe photovoltaic array mounting options.]
4.7 | Materials
4.7.1 | Restricted materials
| The product and its constituent components, must not contain lead (except in batteries), mercury, cadmium, hexavalent chromium, polybrominated biphenyls (PBB) or polybrominated biphenyl ethers (PBDE). | [Please refer to document showing compliance with this clause.]
4.8 | Warranty
| Any component which fails due to defective design, materials or workmanship must be covered by a replacement warranty. The minimum periods for the warranty must be as follows: • 25 years for the solar module power output. • Five years for the battery, with a full like-for-like replacement warranty for the first 12 months and pro-rated financial compensation, based on the purchase cost of the battery set, for the remaining period. • Two years for all other components. | [Please refer to a statement of the warranty in the user manual and how a purchaser may make a claim against this warranty.]
4.9 | Servicing provision
4.9.1 | Maintainability
| Installations are to be substantially maintenance-free other than PV array cleaning, routine battery servicing (flooded batteries) and periodic battery replacement. Legal manufacturers must publish a recommended battery replacement plan. | [Please state how the system has been designed for minimal maintenance. For Type 1 systems, describe battery maintenance and replacement, including frequency.]
4.9.2 | Essential spare parts and user maintenance tools/supplies:
| Based on product design and requirements, the type and quantity of spare parts, basic installation tools/supplies, user and technician maintenance manuals (see Section 4.11 Instructions), must be determined and agreed upon in advance of order placement. As a minimum, each solar power system must be supplied with 10 spare fuses of all fuse size and types used in the system. The spares fuses are to be attached within the appliance or within accessible power system enclosures. | [Please describe (including diagrams or photos) the maintenance kit in accordance with this clause. Show the location of the spare fuses.]
| Legal manufacturers must publish a list of spare parts recommended for five years of operation for purchases of 10 and 50 power systems. Legal manufacturers must ensure supply of spare parts for a minimum of five years from the time of cessation of the last production of equipment. Spare parts are to be provided in kit form for storage in appropriate quantities at National or Sub-national level in the purchasing country, as agreed with the purchasing agency. | [Please declare compliance with this clause and how spare parts may be obtained e.g. via local agents.]
| Legal manufacturers must supply sufficient spare parts for the first five years of operation. Spare parts are to be provided in kit form for storage in appropriate quantities at central or regional level in the purchasing country, as agreed with the purchasing agency. | [Please declare compliance with this clause and how spare parts may be obtained e.g. via local agents.]
4.1 | Disposal and recycling
| The legal manufacturer is to provide information to the buyer on the toxic and hazardous materials contained within the components and suggestions for resource recovery/recycling and/or environmentally safe disposal. For the European Union WEEE compliance in accordance with European Union Directive2002/96/EC is mandatory. | [Please refer to a statement hazardous material in the user manual and the provision for disposal/recycling.]
4.11 | Instructions
Each solar power system must include a separate user manual and installation technician manual in Arabic, English, French, Mandarin Chinese, Russian and Spanish. An English version of all instructions and manuals are required to be supplied at time of laboratory testing. Instructions to include easy to understand visuals whenever possible to avoid reliance on text.
The user manual must include the following information:
| • Health and safety guidance, | [Please refer to a particular section in the user manual.]
| • Basic operations description, | [Please refer to a particular section in the user manual.]
| • Cool-down time instructions to assure both acceptable temperatures and full autonomy is achieved, and | [Please refer to a particular section in the user manual.]
| • Routine maintenance tasks (e.g. daily, weekly, and monthly). | [Please refer to a particular section in the user manual.]
The technician installation manual must include the following information:
| • Health and safety guidance, | [Please refer to a particular section in the installation manual.]
| • Compatible types of appliances, | [Please refer to a particular section in the installation manual.]
| • Detailed operations description, | [Please refer to a particular section in the installation manual.]
| • Cool-down time instructions to assure both acceptable temperatures and full autonomy is achieved, | [Please refer to a particular section in the installation manual.]
| • Correct handling to avoid component damage and for the safety of handling persons | [Please refer to a particular section in the installation manual.]
| • Solar array site shading assessment tools and procedures, | [Please refer to a particular section in the installation manual.]
| • Electrical wiring diagram, | [Please refer to a particular section in the installation manual.]
| • Installation procedures, | [Please refer to a particular section in the installation manual.]
| • Commissioning procedure, | [Please refer to a particular section in the installation manual.]
| • Technical maintenance tasks (e.g. daily, weekly and monthly), | [Please refer to a particular section in the installation manual.]
| • Periodic preventative maintenance procedures, | [Please refer to a particular section in the installation manual.]
| • Diagnostic and repair procedures, | [Please refer to a particular section in the installation manual.]
| • Itemised list of spare parts including part numbers, | [Please refer to a particular section in the installation manual.]
| • End-of-life resource recovery and recycling procedures, and | [Please refer to a particular section in the installation manual.]
| • User training guidance. | [Please refer to a particular section in the installation manual.]
| Printed user and routine maintenance instructions specifically directed at the health centre or store staff must be pictorial. All key information must be summarised on a single sheet that can be fixed onto the appliance cabinet at time of installation. The sheet should be sufficiently durable to last the life of the appliance and must be in a locally-understood language. | [Please refer to photo which shows compliance with this clause.]
| Installation, repair and servicing instructions must be supplied in printed format and optionally on DVD and/or online, to instruct the installation technician in installation standards and practices specific to the appliance and its power system. In addition, supporting video material supplied on DVD and/or online can be supplied to assist the instructor when delivering on-site user training. | [Please state all sources for installation, repair and servicing instructions.]
4.12 | Training
| Specific user training must be provided to the responsible, on-site appliance user. User training conducted by installers at time of installation is recommended. The legal manufacturer is required to provide specific user training instructions to enable an instructor (e.g. installation technician, health worker supervisor) to provide product specific user instructions to the responsible, on-site appliance user. | [Please describe the different options for training which might be available.]
| Optional training courses are to be offered by the qualified supplier or a designated representative. | [Please describe the different options for training which might be available.]
4.13 | Verification
| In accordance with PQS Verification Protocol E003/PV01-VP1 Legal manufacturer’s will be required to present evidence of conformity. | [Please provide test reports and other documents verifying conformity.]
5. | Packaging
| Materials used for packaging components are to be free of ozone-depleting compounds as defined in the Montreal Protocol. The general specification of shipping containers will be subject to agreement with the individual procurement agencies. | [Please describe the packaging and refer to the user manual section/page number on how this packaging may be disposed / recycled.]