## Sheet1
E003 Refrigerator or combined refrigerator and water-pack freezer: intermittent mains powered, compression cycle submission template (V3.0 RFR 03 DEC 2024)
| Assessment of appliance: | [Brand & Model here] | For dossier review on [TBC]
| Tested in accordance with PQS verification protocol: | E003/RF03-VP.5 (2 DEC 2024)
| To verify compliance with PQS specification: | E003/RF03.7 (2 DEC 2024)
ASSESSOR’S SUMMARY:
Instructions in [square brackets] are for your guidance. Please delete these instructions and replace with evidence or a precise reference (including page and section numbers) showing compliance with that particular PQS specification requirement. In general a "pass" or "complies" statement alone is not acceptable.
One document may provide evidence for several clause requirements. Please carefully reference the sub-sections of any such document so that it is clear that the clause requirement has been met. This will help speed the pre-qualification process.
Greyed-out cells do not require a response.
NOTE: If it is not clear or doubtful whether a requirement has been met, then this is a non-compliance and pre-qualification might be delayed.
|  |  | REVIEWER’S SUMMARY: |  | REVIEWER’S SUMMARY: |  | REVIEWER’S SUMMARY:
Section 7. | Product dossier | Applicant's submission | Reviewer's remarks [Name & Date] | Applicants response [Name & Date] | Reviewer's remarks [Name & Date] | Applicants response [Name & Date] | Reviewer's remarks [Name & Date]
The legal manufacturer or reseller is to provide WHO with a pre-qualification dossier containing the following:
| Dossier examination fee in US dollars. | [Please contact PQS.]
| General information about the legal manufacturer, including name and address. | [Please state or refer to a document.]
| Unique identification reference for the product type. | [Please state.]
| Brand name of the product. | [Please state.]
| Full specifications of the product being offered, covering all the requirements set out in this document, including details of product marking and traceability. | [Please refer to a document.]
| A comprehensive set of photographs including a three-quarter view of the appliance with the door open. Take additional photographs showing all external surfaces of the unit, the interior layout, the compressor or cooling system and a close-up of the thermometer, indicator light(s), the control(s), control panel and any special features. | [Please refer to a document.]
| Certified photocopies of all type-approvals obtained for the appliance and/or its components, including CE marking and the like. | [Please refer to a document.]
| Certified photocopies of the legal manufacturer’s ISO 9001 quality system certification. | [Please refer to a document.]
| Where relevant, certified photocopies of the legal manufacturer’s ISO14001 certification, EMAS registration or registration with an equivalent environmental audit scheme. Conformity with an environmental audit scheme is not mandatory however, preference will be given to manufacturers who are able to demonstrate compliance with good environmental practice. | [Please refer to a document.]
| Laboratory test report(s) proving conformity with PQS specifications. | [Please refer to a document.]
| Indicative cost of the product per unit, per 10 units and per 100 units, EXW (Incoterms 2010). | [Please state or refer to a document.]
Section 4. | Requirements
4.1 | General
| Mains-powered compression-cycle refrigerators, with or without water-pack freezing compartment, are used primarily in areas with an intermittent electricity supply (e.g. eight to 20 hours of reliable electricity per typical day, less than eight hours of reliable electricity per typical day). | [Please state minimum number hours of reliable electricity required per day.]
4.2 | Performance
4.2.1 | Operating temperature range
| The operating temperature range is indicated on the temperature zone rating sticker attached to appliance front (see Annex 1). All appliances must be able to maintain the acceptable temperature range in the vaccine storage compartment when operating in the hot zone (+43°C) and at a minimum rated ambient temperature of +10°C or lower. | [Please state the temperature zone, MRA and show a photo of this label on the appliance.]
4.2.2 | Refrigeration system
Compression-cycle unit operating on alternating current electricity.
4.2.3 | Design of vaccine storage compartment
| The vaccine storage compartment must be designed so that no part, which is outside the acceptable temperature range, can be used to store vaccines either by inadvertent or deliberate misuse. As per the freeze-protection classification features, appliances complying with this requirement without demanding any intervention from the user will be published as Grade A. Others will be published as Grade B or Grade C depending on the level of interventions required. The vaccine storage compartment must provide some means (such as baskets) to enforce physical separation between the vaccines and any surfaces that potentially have condensate on them (such as the floor, ceiling and/or walls of the compartment). | [Please state the freeze-protection grade.] [If Grade A, are baskets or similar available as a storage convenience option? [If not Grade A, what provision is made to ensure vaccines do not come into contact with cold surfaces?]
4.2.4 | Vaccine freeze protection classification
| The vaccine freeze protection classification is indicated on the freeze protection classification sticker attached to appliance front (see Annex 4). The amount of user intervention required to ensure that the vaccines will not be exposed to freezing temperatures when the appliance is used within its nominated temperature range and minimum rated ambient temperature will be classified and reported as follows: | [Please refer to a photo of the freeze protection classification sticker.]
| Grade A, user-independent freeze protection (UIFP): when there is no user intervention required to ensure that the vaccines will not be exposed to freezing temperatures whatever the position of the vaccine in the vaccine storage compartment. | [Please state the freeze-protection grade and refer to the test evidence that demonstrates this grade.]
Grade B, user-dependent freeze protection (UDFP): when the user must comply with a procedure provided by the legal manufacturer and requiring one level of user intervention (e.g., the requirement to use baskets to avoid vaccine freezing temperatures constitute one level of user- intervention).
Grade C, user-dependent freeze protection (UDFP): when the user must comply with a procedure provided by the legal manufacturer requiring two or more levels of user interventions (e.g., a refrigerator that not only requires the use of baskets but also requires use of removable thermal barriers constitutes two levels of user intervention).
4.2.5 | Water-pack freezing capacity (combined units only)
| In combined units with freezer compartment, a minimum of 1.6 kg and not less than 2.4 kg per 50 litres of gross freezer volume must be frozen per 24 hours. Additionally, a new methodology for measurement of ice production is presented in Annex 3. This will be reflected in the verification protocol for freezers. | [Please state the freezing capacity and refer to the test evidence that demonstrates this.]
4.2.6 | Water-pack storage compartment capacity (combined units only)
| The freezer compartment to hold a minimum of 3.2 kg of fully frozen water-packs and at least twice the daily water-pack freezing capacity determined by E003/RF03-VP.5. The water-packs must comply with E005/IP01. | [Please state the water-pack storage capacity and refer to the test evidence that demonstrates this.]
4.2.7 | Temperature control
| Refrigerator compartment: The zone within the vaccine compartment that is designated for vaccine storage must remain within the acceptable temperature range during any continuous ambient temperature test(s) or day/night cycling temperature test(s). Combined units must achieve this performance with or without water-packs in the water-pack compartment. | [Please refer to the test evidence that demonstrates correct internal temperatures.]
| Water-pack freezing compartment: The water-pack freezing compartment (if present) must remain below -3°C under the same ambient conditions and the minimum weight of water-packs described in clause 4.2.5 must remain fully frozen at the end of the power-off cycle. | [Please refer to the test evidence that demonstrates correct water-pack temperatures.]
4.2.8 | Thermostat
| The thermostat must be set to prevent freezing in any part of the vaccine storage compartment. The thermostat must be effective throughout the ambient operating temperature range (down to the minimum rated ambient temperature – see clause 4.2.13). It must be designed so that it cannot be adjusted by the user. A means for adjustment by a technician is acceptable provided the device is protected from user interference (e.g. by location within the appliance cabinet). Alternatively, programmable thermostats may be password-protected. Bulb and capillary tube thermostats are not acceptable. | [Please refer to photo(s) or diagram(s) that demonstrates compliance to this clause.]
4.2.9 | Temperature monitoring and thermometer
| The appliance shall have a data logger integrated within the appliance with connections for data transfer and power known as the Machine-to-Machine interface (M2M) that supports the functionality of external devices, including external EMDs. The data logger and M2M support data recording and local access to data. This data logger must conform to the requirements of E006/DL01 Data logger and machine-to-machine interface for Equipment Monitoring Systems. If the appliance is a combined refrigerator and vaccine freezer, both compartments shall be monitored by the data logger. | [Please provide evidence a E006/DL01 datalogger is integrated into the appliance. Please confirm the datalogger is approved for use with that appliance.]
| The appliance shall have a display to show internal cabinet temperature and positioned so that it can be easily read by a health worker without opening the appliance. This electronic thermometer display may be configured as part of an electronic monitoring system (EMS) or as part of a thermometer conforming to E006/TH06. The internal sensor shall be located so that the temperature reading gives the coldest temperature in the vaccine refrigerator compartment. | [Please state whether the appliance has a display per E006/TH06 or integrated EMD (E006/EM01). Please state which configuration per E003/RF05.7 Clause 4.2.9 Table 1 the appliance complies with.]
| The EMS shall last throughout the 10-year expected lifetime of the refrigerator, so that the data logger and monitor is always active on-site. If any data logger components do not have a 10-year lifetime (e.g. batteries), the refrigerator manufacturer must also provide the consignee with replacement components to ensure the EMS will operate for at least 10-years. Use, maintenance and any replacement of data logger components e.g. rechargeable battery, shall be stated in the user instructions. | [Please confirm the EMS has an expected lifetime of 10 years. Please state the epected lifetime of the EMS batteries.]
| A 30-day temperature recorder (30DTR) may be provided in lieu of an EMD and shall conform to E006/TR06. If a 30DTR is provided, the refrigerator manufacturer must also provide the consignee with replacement temperature monitoring devices throughout the 10-year lifetime. The expected lifetime, along with the warranty, of any monitoring device provided shall be stated in the user manual. | [Please state if a 30DTR is being used instead of an EMD.]
| If a manufacturer ships the E-EMD or 30DTR separately, they should arrive at the same destination as the refrigerator shipment, on the same date, and addressed to the same consignee. This is necessary to allow for smooth in-country receipt and assembly of the refrigerators and temperature monitoring devices. Separate shipments of E-EMD or 30DTR must be tracked and coordinated by the refrigerator manufacturer. | [Please state if an E-EMD or 30DTR is being used, and if they are being shipped with the appliance.]
4.2.10 | Humidity control (Optional)
| Refrigerator compartment: The refrigerator compartment must have humidity control. The zone within the vaccine compartment that is designated for vaccine storage must remain within acceptable compartment humidity levels during humidity control testing as specified in the verification protocol. Combined units must achieve this performance with or without water-packs in the water-pack compartment. | [Please state if the appliance has humidity control; and if so, the level of humidity control as shown by testing in accordance with E003/RF05-VP.5 Test 3.]
Water-pack freezing compartment: Not applicable
4.2.11 | Indicator light
| A minimum of one green LED indicator light is required to be located on the front or top of the appliance to alert users that the cooling system is actively operating. A constant green LED light is required to indicate that the compressor or cooling system is active and the light must go off when the compressor or cooling system is off. | [Please refer to photo(s) or diagram(s) that demonstrates compliance to this clause.]
| Optionally, additional indicator lights may be added to indicate other operating conditions including temperature and faults. | [Please refer to photo(s) or diagram(s) that demonstrates compliance to this clause.]
4.2.12 | Holdover times
Holdover times of refrigerators will henceforth be categorized as follows:
| • Short: Holdover 20 hours ≤ 48 hours. | [Please refer to the test evidence that demonstrates the stated holdover time.]
• Medium: Holdover 48 hours ≤120 hours.
• Long: Holdover 120+ hours.
4.2.13 | Minimum rated ambient temperature
| All models must be able to operate at a continuous minimum ambient temperature of +10.0°C or lower while maintaining the acceptable temperature range. The maximum performance rating is achieved if the vaccine load remains within the acceptable temperature range at -10°C. A freeze-prevention circuit may be required to protect against freezing at low ambient temperatures. | [Please refer to the test evidence that demonstrates correct internal temperatures at the stated minimum rated ambient temperature.]
4.2.14 | Power system requirements and consumption
| Direct supply of mains electricity. Options for 220-240 volt 50/60 Hz and 100-127 volt 50/60 Hz are to be offered. Performance must be identical for all options, regardless of the nominal voltage and frequency rating of the appliance. | [Please state which voltage and frequency options are available.]
| At 22% below manufacturers stated voltage, 10 out of 10 cold starts and 10 out of 10 hot starts must all be successful. | [Please refer to the test evidence that demonstrates compliance with this clause.]
| No standard set, however electricity consumption will be reported. | [Please state typical electricity consumptions and the test references that confirm this.]
4.2.15 | Condensation management and defrosting
The environmental conditions within the vaccine storage compartment must be designed so that vaccine primary containers and vaccine cartons are not exposed to levels of humidity which may cause damage to cartons or primary container labels or create a risk of mould growth.
| To alleviate humidity damage, legal manufacturers are to include refrigerator design features and/or provide containers for vaccine storage. Condensate and defrost drainage must be provided in all refrigerator and freezer compartments. If used, the defrost switch (or switches) must be accessible to the user without tools but must be protected from accidental changes in position. | [Please state design features or provision which alleviate humidity in the cabinet.]
4.2.16 | Lock
| The door or lid must be fitted with a lock. Two keys are to be supplied with every unit. | [Please confirm compliance to this clause.]
4.2.17 | Corrosion resistance
| The legal manufacturer must certify compliance that internal and external cabinet, lid and frame are protected against corrosion as appropriate to EN ISO 6270-1 / ASTM D2247 / EN 13523-26, EN ISO 6270-2 / EN 13523-25, ISO 6272 / EN 13523-5 and ISO 2409: 2013. | [Please refer to a document that shows test evidence for compliance to this clause. i.e. precisely which components have which proven level of corrosion protection.]
4.2.18 | Electrical safety rating
| The legal manufacturer must certify compliance with IEC 60335-1, IEC 60335-2-24 and IEC 60364-1. | [Please refer to a document that shows evidence for electrical safety compliance.]
4.2.19 | Markings and labelling
| Compressors must be marked with the blue identifying symbol shown in Annex 2. In addition, the cabinet must be permanently marked, near the compressor position, with the chemical name of the refrigerant or with the refrigerant number, formula or proportion (for blended refrigerants). | [Please refer to a photo of the blue label on the compressor.]
| Appliances operating on R600a must be marked with the warning symbols shown in Annex 2. | [Please refer to a photo of the warning label on the appliance. Note this is required for all flammable refrigerants.]
| All appliances must label hazardous materials and include a Safety Data Sheet. Label and Safety Data Sheet must comply with the Globally Harmonized System for the Classification and Labelling of Chemicals GHS Rev.5. | [Please refer to a copy of the Safety Data Sheet.]
| The appliance must carry the following additional information fixed to the front of the cabinet: manufacturer and model number (unless already located on the front of the unit), serial number, date of manufacture, PQS identification number, applicable service phone number and website URL. This label to remain readable for the expected life of the appliance. | [Please refer to a photo of this label.]
| All appliances require an asset identification label (bar code) as specified in WHO/PQS/E003/BC01.1: Global asset identification. Effective date June 30, 2020. | [Please refer to a photo of the bar code.]
4.2.20 | Vaccine storage advice
All units must carry a factory-fitted non-removable label, designed to last the lifetime of the appliance and carrying the following information:
| • Vaccine refrigerators: Vaccine storage instructions and the appropriate temperature zone symbol as Annex 1. | [Please refer to a copy of these instructions.]
| • Combined units: Vaccine storage instructions, water-pack freezing instructions and the appropriate temperature zone symbol as Annex 1. | [Please refer to a copy of these instructions.]
| The instructions should be fixed to the lid of chest refrigerators and near the top of the door on upright refrigerators. Instructions should be in one of the languages specified in clause 4.11, as indicated by the purchaser at the time of ordering. If the appliance is graded other than “A” and removable baskets are required to avoid freezing temperatures, fix a multi-lingual warning within the refrigerator instructing users to Store vaccine in baskets only or other appropriate instruction. | [Please state ambient temperature range during transport and storage.]
4.2.21 | EMC compliance
| The legal manufacturer must certify EMC compliance with the latest edition of IEC 61000-6-1 and IEC 61000-6-3. | [Please refer to a document that shows evidence for EMC compliance.]
4.3 | Environmental requirements
4.3.1 | Ambient temperature range during transport and storage
| Ambient temperature range during transport and storage is -30°C to +70°C when the product is inactivated. | [Please state ambient temperature range during transport and storage.]
4.3.2 | Ambient humidity range during transport, storage and use:
| The ambient humidity range during transport, storage and use is 5% to 95% RH, non-condensing. | [Please state ambient humidity range during transport, storage and use.]
4.4 | Physical characteristics
4.4.1 | Overall dimensions
| To allow for manoeuvring around corners and through corridors and doorways, the minimum dimension of the product (either length, width or height) should not exceed 710 mm; exceptionally a minimum dimension up to 830 mm can be accepted, but this may restrict the number of sites where the appliance can be installed. The maximum dimension should not exceed 1700 mm and the maximum diagonal (corner to corner) dimension must not exceed 1850 mm. All dimensions must be clearly stated on the PQS information sheet. | [Please state the external dimensions.]
4.4.2 | Weight
| Mechanical lifting equipment will often not be available at the installation sites. It is recommended that the refrigerator and any associated components should be designed for lifting in such a way that no single worker is required to carry more than 25 kg while working on their own, or in a group. Mechanical lifting equipment at the installation sites is necessary when component weight exceed practical weight limits of the workforce. | [Please state how the appliance may be easily lifted or manoeuvred into position on site.]
4.5 | Interface requirements
4.5.1 | Electrical components
Every appliance must be provided with either an integrated or a standalone voltage stabilizer. The voltage stabilizer must be certified by WHO PQS as complying with PQS specification E007/VS01.2.
| All electrical components must be compatible with voltage stabilizers that use tap-changing technology. If the product contains components that are incompatible with this type of voltage stabilizer it must be supplied with an appropriate device of equivalent performance to those voltage stabilizers that are currently prequalified in PQS section E007. A warning must be affixed to the unit stating the type(s) of voltage stabilizer that may be used, and the user’s manual and spare parts list must clearly record this warning. | [Please state which PQS approved voltage stabilizer will be supplied with the appliance.]
4.5.2 | Power lead
| The appliance must be supplied with a power lead with a sealed-on plug compatible with the electrical socket standard in the country where the equipment must be installed. The power lead must be at least 1.5 metre and not more than 2.0 metre in length. | [Please declare compliance with this clause.]
4.6 | Human factors
4.6.1 | Generally
| The product must be usable by the widest practicable range of active health workers, regardless of age, gender, size or minor disability, including colour blind users and long-sighted people without glasses, in accordance with the general principles laid out in ISO 20282-1:2006. | [Please state how your appliance complies with this clause.]
4.6.2 | Control panel and thermometer
| Controls, thermometer and other visual displays may be positioned on the front of the unit preferably as close to eye level as possible. Alternatively, they may be mounted on top of the unit at a height not exceeding 1.3 metres. If a low-level position is essential, the display must be aligned so that it can easily be read without the user having to squat or kneel down. The on-off and/or defrost switch, if present, should be recessed or otherwise protected so that it is not possible inadvertently to activate it. | [Please refer to a photo that shows compliance with this clause.]
4.6.3 | PQS stickers
In addition to the PQS temperature zone sticker the device should carry the following additional information:
| • Manufacturer and model number (unless already located on the front of the unit), serial number, date of manufacture, PQS identification number, applicable service phone number and website URL fixed to the front of the cabinet. | [Please refer to a copy of this label and also show its position on the appliance.]
| • An operations and maintenance pictogram fixed to the lid or near the top front of chest refrigerators and near the top of the door on upright refrigerators. | [Please refer to a copy of this label and also show its position on the appliance.]
PQS stickers should remain readable for the expected age of the equipment.
4.7 | Materials and construction
4.7.1 | Refrigerant
| Hydrocarbon (HC) refrigerants such as R600a or other refrigerants with a global warming potential (GWP) ≤ 11 and zero ozone-depletion potential (ODP) are required. The suitability of alternative refrigerant gases will continue to be assessed and preference will be given to appliances that use gases with low GWP. | [Please state type of refrigerant(s) and their GWP.]
4.7.2 | Thermal insulation foaming agents
| Any gas complying with the limitations and deadlines set by the Montreal/Kigali Protocols on the elimination of ozone-depleting chemicals. | [Please state type of foaming agent, it’s compliance with the Montreal Protocol and its GWP.]
4.7.3 | Other restricted materials
| The product and its constituent components, including batteries, must not contain lead, mercury, cadmium, hexavalent chromium, polybrominated biphenyls (PBB) or polybrominated biphenyl ethers (PBDE). | [Please refer to a declaration of compliance with this clause.]
4.7.4 | PCM
| Integrated thermal buffer materials may be used to prevent freezing temperatures from propagating to the vaccine storage compartment or for other thermal purposes. The buffer material may be PCM-based but if so, must comply with WHO/PQS/ E005/PCMC0.1– PCM specification for phase-change material containers. | [Please state whether your appliance uses a thermal buffer an state the compliance of PCMs if any.]
4.8 | Warranty
| The appliance to be covered by a three-year replacement warranty from receipt of appliance in the event of any component failure arising from defective design, materials or workmanship. | [Please refer to a declaration of compliance with this clause.]
The warranty period begins on the date of shipment from the legal manufacturer.
4.9 | Servicing provision
| The product must be designed, and components selected, with the aim of achieving a zero-repair life of not less than 10 years. | [Please state how the appliance is designed to last at least 10 years.]
4.9.1 | Essential spare parts and user maintenance tools/supplies
| All sites are to be supplied with one complete user maintenance kit consisting of all necessary operations and maintenance tools as proposed by the legal manufacturer. Based on product design and requirements the type and quantity of spare parts, basic installation tools/supplies, user and technician maintenance manuals (see clause 4.11), must be determined and agreed upon in advance of order placement. As a minimum, each appliance to be supplied with 10 spare fuses of all fuse size and type used in the appliance. The spares fuses are to be attached within or on the appliance. | [Please describe (including diagrams or photos) the maintenance kit in accordance with this clause. Show the location of the spare fuses.]
| Legal manufacturers are to publish a list of spare parts recommended for purchases of 10 and 50 appliances and power systems. The legal manufacturer must ensure supply of spare parts for a minimum of five years from the time of cessation of the last production of equipment. Spare parts are to be provided in kit form for storage in appropriate quantities at National or Sub-national level in the purchasing country, as agreed with the purchasing agency. | [Please declare compliance with this clause and how spare parts may be obtained e.g. via local agents.]
4.10 | Disposal and recycling
| The manufacturer must provide information to the buyer on the hazardous materials contained within the system and suggestions for resource recovery/recycling and/or environmentally safe disposal. For the European Union WEEE compliance in accordance with European Union Directive 2002/96/EC is mandatory. | [Please refer to a document giving details of any hazardous materials.] [Please refer to information explaining resource recovery/recycling and/or environmentally safe disposal. e.g. a section in the user manual.]
4.11 | Instructions
Printed user, installation and routine maintenance instructions specifically directed at the health centre or store staff must be pictorial. All key information should be summarized on a single sheet fixed onto the appliance cabinet sufficiently durable to last the life of the product and must be available in Arabic, Mandarin Chinese, English, French, Russian and Spanish. In addition, supporting video material supplied on DVD and/or online can be supplied to assist the instructor when delivering on-site user training.
| The manufacturer can maintain a core set of manuals in one language, but must be able to provide translations in any of the aforementioned UN languages if requested. The manufacturer can also store the manuals on their websites, with links that are shareable when requested. | [Please state all sources for installation, repair and servicing instructions.]
The user, installation and maintenance documents must address five aspects:
1. Introductory information. Must include:
| - Title page with image of unit, supplier name, supplier model number, PQS code and version number | [Please refer to a particular section in the installation manual.]
| - Table of contents | [Please refer to a particular section in the installation manual.]
| - General information on unit, its functionality and intended use | [Please refer to a particular section in the installation manual.]
| - Relevant warnings related to transportation, any corrosive or toxic substances in the construction of the appliance, power source or disposal. | [Please refer to a particular section in the installation manual.]
2. Model specifications and details. Must include:
| - Detailed technical specifications, including wiring diagram | [Please refer to a particular section in the installation manual.]
| - Parts and equipment list | [Please refer to a particular section in the installation manual.]
| - Detailed technical specifications, including wiring diagram | [Please refer to a particular section in the installation manual.]
| - Safety procedures, including warranty information and supplier contact information | [Please refer to a particular section in the installation manual.]
| - Directions for safe transportation. | [Please refer to a particular section in the installation manual.]
3. Installation and operation. Must include:
| - Detailed installation procedure, including installation checklist | [Please refer to a particular section in the installation manual.]
| - Cool-down time to achieve both acceptable temperature range and full holdover time. | [Please refer to a particular section in the installation manual.]
| - Detailed operational procedures covering both vaccine storage, as well as ice-pack / cold-pack preparation | [Please refer to a particular section in the installation manual.]
| - Disposal guidelines. | [Please refer to a particular section in the installation manual.]
4. Maintenance. Must include:
| - Detailed guidance on preventative maintenance, including checklists and standard operating procedures (SOPs) | [Please refer to a particular section in the installation manual.]
| - Troubleshooting guide for corrective maintenance, including table detailing common issues and step-by-step remedial actions | [Please refer to a particular section in the installation manual.]
| - Typical replacement cycle for spare parts | [Please refer to a particular section in the installation manual.]
| - Recommended replacement dates for temperature monitoring devices (see clause 4.2.9). | [Please refer to a particular section in the installation manual.]
5. Format and usability. Additional guidance:
| - Include clear graphics to illustrate tasks, with multiple view-points (e.g. top, side) and clear labelling | [Please refer to a particular section in the installation manual.]
| - Publish in English, with translations readily available in all UN languages (Arabic, English, French, Mandarin, Russian, and Spanish). Translations to other languages specific to certain countries are to be provided if requested by the buyer | [Please refer to a particular section in the installation manual.]
| - Be specific to a given model and avoid covering multiple devices in same document | [Please refer to a particular section in the installation manual.]
| - Have a clear and consistent structure that covers installation, operation and maintenance and repair | [Please refer to a particular section in the installation manual.]
| - Be accessible and downloadable from a central repository. | [Please refer to a particular section in the installation manual.]
| Installation, repair and servicing instructions must be supplied in printed format, and optionally on DVD and/or online to instruct the installation teams in installation standards and practices specific to the product and its power system. | [Please state all sources for installation, repair and servicing instructions.]
4.12 | Training
| Not mandatory. | [Please state if there is a training program.]
4.13 | Verification
| In accordance with PQS Verification Protocol E003/RF03 -VP.5 | [Please refer to the laboratory test report.]