## Sheet1
E003 Submission to PQS for approval: SDD Refrigerator or combined refrigerator and water-pack freezer submission template (V3.0 RFR 03 DEC 2024)
| Assessment of appliance: | [Brand & Model here] | For dossier review on [TBC]
| Tested in accordance with PQS verification protocol: | E003/RF05-VP.6 (2 December 2024)
| To verify compliance with PQS specification: | E003/RF05.8 (2 December 2024)
ASSESSOR’S SUMMARY:
Instructions in [square brackets] are for your guidance. Please delete these instructions and replace with evidence or a precise reference (including page and section numbers) showing compliance with that particular PQS specification requirement. In general a "pass" or "complies" statement alone is not acceptable.
One document may provide evidence for several clause requirements. Please carefully reference the sub-sections of any such document so that it is clear that the clause requirement has been met. This will help speed the pre-qualification process.
Greyed-out cells do not require a response.
NOTE: If it is not clear or doubtful whether a requirement has been met, then this is a non-compliance and pre-qualification might be delayed.
|  |  | REVIEWER’S SUMMARY: |  | REVIEWER’S SUMMARY: |  | REVIEWER’S SUMMARY:
Section 7. | Product dossier | Applicant's submission | Reviewer's remarks [Name & Date] | Applicants response [Name & Date] | Reviewer's remarks [Name & Date] | Applicants response [Name & Date] | Reviewer's remarks [Name & Date]
The legal manufacturer or reseller is to provide WHO with a pre-qualification dossier containing the following:
| Dossier examination fee in US dollars. | [Please contact PQS.]
| General information about the legal manufacturer, including name and address. | [Please state or refer to a document.]
| Unique identification reference for the product type. | [Please state.]
| Brand name of the product. | [Please state.]
| Full specifications of the product being offered, covering all the requirements set out in this document, including details of product marking and traceability. | [Please refer to a document.]
| Full details of the recommended compatible Type 2 solar power system (see specification E003/PV01). | [Please refer to a document.]
| A comprehensive set of photographs including a three-quarter view of the appliance with the door open. Take additional photographs showing all external surfaces of the unit, the interior layout, the compressor or cooling system and a close-up of the thermometer, indicator light(s), the control(s), control panel and any special features. | [Please refer to a document.]
| Certified photocopies of all type-approvals obtained for the appliance and/or its components, including CE marking and the like. | [Please refer to a document & section.]
| Certified photocopies of the legal manufacturer’s ISO 9001 quality system certification. | [Please refer to a document.]
| Where relevant, certified photocopies of the legal manufacturer’s ISO 14001 certification, EMAS registration or registration with an equivalent environmental audit scheme. Conformity with an environmental audit scheme is not mandatory; however, preference will be given to manufacturers who are able to demonstrate compliance with good environmental practice. | [Please refer to a document.]
| Where available, laboratory test report(s) proving conformity with the PQS equipment performance specifications. | [Please refer to a document.]
| Indicative cost of the appliance per unit, per 10 units and per 100 units, EXW (Incoterms 2010) including appliance and solar power system. | [Please state or refer to a document.]
Section 4. | Requirements
4.1 | Solar powered vaccine refrigerators or combined vaccine refrigerator and water-pack freezers (appliances) are used primarily in areas without any electricity, or where there is less than eight hours of reliable electricity over a typical day.
| Reliability, durability and effective maintenance are essential for a successful installation. The associated power system must be designed to match both the appliance power consumption and local climate conditions (i.e. ambient temperatures and solar radiation resource). All appliances must fully meet the stated performance requirements with a solar radiation reference period of 3.5 kWh/m²/day and hot zone (+43°C) test temperatures. In addition, manufacturers may offer appliances suitable for hot zone and a lower temperature zone and/or lower solar radiation reference period. | [Please state (or refer a document) how the power system matches the appliance power consumption.]
| Net vaccine storage capacity bands of the refrigerator are based on the capacity bands prescribed and utilized by UNICEF SD. These are: < 30 L; 30 L < 60 L; 60 L < 90 L; 90 L < 120 L; 120 L < 150 L, ≥150 L and above. | [Please state the vaccine storage capacity.]
4.2 | Performance
4.2.1 | Operating temperature range
| The operating temperature range is indicated on the temperature zone rating sticker that is attached to the front of the appliance (see Annex 1). All appliances must be able to maintaining the acceptable temperature range in the vaccine storage compartment when operating in the hot zone (+43°C) and at a minimum rated ambient (MRA) temperature of +10°C or lower. | [Please state the temperature zone, MRA and show a photo of this label on the appliance.]
4.2.2 | Refrigeration cycle
| Refrigeration cycle must be an electrically powered compression-cycle unit or thermoelectric appliance operating with input of direct current (DC) electricity generated with solar energy. | [Please state.]
4.2.3 | Design of the vaccine storage compartment
| The vaccine storage compartment must be designed so that no part which is outside the acceptable temperature range can be used to store vaccines, either by inadvertent or deliberate misuse. As per the freeze-protection classification features, appliances complying with this requirement without demanding any intervention from the user will be published as Grade A. Others will be published as Grade B or Grade C depending on the level of interventions required. | [Please state which freeze protection grade, the test reference that proves this and if not Grade A, what provision is made to ensure vaccine does not come into contact with cold walls in the appliance.]
| Areas of an otherwise acceptable appliance which are too warm or too cold must be excluded from vaccine storage by design. The vaccine storage compartment must provide some means (such as baskets) to enforce physical separation between the vaccines and any surfaces that potentially have condensate on them (such as the floor, ceiling and/or walls of the compartment). | [Please state how the vaccine storage is prohibited from areas outside 2°C – 8°C.]
4.2.4 | Vaccine freeze protection classification
The vaccine freeze protection classification is indicated on the freeze protection classification sticker attached to appliance front (see Annex 3). The amount of user intervention required to ensure that the vaccines will not be exposed to freezing temperatures, when the appliance is used within its nominated temperature range and minimum rated ambient temperature, will be classified and reported as follows:
| • Grade A, user-independent freeze protection (UIFP): when there is no user intervention required to ensure that the vaccines will not be exposed to freezing temperatures whatever the position of the vaccine in the vaccine storage compartment. | [Please state the freeze protection grade and show a photo of this label on the appliance.]
| • Grade B, user-dependent freeze protection (UDFP): when the user must comply with a procedure provided by the legal manufacturer and requiring one level of user intervention (e.g., the requirement to use baskets to avoid vaccine freezing temperatures constitute one level of user- intervention). | [Please state the freeze protection grade and show a photo of this label on the appliance.]
| • Grade C, user-dependent freeze protection (UDFP): when the user must comply with a procedure provided by the legal manufacturer requiring two or more levels of user interventions (e.g., a refrigerator that not only requires the use of baskets but also requires use of removable thermal barriers constitutes two levels of user intervention). | [Please state the freeze protection grade and show a photo of this label on the appliance.]
4.2.5 | Water-pack freezing capacity (combined units only)
| In combined units with freezer compartment, a minimum of 1.6 kg and not less than 2.4 kg per 50 litres of gross freezer volume must be frozen per 24 hours. Some overnight ice loss is acceptable – see clause 4.2.7. | [Please state the water-pack freezing capacity and the test reference that proves this.]
4.2.6 | Water-pack storage compartment capacity (combined units only)
| The freezer compartment to hold a minimum of 3.2 kg of fully frozen water-packs and at least twice the daily water-pack freezing capacity determined by E003/RF05 VP.6. The water-packs must comply with E005/IP01. These may be frozen gradually over several days. Some overnight ice loss is acceptable – see clause 4.2.7. | [Please state the water-pack storage capacity and the test reference that proves this.]
4.2.7 | Temperature control
| Refrigerator compartment: The zone within the vaccine compartment that is designated for vaccine storage must remain within the acceptable temperature range during any continuous ambient temperature test(s) or day/night cycling temperature test. Combined units must achieve this performance with or without water-packs in the water-pack compartment. | [Compliance to this clause must be demonstrated in the laboratory testing.]
| Water-pack freezing compartment: The compartment (if present) must be capable of producing fully frozen water-packs for use by health workers at the beginning of each working day. Under the water-pack freezing tests, the temperature of the water-pack freezing compartment is permitted to exceed 0°C during the 12-hour night phase and the first three hours of the 12-hour solar phase. The net amount of fully frozen water-packs remaining under worst -case overnight test conditions and day/night cycling per E003/RF05 VP.6 will be visually estimated and reported. | [Compliance to this clause must be demonstrated in the laboratory testing. Please make precise reference to these tests.]
4.2.8 | Thermostat
| The thermostat must be set to prevent freezing in any part of the vaccine storage compartment. The thermostat must be effective throughout the ambient operating temperature range (down to the minimum rated ambient temperature – see clause 4.2.13). It must be designed so that it cannot be adjusted by the user. A means for adjustment by a technician is acceptable provided the device is protected from user interference (e.g. by location within the appliance cabinet). Alternatively, programmable thermostats may be password-protected. Bulb and capillary tube thermostats are not acceptable. | [Compliance to this clause must be demonstrated in the laboratory testing. Please state how thermostatic settings are adjusted by a technician.]
4.2.9 | Temperature monitoring and thermometer
| The appliance shall have a data logger integrated within the appliance with connections for data transfer and power known as the Machine-to-Machine interface (M2M) that supports the functionality of external devices, including external EMDs. The data logger and M2M support data recording and local access to data. This data logger must conform to the requirements of E006/DL01 Data logger and machine-to-machine interface for Equipment Monitoring Systems. If the appliance is a combined refrigerator and vaccine freezer, both compartments shall be monitored by the data logger. | [Please provide evidence a E006/DL01 datalogger is integrated into the appliance. Please confirm the datalogger is approved for use with that appliance.]
| The appliance shall have a display to show internal cabinet temperature and positioned so that it can be easily read by a health worker without opening the appliance. This electronic thermometer display may be configured as part of an electronic monitoring system (EMS) or as part of a thermometer conforming to E006/TH06. The internal sensor shall be located so that the temperature reading gives the coldest temperature in the vaccine refrigerator compartment. | [Please state whether the appliance has a display per E006/TH06 or integrated EMD (E006/EM01). Please state which configuration per E003/RF05.8 Clause 4.2.9 Table 1 the appliance complies with.]
| The EMS shall last throughout the 10-year expected lifetime of the refrigerator, so that the data logger and monitor is always active on-site. If any data logger components do not have a 10-year lifetime (e.g. batteries), the refrigerator manufacturer must also provide the consignee with replacement components to ensure the EMS will operate for at least 10-years. Use, maintenance and any replacement of data logger components e.g. rechargeable battery, shall be stated in the user instructions. | [Please confirm the EMS has an expected lifetime of 10 years. Please state the epected lifetime of the EMS batteries.]
| A 30-day temperature recorder (30DTR) may be provided in lieu of an EMD and shall conform to E006/TR06. If a 30DTR is provided, the refrigerator manufacturer must also provide the consignee with replacement temperature monitoring devices throughout the 10-year lifetime. The expected lifetime, along with the warranty, of any monitoring device provided shall be stated in the user manual. | [Please state if a 30DTR is being used instead of an EMD.]
| If a manufacturer ships the E-EMD or 30DTR separately, they should arrive at the same destination as the refrigerator shipment, on the same date, and addressed to the same consignee. This is necessary to allow for smooth in-country receipt and assembly of the refrigerators and temperature monitoring devices. Separate shipments of E-EMD or 30DTR must be tracked and coordinated by the refrigerator manufacturer. | [Please state if an E-EMD or 30DTR is being used, and if they are being shipped with the appliance.]
4.2.10 | Humidity control
NOTE: As of the publishing of this specification version, refrigerators will be required to be tested for humidity control as part of the laboratory verification protocol, but compliance with the requirement directly below will not be required until publishing of an updated version of this specification document. The intended time frame will be to require compliance as of January 2023, pending review of equipment performance tests up to that time and publication of updated specifications and verification protocols.
| Refrigerator compartment: The refrigerator compartment must have humidity control. The zone within the vaccine compartment that is designated for vaccine storage must remain within acceptable compartment humidity levels during humidity control testing as specified in the verification protocol. Combined units must achieve this performance with or without water-packs in the water-pack compartment. | [Please state the level of humidity control as shown by testing in accordance with E003/RF05-VP.6 Test 3, or the development in the humidity control so far.]
Water-pack freezing compartment: Not applicable
4.2.11 | Indicator light
| A minimum of one green LED indicator light is required to be located on the front or top of the appliance to alert users that the cooling system is actively operating. A constant green LED light is required to indicate that the compressor or cooling system is active and the light must go off when the compressor or cooling system is off. | [Please refer to a photo with these indicator light(s).]
| Optionally, additional indicator lights may be added to indicate other operating conditions including temperature and faults. | [Please refer to a photo with these indicator light(s).]
4.2.12 | Autonomy
| All solar direct drive refrigerators must prequalify with at least a minimum autonomy of three days at the minimum solar radiation reference period and when tested for the hot zone. Autonomy (refrigerator) is determined by testing in accordance with E003/RF05-VP.6. This prequalification testing determines the autonomy of the refrigerator, and the prequalified appliance will be limited to installation sites where the refrigerator autonomy meets or exceeds the installation site autonomy requirements. (Refrigerator autonomy must always meet or exceed the autonomy requirement of the installation site) | [Please state the autonomy result and give a precise reference to the laboratory test report where this is demonstrated.]
| Site autonomy is determined by the methods described in E003/PV01 clause 4.1.2. Installation site conditions may require autonomy (refrigerator) of more than three days. | [With reference to E003/PV01.4 Clause 4.1.2, please state how the solar power system meets the requirements of this clause.]
| All solar direct drive freezers must prequalify with a minimum autonomy of overnight (0.5 days) at the minimum solar radiation reference period and when tested for the hot zone. The minimum autonomy (freezer) is confirmed by testing in accordance with E003/RF05-VP.6 | [Please state the night time autonomy result and give a precise reference to the laboratory test report where this is demonstrated.]
4.2.13 | Minimum rated ambient temperature
| All models must be able to operate at a continuous minimum ambient temperature of +10.0°C or lower whilst maintaining the acceptable temperature range. The maximum performance rating is achieved if the vaccine load remains within the acceptable temperature range at -10°C. A freeze-prevention circuit may be required to protect against freezing at low ambient temperatures. | [Please state the minimum rated ambient and the test reference that proves this.]
4.2.14 | Power system requirements and consumption
| Compatible Type 2 solar power system to directly power the appliance is specified in E003/PV01, with power capacity (watts) equal to or greater than the power system capacity used for PQS prequalification testing. If supplied, the legal manufacturers must certify compliance. | [Even if the solar power system is not provided, please state the specifications required to power the appliance.]
| Solar module voltage up to 45 volts open circuit (Voc at Standard Test Condition of solar radiation 1000W/m², cell temperature +25°C, air mass 1.5) is acceptable provided all electrically powered system components are integrated in such a way that performance and component life is not reduced by voltage input from the solar array. | [Please state open circuit voltage of solar module.]
| No standard is set for power consumption; however, performance data will be published including power consumption (watt-hours/day) and the cooling system percentage on-time over the 12-hour solar phase. | [Please state the energy consumption / day and the test references that show this.]
4.2.15 | Condensation management and defrost
| The conditions within the zone designated for vaccine storage must be designed so that vaccine primary containers and vaccine cartons are not exposed to levels of humidity which may cause damage to cartons or primary container labels or create a risk of mould growth. | [Please describe the design features for humidity control and reference the particular section in the user manual. Also,describe how any defrost switches are safe from accidental operation.]
To alleviate humidity damage legal manufacturers are to include refrigerator design features and/or provide containers for vaccine storage. Condensate and defrost drainage must be provided in all refrigerator and freezer compartments. If used, the defrost switch (or switches) must be accessible to the user without tools but must be protected from accidental changes in position.
4.2.16 | Lock
| The door or lid must be fitted with a lock. Two keys are to be supplied with every unit. | [Please declare compliance.]
4.2.17 | Corrosion resistance
| The legal manufacturer must certify compliance that internal and external cabinet, lid and frame are protected against corrosion as appropriate to EN ISO 6270-1 / ASTM D2247 / EN 13523-26, EN ISO 6270-2 / EN 13523-25, ISO 6272 / EN 13523-5 and ISO 2409: 2013. | [Please state the level of protection for various components (internal and external) and reference the test information that proves this.]
4.2.18 | Electrical safety rating
| The legal manufacturer must certify compliance with IEC 60335-1, IEC 60335-2-24 and IEC 60364-1. | [Please reference the test information that proves compliance to this clause.]
4.2.19 | Markings and labelling
| If used, compressors must be marked with the blue identifying symbol shown in Annex 2. In addition, the cabinet must be permanently marked, near the compressor position, with the chemical name of the refrigerant or with the refrigerant number, formula or proportion (for blended refrigerants). | [Please refer to photos of the compressor label and the refrigerant label (on the appliance) in accordance with this clause.]
| Appliances operating on R600a must be marked with the warning symbols shown in Annex 2. | [Please refer to a photo of the warning label on the appliance. Note this is required for all flammable refrigerants.]
| Appliances not utilizing compressor-based cooling methods must be marked identifying any refrigerant and/or heat transfer fluids used including, but not limited to, the chemical name, formula or proportion (for blended refrigerants). | [If applicable, please refer to photos of the refrigerant label (on the appliance) in accordance with this clause.]
| All appliances must label hazardous materials and include a Safety Data Sheet. Label and Safety Data Sheets must comply with the Globally Harmonized System for the Classification and Labelling of Chemicals GHS Rev.5. | [Please refer to a copy of the Safety Data Sheet.]
| The appliance must carry the following additional information fixed to the front of the cabinet: manufacturer and model number (unless already located on the front of the unit), serial number, date of manufacture, PQS identification number, applicable service phone number and website URL. This label to remain readable for the expected life of the appliance. | [Please refer to a photo of this label.]
| All appliances require an asset identification label (bar code) as specified in WHO/PQS/E003/BC01.1: Global asset identification. Effective date June 30, 2020. | [Please refer to a photo of the Global asset ID.]
4.2.20 | Vaccine storage advice
All appliances must carry a factory-fitted non-removable label, designed to last the lifetime of the appliance and carrying the following information:
| • Vaccine refrigerators: Vaccine storage instructions. | [Please refer to a copy of this label and also show its position on the appliance.]
| • Combined appliances: Vaccine storage instructions and water-pack freezing instructions. | [Please refer to a copy of this label and also show its position on the appliance.]
| The instructions should be fixed to the lid of chest appliances and near the top of the door on upright appliances. Instructions should be in one of the languages specified in clause 4.11 and as indicated by the purchaser at time ordering. If removable baskets are required to avoid freezing temperatures fix a multilingual warning within the refrigerator instructing users to “Store vaccine in baskets only”, or other appropriate instruction. | [Please refer to a photo of the instructions showing its position on the appliance. Please state whether storing vaccines in baskets or similar is required to avoid freezing temperatures.]
4.2.21 | Electromagnetic compatibility
| The legal manufacturer must certify compliance with the requirements of the latest edition of IEC 61000-6-1 and IEC 61000-6-3. | [Please refer to a test report proving EMC compliance.]
4.3 | Environmental requirements
4.3.1 | Ambient temperature range during transport and storage
| The ambient temperature range during transport and storage is -30°C < +70°C when the product is inactivated. | [Please state ambient temperature range during transport and storage.]
4.3.2 | Ambient humidity range during transport, storage and use
| The ambient humidity range during transport, storage and use is 5% < 95% RH, non-condensing | [Please state ambient humidity range during transport, storage and use.]
4.4 | Physical characteristics
4.4.1 | Overall dimensions
| To allow for manoeuvring through corners, corridors and doorways, the minimum dimension of the product (either length, width or height) should not exceed 710 mm; exceptionally, a minimum dimension up to 830 mm can be accepted, but this will restrict the number of sites where the appliance can be installed. The maximum dimension must not exceed 1700 mm and the maximum diagonal (corner to corner) dimension must not exceed 1850 mm. | [Please state the external dimensions.]
4.4.2 | Weight
| Mechanical lifting equipment will typically not be available at the installation sites. It is recommended that the appliance and any associated components be designed for lifting in such a way that no single worker is required to carry more than 25 kg whilst working on their own, or in a group. | [Please state how the appliance may be easily lifted or manoeuvred into position on site.]
4.5 | Interface requirements
4.5.1 | Electrical components
| All electrical components must be compatible with a Type 2 solar power system as specified in specification E003/PV01. The appliance must be equipped with a locking female and male coupler system (“plug and play”) that is compatible with the solar array interconnection cable and polarity. | [Please describe (including a diagram or photo) the connecting components in accordance with this clause.]
4.5.2 | Power switch
| Each appliance to provide an on and off power switch that is readily accessible to the user, either on the outside of the appliance cabinet or in a wall-mounted switch within one metre of the appliance. Plug and play solar array cable connectors (i.e. locking female and male coupler system) are not acceptable for on and off switching of the solar power system or appliance. | [Please describe (including a diagram or photo) the power switching in accordance with this clause.]
4.6 | Human factors
4.6.1 | General design
| The product must be usable by the widest practicable range of active health workers, regardless of age, gender, size or minor disability, including colour blind users and long-sighted people without glasses, in accordance with the general principles laid out in ISO 20282-1: 2006. | [Please state how your appliance complies with this clause.]
4.6.2 | Control panel, indicator light and thermometer
| Control panel, indicator light, thermometer and other visual displays may be positioned on the front of the unit preferably as close to eye level as possible. Alternatively, they may be mounted on top of the unit at a height not exceeding 1.3 metres. If a low-level position is essential, the display should be aligned so that it can easily be read without the user having to squat or kneel down. The on and off and/or defrost switch, if present, should be recessed or otherwise protected from accidental change in position. | [Please refer to a photo that shows compliance with this clause.]
4.7 | Materials
4.7.1 | Refrigerant
| Hydrocarbon (HC) refrigerants such as R600a or other gases with global warming potential (GWP) ≤ 11 and zero ozone-depletion potential (ODP) are required. Existing appliances with HCFC refrigerants including R134a will be phased out over a transition period of four years from issue date of this specification. The suitability of alternative refrigerant gases will continue to be assessed and preference will be given to appliances that use gases with low GWP. | [Please state type of refrigerant(s) and their GWP.]
4.7.2 | Thermal insulation foaming agents
| Any gas complying with the limitations and deadlines set by the Montreal Protocol on the elimination of ozone-depleting chemicals. | [Please state type of foaming agent, it’s compliance with the Montreal Protocol and its GWP.]
4.7.3 | Other restricted materials
| The appliance and its constituent components, must not contain lead, mercury, cadmium, hexavalent chromium, polybrominated biphenyls (PBB) or polybrominated biphenyl ethers (PBDE). | [Please declare compliance with this clause.]
4.7.4 | PCM
| PCM - Integrated thermal buffer materials may be used to prevent freezing temperatures from propagating to the vaccine storage compartment or for other thermal purposes. The buffer material may be PCM-based but if so, must comply with WHO/PQS/ E005/PCMC0.1– PCM specification for phase-change material containers. | [Please state whether your appliance uses a thermal buffer an state the compliance of PCMs if any.]
4.8 | Warranty
| The appliance must be covered by a three-year replacement warranty in the event of any component failure arising from defective design, materials or workmanship. | [Please refer to a statement of the warranty in the user manual and how a purchaser may make a claim against this warranty.]
4.9 | Servicing provision
| The appliance and solar power system must be designed to achieve a low-maintenance life of not less than 10 years apart from routine defrosting, cleaning and solar array cleaning and shading prevention. | [Please state how the appliance is designed to last at least 10 years.]
4.9.1 | Essential spare parts and user maintenance tools/supplies:
| All sites are to be supplied with one complete user maintenance kit consisting of all necessary operations and maintenance tools as proposed by the legal manufacturer. Based on product design and requirements the type and quantity of spare parts, basic installation tools/supplies, user and technician maintenance manuals (see clause 4.11), must be determined and agreed upon in advance of order placement. As a minimum, each appliance to be supplied with 10 spare fuses of all fuse size and type used in the appliance. The spares fuses are to be attached within or on the appliance. | [Please describe (including diagrams or photos) the maintenance kit in accordance with this clause. Show the location of the spare fuses.]
| Legal manufacturers are to publish a list of spare parts recommended for purchases of 10 and 50 appliances and power systems. The legal manufacturer must ensure supply of spare parts for a minimum of five years from the time of cessation of the last production of equipment. Spare parts are to be provided in kit form for storage in appropriate quantities at National or Sub-national level in the purchasing country, as agreed with the purchasing agency. | [Please declare compliance with this clause and how spare parts may be obtained e.g. via local agents.]
4.1 | Disposal and recycling
| The legal manufacturer must provide information to the buyer on the hazardous materials contained within the system and suggestions for resource recovery/recycling and/or environmentally safe disposal. For legal manufacturers from the European Union WEEE compliance in accordance with European Union Directive 2002/96/EC is mandatory. | [Please refer to a statement hazardous material in the user manual and the provision for disposal/recycling.]
4.11 | Instructions
Each appliance to include a separate user manual and technician installation manual in Arabic, English, French, Mandarin Chinese, Russian and Spanish. An English version of all instructions and manuals are required to be supplied at time of laboratory testing.
| Instructions to include easy to understand visuals whenever possible to avoid reliance on text. | [Please note this clause.]
The user manual must include the following information:
| • Health and safety guidance, | [Please refer to a particular section in the user manual.]
| • Basic operations description, and | [Please refer to a particular section in the user manual.]
| • Cool-down time instructions to assure both acceptable temperatures and full autonomy is achieved, and | [Please refer to a particular section in the user manual.]
| • Preventive maintenance tasks (e.g. daily, weekly, and monthly). | [Please refer to a particular section in the user manual.]
The technician installation manual must include the following information:
| • Health and safety guidance, | [Please refer to a particular section in the installation manual.]
| • Detailed operations description, | [Please refer to a particular section in the installation manual.]
| • Cool-down time instructions to assure both acceptable temperatures and full autonomy is achieved, | [Please refer to a particular section in the installation manual.]
| • Correct handling to avoid appliance damage and for the safety of handling persons | [Please refer to a particular section in the installation manual.]
| • Installation procedures, | [Please refer to a particular section in the installation manual.]
| • Compatible solar power system voltage, | [Please refer to a particular section in the installation manual.]
| • Preventive maintenance tasks (e.g. daily, weekly and monthly), | [Please refer to a particular section in the installation manual.]
| • Corrective maintenance procedures, | [Please refer to a particular section in the installation manual.]
| • Diagnostic and repair procedures, | [Please refer to a particular section in the installation manual.]
| • Itemised list of spare parts including part numbers, | [Please refer to a particular section in the installation manual.]
| • End-of-life resource recovery and recycling procedures, and | [Please refer to a particular section in the installation manual.]
| • User training guidance | [Please refer to a particular section in the installation manual.]
| Printed user-operations and maintenance instructions specifically directed at the health centre or store staff must be pictorial. All key information must be summarized on a single pictogram sheet fixed onto the appliance lid or near the top front of chest appliances and near the top of the door on upright appliances. The sheet must be sufficiently durable to last the life of the appliance. | [Please refer to photo which shows compliance with this clause.]
| Installation, repair and servicing instructions must be supplied in printed format, and optionally on DVD and/or online to instruct the installation technician installation standards and practices specific to the appliance and its power system. | [Please state all sources for installation, repair and servicing instructions.]
4.12 | Training
| If requested, all legal manufacturers are required to have the capability of providing in-person training in the countries where their product is deployed. Training may be delivered by the legal manufacturer, manufacturer’s representative or reseller. The legal manufacturer must provide user training guidance to enable installation technicians to present user training. In addition, supporting video material supplied on DVD and/or online can be supplied to assist the instructor when delivering on-site user training. | [Please describe the different options for training which might be available.]
4.13 | Verification
| In accordance with PQS Verification Protocol E003/RF05-VP.6 | [Please refer to the laboratory test report.]