## Sheet1
E003 Transportable, powered vaccine storage appliances submission Template (V1.2 19 FEB 2025)
| Assessment of appliance: | [Brand & Model here] | For dossier review on [TBC]
| Tested in accordance with PQS verification protocol: | E003/TS01.1 (30 April 2021)
| To verify compliance with PQS specification: | TS01-VP.1 (30 April 2021)
ASSESSOR’S SUMMARY:
Instructions in [square brackets] are for your guidance. Please delete these instructions and replace with evidence or a precise reference (including page and section numbers) showing compliance with that particular PQS specification requirement. In general a "pass" or "complies" statement alone is not acceptable.
One document may provide evidence for several clause requirements. Please carefully reference the sub-sections of any such document so that it is clear that the clause requirement has been met. This will help speed the pre-qualification process.
Greyed-out cells do not require a response.
NOTE: If it is not clear or doubtful whether a requirement has been met, then this is a non-compliance and pre-qualification might be delayed.
|  |  | REVIEWER’S SUMMARY: |  | REVIEWER’S SUMMARY: |  | REVIEWER’S SUMMARY:
Section 7. | Product dossier | Applicant's submission | Reviewer's remarks [Name & Date] | Applicants response [Name & Date] | Reviewer's remarks [Name & Date] | Applicants response [Name & Date] | Reviewer's remarks [Name & Date]
The legal manufacturer or reseller is to provide WHO with a pre-qualification dossier containing the following:
| Dossier examination fee in US dollars. | [Please contact PQS.]
| General information about the legal manufacturer, including name and address. | [Please state or refer to a document.]
| Unique identification reference for the product type. | [Please state.]
| Brand name of the product. | [Please state.]
| Full specifications of the product being offered, covering all requirements set out in this document, including details of product marking and traceability. Also including the narrative description as applicable and required in Section 4.2.17 regarding product safety. | [Please refer to a document.]
| Designation of appliance type (lightweight appliance, medium-weight appliance, or heavy-weight appliance) and freeze protection classification (Grade A, B, or C) to be verified by PQS and the testing organization. | [Please state which type and which freeze protection classification.]
| Full details of the required, dedicated, compatible solar power system per specification WHO/PQS/E003/PV01 (if applicable). | [If applicable, please refer to a document.]
| A comprehensive set of photographs including photographs of the appliance with the door or lid open, the appliance closed, all external surfaces of the appliance, the interior layout, the compressor or cooling system, and close-ups of the thermometer, indicator light(s), control(s), control panel, securing mechanism (if applicable), and any other special features. | [Please refer to a document & section.]
| Certified photocopies of all type-approvals obtained for the appliance. Appliances shall carry the CE mark, UL mark and/or equivalent internationally accepted evidence of conformity assessment. | [Please refer to a document.]
| Certified photocopies of the legal manufacturer’s current ISO 9001 quality system certification. | [Please refer to a document.]
| Where relevant, certified photocopies of the legal manufacturer’s ISO14001 certification, EMAS registration or registration with an equivalent environmental audit scheme. Conformity with an environmental audit scheme is not mandatory however, preference will be given to manufacturers who are able to demonstrate compliance with good environmental practice. | [Please refer to a document.]
| Laboratory test report(s) proving conformity with PQS specifications. | [Please refer to a document.]
| Indicative cost of the product per unit, per 10 units and per 100 units, EXW (Incoterms 2010). | [Please state or refer to a document.]
Section 4. | Requirements
4.1 | General
Transportable, powered vaccine storage appliances may be used in multiple settings for transport from locations with reliable power or electricity supply (i.e. 20 or more hours of reliable electricity per typical day) to areas with less reliable, or no electricity. They may also be used for short-term storage in locations that do not or cannot support a more permanent refrigeration point in the cold chain. Appliances, including any stationary components, shall carry the CE mark, UL mark, and/or equivalent internationally accepted evidence of conformity assessment. The mass, volume, and intended use of the appliance should fall within one of the three following categories for the purposes of this specification and the corresponding verification protocol WHO/PQS/E003/TS01-VP.1:
| Lightweight appliance: intended primarily for short-term vaccine transport with or without a vehicle for transportation and sized for carrying by one person likely to a single destination - having a maximum loaded mass of less than 8 kg and a vaccine net storage capacity of greater than 1.5 litres | [Please state which type for your appliance.]
Medium-weight appliance: intended primarily for longer-range transportation, possibly to multiple destinations by a vehicle (e.g. truck, motorbike, donkey, camel), and movable for at least short distances by a single person - having a maximum loaded mass of less than 25 kg and a vaccine net storage capacity of greater than 1.5 litres.
Heavy-weight appliance: intended primarily for temporary, short- term vaccine storage and transport by powered vehicle and movable by two or more people to and from a vehicle - having a maximum loaded mass of less than 50 kg and a vaccine net storage capacity of greater than 5 litres.
4.2 | Performance
| If the appliance is intended to be powered, charged, or otherwise prepared for use by more than a single type of power source or process, performance will be tested and verified separately with each source or process. | [Please state the type of power source(s).]
4.2.1 | Operating temperature range
| All appliances shall be suitable for operation in the hot zone (+43°C maximum temperature) and at a minimum rated ambient temperature of +10°C or lower, while maintaining the acceptable temperature range in the vaccine storage compartment. | [Please refer to a document & section that shows evidence for the ambient temperature range.]
4.2.2 | Refrigeration cycle
| Refrigeration can be accomplished using a compression cycle with one or two compressors or with thermoelectric components. Other means of cooling (e.g. absorption-cycle, evaporative cooling), may also be acceptable but shall be discussed with PQS prior to prequalification testing and approval. However, absorption-cycles or other appliances that require an open flame or other combustion during use or preparation are not allowed due to safety concerns specific to transportable appliances. | [Please state which type for your appliance.]
4.2.3 | Design of vaccine storage compartment
| The vaccine storage compartment shall be designed so that no area or surface that reaches temperatures outside of the acceptable temperature range in normal use or during testing can be used to store vaccines either by inadvertent or deliberate misuse. | [Please refer to a document & section that shows evidence for the internal temperature range.]
| As per the classification of freeze protection features (Section 3, Terms and Definitions) appliances complying with this requirement without demanding any intervention from the user will be published as Grade A. Others will be published as Grade B or Grade C depending on the number of user-interventions required. | [Please state which freeze protection classification.]
4.2.4 | Vaccine freeze protection classification
| The classification shall be designated by the legal manufacturer and indicated on the freeze protection classification sticker attached to the appliance (see Annex 3). The number of user-interventions required to ensure that the vaccines will not be exposed to freezing temperatures when the appliance is used within the hot zone temperature range and minimum rated ambient temperature will be classified and reported as Grade A, Grade B, or Grade C. These grades are defined in Section 3 Terms and Definitions. | [Please refer to a document that shows evidence for the freeze protection classification.]
4.2.5 | Temperature control
| The entire vaccine load shall remain within the acceptable temperature range during any continuous ambient temperature test(s) or cycling temperature test(s) and down to the minimum rated ambient temperature in the minimum rated ambient temperature test (see Section 4.2.13). For appliances that are intended to be powered, charged, or otherwise prepared for use by more than a single power source or process, this requirement applies to operation when powered, charged, or prepared by each possible power source or process independently. | [Please refer to a document & section that shows evidence for the internal temperature range.]
4.2.6 | Cool-down time
| The appliance shall be able to cool the vaccine storage compartment to +8°C or below within four hours or less without any external power input. Internal or integrated power sources such as batteries, rechargeable chemical mechanisms, or others can be used for this cooling task. This functionality will be tested at +43°C after following any instructions for charging or otherwise preparing the appliance and after the appliance indicates that it is prepared and safe for use. The cool-down time will be established in verification testing per the verification protocol WHO/PQS/E003/TS01-VP. | [Please refer to a document & section that shows evidence for the cool-down time.]
4.2.7 | Use-Reuse cycle
| The appliance shall be capable of being charged, conditioned, or otherwise prepared for use within 12 hours of a previous use, while receiving any manufacturer-specified minimum power requirements for the appliance. It shall also be possible to delay the use of the appliance by a minimum of 24 hours after being fully charged or prepared for use, and meet all performance requirements (e.g. cool-down time, independence, etc.). | [Please refer to a document & section that shows evidence for use-reuse cycle parameters.]
4.2.8 | Indicator of proper preparation for use
| If the appliance needs to be charged, conditioned, or otherwise prepared before being used to transport vaccines, there shall be a clear indication for the user to assess whether the appliance is ready. Preferred designs will hinder the user from loading or using the appliance unless it is properly prepared. However, if this functionality is not present, an obvious indication to the user will be acceptable. This safe-for-use indication may be an obvious light or display, temperature reading, coloured indicator, or other form. | [Please refer to a document & section that shows how the design prevents reuse or has a safe-for-use indicator e.g. user manual.]
4.2.9 | Thermostats
| The thermostat or other control mechanism shall be set to prevent freezing in any part of the vaccine storage compartment. It shall be effective throughout the ambient operating temperature range (down to the minimum rated ambient temperature – see Section 4.2.13). It shall be designed so that it cannot be adjusted by the user. A means for adjustment by a technician is acceptable provided the device is protected from user interference (e.g. by location within the appliance cabinet). Alternatively, programmable thermostats may be password-protected. Bulb and capillary tube thermostats are not acceptable. | [Please refer to a document & section that shows the thermostat and technician adjustment if applicable e.g. user manual.]
4.2.10 | Temperature monitoring and thermometer
The appliance shall be equipped with a permanent, externally readable thermometer or display of one of the following types:
| ● A cabinet-mounted electronic thermometer conforming to PQS specification WHO/PQS/E006/TH06 of Type A or B as referred to in that specification. Referring to version TH06.2 of that specification, the requirements on the power source of the appliance can be ignored. | [Please give details of the type and power source.]
● An electronic thermometer powered by a photovoltaic cell conforming to PQS specification WHO/PQS/E006/TH06 which forms part of the device. This type draws no power from the appliance in which it is installed and is referred to as Type C in that specification. Referring to version TH06.2 of that specification, the requirements on the power source of the appliance can be ignored.
| The location at which the temperature is monitored for display shall be in the zone of the compartment with the lowest temperature during appliance operation. If the appliance is or can be housed in a secondary carrier for transport (e.g. a backpack or satchel), it is allowable for the temperature display to be covered by the secondary carrier as long as the user is not required to open the primary appliance to view the temperature reading. | [Please refer to a photo that shows the location of the temperature sensor.]
| OPTIONAL: Additionally, some appliances may be equipped with a temperature monitoring functionality that supports the transfer of data to another system for analysis purposes. This functionality is optional and can be combined with, or separate from the external monitor or display. However, if it is included, it must comply with the following: | [Please state if this is included.]
| ● If the temperature monitoring device is separate and not permanently physically integrated with the appliance, it shall be a 30-day temperature logger, prequalified by WHO PQS as complying with WHO/PQS/E006/TR06. The sensor shall be located in an integrated holder within the vaccine storage compartment with or without an external sensor lead. The holder shall be positioned so that the device can easily be read by the health worker, and shall be located so that temperature readings are taken in the zone of the compartment with the lowest temperature during appliance operation. | [Please give details of compliance.]
| ● If the legal manufacturer ships temperature monitoring devices with appliances, the temperature monitoring device and its battery shall not be adversely impacted by the shipping and storage conditions as specified in Section 4.3. | [Please give details of compliance e.g. a battery specification document as part compliance.]
| ● If the legal manufacturer ships temperature monitoring devices separately, they should arrive at the same destination as the appliance shipment, within the same week, and addressed to the same consignee. This is necessary to allow for coordinated in- country receipt and assembly of the appliances and temperature monitoring devices. Separate shipments of temperature monitoring devices shall be tracked and coordinated by the legal manufacturer. | [Please declare compliance to delivery requirements.]
| ● The legal manufacturer shall also provide the consignee with the expected number of replacement temperature monitoring devices throughout the 10-year expected lifetime of the appliance, so that the temperature monitoring device is always active on-site (i.e. if the monitoring devices have a design life of two years, a total of five devices will need to be supplied at two-year intervals). This cost is expected to be included in the upfront price of the appliance that is quoted to WHO PQS and purchasers. Scheduled replacement dates of the temperature monitoring devices must also be mentioned by the legal manufacturer in the instructions (see Section 4.11). | [Please give details of compliance.]
4.2.11 | Independence
| The appliance shall be able to maintain temperatures in the vaccine storage compartment between +2°C and +8°C for a minimum of 12 hours without any external power input or user interaction. This applies while in a location with a constant ambient temperature of +43°C. This independence time will be established in testing per verification protocol WHO/PQS/E003/TS01-VP after the appliance has initially cooled-down with all external power sources disconnected or switched off. Independence may include both powered, active cooling using energy stored by the device and passive cooling. (Holdover time is distinct from independence in that it only includes passive cooling capacity, while independence may include both active cooling and passive cooling.) | [Please refer to a document & section that shows evidence for independence.]
4.2.12 | Holdover time
| It is recommended that the appliance be able to maintain temperatures in the vaccine storage compartment between +2°C and +8°C for an extended period of time without any active cooling or power input. A holdover time of four hours or greater is recommended in case of failure or malfunction of any active cooling components in the appliance. This applies while in a location with a constant ambient temperature of +43°C. This holdover time will be established in testing per verification protocol WHO/PQS/E003/TS01-VP after the appliance has initially cooled down and all external power sources or means for active cooling have been disconnected, exhausted or switched off. (Holdover time is distinct from independence in that it only includes passive cooling capacity, while independence may include both active cooling and passive cooling.) | [Please refer to a document & section that shows evidence for holdover time.]
4.2.13 | Minimum rated ambient temperature
| All models will be tested to verify their stated minimum rated ambient temperature. Appliances shall be able to operate at a continuous minimum ambient temperature of +10°C or lower. A freeze-prevention circuit may be required to protect against freezing at low ambient temperatures. For appliances that can be powered by more than a single power source, this requirement applies to operation while powered by each possible power source independently. This minimum rated ambient temperature will be established in testing per verification protocol WHO/PQS/E003/TS01- VP. | [Please refer to a document & section that shows evidence for minimum rated ambient temperature.]
4.2.14 | Power system requirements and consumption
| If the appliance is designed to receive mains power, 220-240 volt 50/60 Hz and 100-127 volt 50/60 Hz options shall be offered. Thermal performance should be functionally comparable for all options, regardless of the nominal voltage and frequency rating of the appliance. The electrical components and system shall be selected and designed to be capable of withstanding inconsistencies in mains power including both transient and long-term voltage and frequency variability. | [Please state which options are offered.]
| If the appliance is designed to operate using DC it should generally be a 12 - 24 volt DC system. Other voltages may be acceptable but shall be discussed with PQS prior to qualification testing and approval. The electrical components and system shall be selected and designed to be capable of withstanding inconsistencies in the voltage and current that may be delivered to the appliance through vehicles or other DC sources. | [Please state which DC voltage options are offered.]
| There is no standard set for power or total energy consumption, however consumption will be reported in all tests as appropriate. All mains powered appliances shall be capable of starting and operating at 22% below the legal manufacturer’s stated voltage, as tested in the verification protocol WHO/PQS/E003/TS01-VP. | [Please refer to a document & section that shows evidence for low voltage (-22%) starts.]
4.2.15 | Batteries
| For appliances containing or using batteries the legal manufacturer shall supply evidence of compliance with UL 2054 as well as a material safety data sheet. For appliances using solar PV, if batteries are included as part of the stationary components and power system, they shall comply with specification WHO/PQS/E003/PV01. | [Please give details of compliance.]
If the batteries are used to power all or part of the primary appliance cooling system, they shall be rechargeable and:
| ● Replaceable by a technician, but not easily accessible or removable by the user, | [Please refer to a document & section e.g. user manual.]
| ● Of maintenance-free design, and | [Please declare compliance to this clause.]
| ● Robust enough to power the appliance and cool it within the acceptable temperature range for at least five years with expected cycling and recharging. The minimum intended lifetime of the appliance is 10 years so one battery replacement is allowed during that lifetime. | [Please refer to a document that details the battery specification.]
If single-use, non-rechargeable (i.e. primary cell) batteries are used to power secondary functions in the appliance (e.g. temperature monitoring, alarm system, etc.), they shall be:
| ● Replaceable by a user given detailed direction or instructions, | [Please refer to a document & section e.g. user manual.]
| ● A common type of battery that can be easily sourced in the locations where the equipment is procured | [Please give details of battery type.]
| ● Of maintenance-free design, and | [Please declare compliance to this clause.]
| ● Robust enough to power the applicable functions for at least two years with a target lifetime of three years. The minimum intended lifetime of the appliance is 10 years so five or fewer replacements are allowed during the appliance lifetime. | [Please refer to a document that details the battery specification.]
| NOTE: Battery performance can be significantly affected by temperature and state of charge. The usable lifetime must take into account that storage and use temperatures may frequently be elevated. Any manufacturer calculations, design, and testing should consider the hot zone temperature of +43°C and that temperatures in transportation vehicles or with exposure to sun could sometimes exceed this value. | [Please refer to a document that details the calculations for the battery performance in hot zone temperatures.]
Lithium batteries
| Lithium-based batteries can pose fire risks and are not allowed as part of PV systems per WHO/PQS/E003/PV01. Lithium-based batteries are also not allowed as part of the stationary components of the appliance system. Lithium-based batteries may be used in the primary, transportable portion of the appliance system; however, extreme caution, risk assessment and mitigation shall be used by the legal manufacturer in the appliance design, material selection, and production. | [If applicable, please refer to a document that details any lithium battery specification and required precautions in accordance with this clause.]
| The appliance may be exposed to harsh environmental conditions, as noted throughout this specification, that must be considered in terms of appliance and battery safety. Temperatures during transport, especially when located inside vehicles or exposed to direct sunlight, may exceed +43°C or reach below 0°C in certain regions of use. The legal manufacturer shall design the appliance, including necessary battery management systems, to account for safe operation under these conditions. As relevant, standard UL 2595 shall be applied in addition to UL 2054 to ensure the safety of appliance operation. If lithium-based batteries are used, the total capacity shall not exceed 1000 Wh and it is recommended that 300 Wh not be exceeded. Furthermore, the legal manufacturer should consider that shipping and transport of lithium-based batteries are frequently difficult and restricted especially through international borders. It is recommended to consider this in both sizing and selecting the chemistry of appliance batteries. | [Please refer to a document that details the battery specification and how it meets the requirements of this clause.]
4.2.16 | Closure and lock
| The lid, door, or other closure should be fitted with a mechanism to secure it in place so that the appliance does not open if it is dropped onto its side or onto its lid when full. Acceptable closure devices include, but are not confined to, magnetic or mechanical catches. It must not be possible for a catch to open accidentally once engaged due to vibrations or other similar, unintended causes. Mechanical catches shall be recessed so that they are fully protected against damage during transport and storage. Catches shall be maintenance-free, without need for lubrication and shall be secured to the container in a manner that prevents loosening due to vibration and frequent use. If the appliance is designated as medium-weight or heavy- weight, it shall be fitted with either an integrated lock or the door or lid must be fitted with a robust tab or other component to accommodate a padlock. If the lock is integrated or a padlock is shipped with the appliance, two keys shall be supplied with every unit to allow one set of spare key(s). Additionally, if the lock is integrated, the sub-component shall be available as a spare part and be able to be changed by a technician. | [Please give details (including photos) of the lock and catch mechanism.]
4.2.17 | Electrical and other safety considerations
| For appliances with relevant electrical components, the legal manufacturer shall supply evidence of and declare compliance with IEC 60335-1, IEC 60335-2-24 and IEC 60364-1. | [Please refer to a document that shows evidence for electrical safety compliance.]
| If the appliance, including any stationary components, presents additional safety risks to the user or other people likely to be in the vicinity of use such as children and patients, the legal manufacturer shall provide a narrative description of the design features, precautions, mitigants and other steps taken to address these safety risks. Specifically, this shall be provided for PQS review in the dossier for appliances that require potentially hazardous chemicals and/or high pressure (other than those used in common compressor driven cooling systems) and/or the intentional application of heat at temperatures above +50°C. Failure to comply with this initial reporting or to adequately address these safety concerns may result in immediate removal of equipment from PQS listing at the discretion of WHO PQS during, or at any time after, PQS prequalification. | [Please refer to a document & section that details the safety risks and precautions e.g. user manual. In addition, please provide photos of warning labels on the appliance.]
4.2.18 | Electromagnetic compatibility
| The legal manufacturer shall supply evidence of and declare compliance with the requirements of the latest edition of IEC 61000-6-1 and IEC 61000-6-3. | [Please refer to a document that shows evidence for EMC compliance.]
4.2.19 | Shape
| For lightweight and medium-weight appliances, the design should consider and promote comfortable carrying. Sharp edges and corners should be avoided; rounded external corners are preferred. It is also recommended that the design of medium-weight and heavy-weight appliances take into consideration the ergonomics of moving the appliance with multiple people and robust mounting on motorbikes or other modes of transport. The shape and design of appliances that require ventilation should also take into account protection of the ventilation air-flow pathways from accidental blockage, especially if packed in direct contact with other articles during transport. Appliances with such ventilation or other heat dissipation features should be designed (e.g. inclusion of components such as “standoffs”, “spacers”, or “bump stops”) to ensure adequate air circulation. The spacing and design should account for storage or transport on or next to compressible materials, like those often used for seats in vehicles that may inhibit air flow and heat transfer. | [Please refer to photos that show compliance to this requirement.]
4.2.20 | Hinges
| Hinges, where fitted, should allow the lid to open beyond 90° to give full access to the interior of the appliance. Preferably, the hinges should be recessed so that they are fully protected against damage during transport and storage. Hinges shall be maintenance-free, without need for lubrication and must be secured to the container in a manner which prevents loosening due to vibration. | [Please refer to photos that show compliance to this requirement.]
4.2.21 | Carrying components
| All carrying components shall be robustly constructed and firmly attached in order to survive rough handling. Both lightweight and medium-weight appliances shall be fitted with at least one carrying component arranged so that the appliance can be comfortably carried in a substantially upright position. Examples include: | [Please refer to photos that show compliance to this requirement.]
| • Carrying handle: a hinged, sliding or moulded-in handle attached to or forming an integral part of the appliance. | [Please state which type of strap(s) are available.]
• Shoulder strap: an adjustable padded strap arrangement that allows the appliance to be carried over the shoulder.
• Backpack: an adjustable padded strap arrangement that allows the appliance to be carried as a backpack.
| Heavy-weight appliances shall be fitted with at least two handles designed so that they can be used by multiple people to lift and carry the appliance comfortably when it is fully loaded. Hinge mechanisms, if present, shall restrain the handle in a near-horizontal position when the container is being moved or guided. Preferably, the handles should be recessed so that they are fully protected against damage during transport and storage. | [Please refer to photos that show compliance to this requirement.]
| Additionally, both medium-weight and heavy-weight appliances shall include features that facilitate securing, fastening, or otherwise anchoring the appliance to motorbikes or other likely modes of transport to prevent accidental movement during transport. | [Please refer to photos that show compliance to this requirement.]
4.2.22 | IP rating
| Protection of the appliance with lid or door closed and latched must not be less than that which is stipulated by IEC 60529: IP54 while operational and allowing no ingress of water into the vaccine storage compartment. The rating includes the entirety of the transportable portion of the device, not just the vaccine storage compartment. It is recommended that any stationary components be capable of meeting at least IP44. | [Please refer to a document that shows evidence for IP compliance.]
4.2.23 | Robustness and vibration
| The appliance with a full load of vaccine vials and diluent ampoules shall withstand a 0.85 metre drop onto the 11 faces, edges, corners or chimes (circumferential edges) that are most likely to receive impact, and when tested in accordance with ASTM D5276-09. At the end of the test there shall be no damage that significantly affects the performance of the appliance and the lid or door must still close and latch correctly although some damage to the appliance or contents is acceptable. The appliance with a full load of vaccine vials and diluent ampoules must also withstand vibration testing typical of unpaved road conditions in accordance with ASTM D4169-09, with no damage to either the appliance or contents. Drop and vibration tests will be carried out on each appliance per WHO/PQS/E003/TS01-VP. | [Please refer to a document & section that shows evidence for compliance to this clause.]
4.3 | Environmental requirements
4.3.1 | Material properties at ambient temperatures
| Appliance materials and components shall be able to withstand temperatures during transport and storage of -30°C to +70°C when the product is inactivated or turned off. | [Please refer to a document & section that shows evidence for compliance to this clause e.g component specifications.]
4.3.2 | Ambient humidity range during transport, storage and use:
| Appliances may be subject to a wide range of uncontrolled ambient humidity up to condensing, 100% RH and should be designed to remain undamaged and fully functional if exposed to these conditions in hot zone temperatures (up to +43°C). | [Please refer to a document & section that shows evidence for compliance to this clause e.g component specifications.]
4.4 | Physical characteristics
4.4.1 | Overall dimensions
| To allow for manoeuvring through corners, corridors and doorways, the minimum dimension of the product (length, width, or height) should not exceed 710 mm. The maximum dimension shall not exceed 1700 mm. The maximum diagonal dimension (e.g. corner-to-corner distance in any plane of a rectangular prism) shall not exceed 1850 mm. | [Please give the dimensions or refer to a document & section.]
4.4.2 | Weight
| Mechanical lifting equipment will typically not be available at sites. It is recommended that the appliance and any associated components, including stationary components, should be designed for lifting or moving in such a way that no single person is required to carry more than 25 kg, including whilst moving a heavy-weight appliance on their own, or in a group. The maximum loaded mass of the appliance, inclusive of any physically separate components used during transport but excluding any stationary components, shall not exceed the following figures: | [Please state the weight and how the appliance is intended to be carried e.g. refer to a section in the user manual.]
| ● Lightweight appliance: 8 kg | [Please state which category.]
● Medium-weight appliance: 25 kg
● Heavy-weight appliance: 50 kg
4.4.3 | Internal Volume
| The vaccine net storage capacity shall meet the following minimum requirement based on the type designated by the legal manufacturer: | [Please state the internal volume.]
| ● Lightweight appliance: greater than 1.5 litres | [Please state which category.]
● Medium-weight appliance: greater than 1.5 litres
● Heavy-weight appliance: greater than 5 litres
4.5 | Interface requirements
4.5.1 | Voltage stabilizers (if applicable)
| Any appliance designed for use with AC mains power shall be provided with either an integrated or a stand-alone voltage stabilizer or protection. The voltage stabilizer or protection must be certified by WHO PQS as complying with WHO/PQS/E007/VS01. The design of integrated voltage stabilizers must allow for easy access and repair or replacement by technicians in the event of malfunction. Note that: although DC-powered appliance may still benefit from similar electrical protection, only AC-powered appliances are required to have voltage regulation in full compliance with WHO/PQS/E007/VS01. | [Please give details of the stand-alone voltage stabilizer or protection and it’s compliance to this clause.]
4.5.2 | Solar PV interface (if applicable)
| For appliances using solar PV, all electrical components shall be compatible with a solar power system as stipulated in specification WHO/PQS/E003/PV01. The appliance is to be equipped with a locking female and male coupler system (“plug and play”) that is compatible with the solar array interconnection cable and polarity. If the appliance is intended to be powered by (and/or marketed as powered by) solar power, it shall be supplied with a dedicated, applicable, solar PV system meeting all requirements in WHO/PQS/E003/PV01. | [Please give details of the solar power system and it’s compliance to this clause.]
| The design of the appliance and related system components should account for repeated connection and disconnection by the user to transport the appliance. Disconnecting from a solar PV system or other electricity source while under electrical load using the locking coupler system is generally not recommended for this repeated action. The legal manufacturer shall design the appliance and system to clearly lead the user to connect and disconnect the transportable appliance from the PV system in a safe and robust way. | [Please give details of the connection and disconnection e.g. refer to the user manual & section and how this complies with this clause.]
4.5.3 | Power switch
| The appliance shall be equipped with a mechanism specifically designed to turn all active cooling on and off (e.g. power switch or button) that is readily accessible to the user but protected from accidental activation during transport or other conditions. If the appliance system includes stationary components with an on and off mechanism, there shall also be an additional mechanism included on the non-stationary components that is readily accessible to the user but protected from accidental activation during transport or other conditions. Cable or cord connectors (i.e. locking female and male coupler system) are not acceptable for switching active cooling components or the appliance itself on and off. | [Please give details of the power switch e.g. refer to the user manual section, and how this complies with this clause.]
4.5.4 | Power lead
| If the appliance uses external electrical power, then it shall be supplied with electrical power lead(s) with sealed-on plug(s) that are compatible with the electricity socket standard for the intended country of use and intended power source(s). The plug should be break-proof if dropped. If the appliance is intended to be charged or operated by using a 12-volt “cigarette lighter” plug, it is recommended that the connector seat firmly in the 12-volt port and include a latching or locking mechanism to avoid accidental disconnection or loosening. The connector may also include an LED or other indicator that the plug is correctly electrically connected in the port. | [Please give details of the power lead e.g. refer to the user manual section, and how this complies with this clause.]
4.5.5 | Alerts and alarms
| The appliance shall include an alert system. At a minimum, the alert shall include upper and lower settings as noted in WHO/PQS/E006/TH06 for vaccine refrigerators: | [Please summarize the type of alarm and visual warning e.g refer to the user manual & section.]
| a. Low alarm setting: Exposure to a single temperature event of - 0.5°C or below for 60 minutes. | [Please refer to a document & section that show compliance to this clause.]
| b. High alarm setting: Exposure to a single temperature event of +10°C or above for 10 hours. | [Please refer to a document & section that show compliance to this clause.]
| Note that the minimum sound requirement of 100 dB(A) for optional audible alarms required in WHO/PQS/E006/TH06 is not required for these portable appliances which may be located very close to users’ ears during use. A maximum of 80 dB(A) is recommended. The legal manufacturer should consider this proximity if designing optional, audible alarm systems. It is recommended that both an audible and visual alarm are present and that the alerts should be triggered by the same temperature measurement information that is visible to the user on the display required in Section 4.2.10 of this specification, in order to avoid apparent discrepancies to the user. | [Please state the type of alarm and the sound level in dB(A) with visual warning e.g refer to the user manual & section.]
| Since appliances are not expected to run and maintain temperatures continuously, provision shall be made for the user to silence or disable the alarms while the appliance is not in use. While in use, the legal manufacturer should consider the implication of simpler vs more difficult ways to silence alarms or alerts. It is recommended that the provision for completely disabling alarms is clearly labelled and either located or programmed in a way that will make it unlikely to be changed inadvertently. | [Please give details of how to enable this function e.g refer to the user manual & section.]
| Incorporation of an additional method to alert the user of temperature exceeding +8°C as an early warning is preferred but not required in order to help reduce the frequency of larger excursions that could be caused by users accidentally leaving the appliance door or lid open. Similarly, a door or lid open alert or alarm is also recommended. | [Please give details of this option e.g. refer to the user manual & section.]
| Additionally, if the appliance is primarily powered by an electrical battery, the appliance shall have a charge status indicator. This indicator should, at a minimum, give the user some information regarding the amount of remaining energy in the battery. Indications of remaining cooling ability or time may also be included, but the manufacture should consider that due to the circumstantial nature of the estimation, this type of indication may prove difficult to implement accurately. | [Please give details of the battery charge indicator e.g. refer to the user manual & section.]
4.5.6 | Dimensional compatibility with vaccine packaging
| Lightweight appliances will generally be used to transport mixed loads of vaccine in individual primary containers. The net dimensions of the vaccine storage compartment should accommodate all types of pre-filled vaccine presentation and the complete range of standard vaccine vials and ampoules up to 50 dose size as well as tubes and pre-filled devices. | [Please give details, e.g. refer to the user manual section with diagram and/or photos of options to pack vaccines.]
| Medium-weight and heavy-weight appliances may additionally be used to transport secondary cartons (see Annex 4). The net dimensions of the vaccine storage compartment (length, breadth, and height) should accommodate the widest possible range of routine vaccines supplied in these forms of packaging. Secondary cartons may be oriented in all three planes – for transport purposes it is not essential that vial caps face upwards. | [Please give details, e.g. refer to the user manual section with diagram and/or photos of options to pack vaccines.]
4.5.7 | Compatibility with distribution method
| Lightweight and medium-weight appliances shall be designed so that they can be securely strapped, upright, to the luggage rack of a bicycle or motorbike. Medium-weight and heavy-weight appliance design should additionally take into consideration secure transport by larger vehicles in truck beds and on larger roof or tail racks. | [Please show details of the design that meets the requirements of this clause.]
4.5.8 | Stationary components
| Some appliances may include stationary components. As these components of the appliance system are, by definition, not intended for transport, these components are not subject to requirements in this specification in sections: | [Please state whether your appliance has any stationary components.]
● 4.2.16 Closure and Lock
● 4.2.19 Shape
● 4.2.21 Carrying components
● 4.2.22 IP rating
● 4.2.23 Robustness and vibration
● 4.5.7 Compatibility with distribution method
4.5.9 | Intended uses
| As identified in the scope, this specification is applicable only for appliances intended for use for temporary storage and transport of vaccines. Use for long-term storage or when powered by systems not intended for the appliance (e.g. health facility solar photovoltaic systems that may not have adequate design capacity) is prohibited. The legal manufacturer shall not design the appliance with the intended use as long-term or permanent storage and shall further warn against such use in both the user manual and in permanent labelling on the appliance. Although specific language is not set for this required labelling, “WARNING: Not for continuous use” or “WARNING: Maximum continuous use = X days” are examples of wording that may be acceptable appliance labels. | [Please show a copy of the permanent warning label on the appliance and refer to the section in the user manual.]
| As an exception, if the legal manufacturer intends the appliance to be used for long-term storage in addition to use as temporary storage or for transportation as specified herein, they must additionally meet all requirements and pass all tests for another applicable, PQS appliance category (e.g. WHO/PQS/E003/RF03 - Refrigerator or combined refrigerator and water-pack freezer: intermittent mains powered, compression cycle, WHO/PQS/E003/RF04 - Refrigerator or combined refrigerator and water-pack freezer: Solar powered with rechargeable battery, etc.), and may be listed in both applicable, verified categories. | [Please state whether your appliance is intended for long-term storage and which other WHO PQS specifications are met.]
| Additionally, If the appliance could be connected to and directly or indirectly powered or charged by external electrical sources, then the legal manufacturer shall further warn against potential depletion of limited power sources in both the user manual and in permanent labelling on the appliance. This has been confirmed as a risk of using electrically powered, transportable appliances within health systems and at health facilities with limited power resources. Although specific language is not required, the warning language should be relevant to the applicable connection possibilities for example: “WARNING: This appliance may deplete or drain vehicle batteries if connected for extended periods of time” or “WARNING: This appliance may deplete or drain power from facility solar (battery) systems that were not intended or designed for this appliance.” | [Please show a copy of the permanent warning label on the appliance and refer to the user manual section.]
4.6 | Human factors
4.6.1 | General design
| The appliance shall be usable by the widest practicable range of active health workers, regardless of age, gender, size, or minor disability, including colour-blind users and long-sighted people without glasses, in accordance with the general usability principles laid out in ISO 20282-1. Lightweight appliances shall be designed to be comfortable for transport by a similar potential range of health workers for several hours when fully loaded. This includes consideration of individuals wearing traditional dress in intended use regions. Backpack units, for example, may be incompatible with some types of clothing. | [Please state, with examples, which general usability principles have been considered in the design of your appliance.]
4.6.2 | Markings and labelling
| Compressors, if present, shall be marked with the blue identifying symbol shown in Annex 2. In addition, the cabinet shall be permanently marked near the compressor position with the chemical name of the refrigerant or with the refrigerant number, formula or proportion (for blended refrigerants). Appliances operating on R600a shall be marked with the warning symbols shown in Annex 2. Appliances not utilizing compressor-based active cooling methods shall be marked identifying any refrigerant and/or heat transfer fluids used including, but not limited to, the chemical name, formula, or proportions (for blended refrigerants). | [if applicable, please state the amount and type of refrigerant and refer to photos of the warning labels on the appliance.]
| All appliances containing hazardous materials must be labelled as such and include a Safety Data Sheet. Labels and Safety Data Sheets shall comply with the Globally Harmonized System for the Classification and Labelling of Chemicals GHS Rev.5. | [If applicable, please state the hazardous materials and refer to the Safety Data Sheet for these. Also refer to the photos of the warning labels.]
| The appliance shall have a non-removable label fixed to the outside, not the bottom, and designed to last the lifetime of the appliance that specifies PQS temperature zone as shown in Annex 1. The appliance shall carry the following additional information similarly permanently fixed to the outside, not the bottom, of the appliance: legal manufacturer and model number, serial number, date of manufacture, PQS identification number, applicable service phone number and website URL and applicable required warning statements as described in Section 4.5.9 of this document. These labels and information shall remain readable for the expected life of the appliance. | [Please refer to a copy of this label and photos as fixed to the appliance.]
4.6.3 | Vaccine storage advice
| All appliances shall carry a factory-fitted, non-removable label designed to last the lifetime of the appliance that specifies the vaccine storage instructions. The instructions should be fixed internally or externally to either the lid of appliances that open at the top or near the top of the door on appliances that open on a side. Instructions should be in one of the languages specified in Section 4.11, as indicated by the purchaser at the time of ordering. If the appliance is graded other than Grade A and removable baskets or other, similar components are supplied, a warning shall be affixed within the appliance instructing users to “Store vaccine in baskets only” or other appropriate instruction in the same language. | [Please refer to a copy of this label and photos as fixed to the appliance.]
4.7 | Materials
4.7.1 | Casing material selection
| Some materials are degraded by exposure to ultraviolet radiation. Additionally, certain materials can release toxic halogen gases or other toxins when exposed to sunlight and ultraviolet radiation. Therefore, casing materials that will potentially be exposed to ambient light during use or storage shall resist ultraviolet degradation caused by long-term exposure as well as the release of halogen gases or other toxins. This applies to external surfaces and joints as well as internal surfaces and joints that may be exposed to ambient light during storage if the appliance is left open (e.g. the vaccine storage compartment). Neither chlorinated plastics, nor composites containing epoxy, polyurethane, phenol formaldehyde or urea formaldehyde resins are permitted for these casing materials. Additionally, all joints between moulded components shall be water and vapour proof, easy to clean and selected with environmentally safe end-of-life decommissioning and recycling in mind. | [Please state the casing materials and how they meet the requirements of this clause.]
4.7.2 | Thermal insulation foaming agents
| Any gas complying with the limitations and deadlines set by the Montreal Protocol on the elimination of ozone-depleting chemicals may be used. Insulation materials with a low global warming potential (GWP) are preferred. If foam insulation materials are used, they shall be warranted to have a Long-Term Thermal Resistance (LTTR) no worse than 7% lower than the rated ‘R’ value at the time of foaming, when subjected to an accelerated aging test in accordance with ASTM C1303 / C1303M-12. | [If applicable, please state the foaming agent(s) and their GWP.]
4.7.3 | Refrigerant
| If refrigerants are used, hydrocarbon (HC) refrigerants such as R600a or other gases with global warming potential (GWP) ≤ 11 and zero ozone depletion potential (ODP) are acceptable. CFC (chlorofluorocarbon) and HCFC (hydrochlorofluorocarbon) gases are not acceptable. The suitability of alternative refrigerants will continue to be assessed and preference will be given to products that use gases with low global warming potential (GWP). | [If applicable, please state the amount and type of refrigerant.]
4.7.4 | Other restricted materials
| The appliance and its constituent components, including batteries, shall not contain lead, mercury, cadmium, hexavalent chromium, polybrominated biphenyls (PBB), or polybrominated diphenyl ethers (PBDE). | [Please declare that these materials are not included.]
4.7.5 | Vacuum panels
| Vacuum panel construction is an acceptable alternative to foam insulation provided the panels are warranted to retain an acceptable level of vacuum for a minimum period of 10 years under the intended conditions of use. The panels shall retain their thermal properties well enough that the appliance continues to operate to all performance specifications in this document over the entire period. | [Please state whether VIPs are used in the design and show evidence for the integrity of these VIPs.]
4.7.6 | PCM
| Integrated thermal buffer materials may be used to prevent freezing temperatures from propagating to the vaccine storage compartment or for other thermal purposes. The buffer material may be PCM-based but, if so, the material and container shall comply with WHO/PQS/E005/PCMC0.1 – Phase change material containers. | [Please explain if there are any thermal buffers and the precautions if these are PCM-based.]
4.7.7 | Corrosion resistance
| The legal manufacturer shall supply evidence of and declare compliance that the internal and external cabinet, lid, frame, hinges, stays, catches or handles and fixings are protected against corrosion as per EN ISO 6270-1 / ASTM D2247 / EN 13523-26, EN ISO 6270-2 / EN 13523-25, ISO 6272 / EN 13523-5 and ISO 2409 as appropriate to the materials used. | [Please refer to evidence of how various components meet the requirements of this clause. e.g. test report(s) or suppliers specification.]
| It should be expected that the casing and other parts of the appliance may be exposed to condensate when operating in a warm, humid environment. Critical components such as thermostatic controls, electronics, wiring and connections, and displays shall not be located such that condensate or other moisture will flow toward them or accumulate on or near them. All such components shall not be vulnerable to moisture leading to potential corrosion or other related faults or failures. The susceptibility of modified or combined materials (e.g. after welding, brazing, bending, cutting, etc.) to corrosion should be considered in appliance design, material selection, and for possible post-process treatments or coatings as applicable to avoid corrosion or other related faults and failures. Because the final use location of appliances will often not be known in advance, design and material selection should include the possibility of installation in and exposure to corrosive, coastal, ambient environments. | [Please explain with reference how your appliance is designed and protected to minimized corrosion.]
4.7.8 | Chemical resistance
| The external and internal surfaces of the appliance shall be resistant to chemicals commonly used for disinfecting (e.g. sodium hypochlorate, 5.25% in water). | [Please refer to material specification or test evidence.]
4.8 | Warranty
| The product is to be covered by a minimum two-year replacement warranty in the event of any component failure arising from defective design, materials or workmanship. The legal manufacturer shall state the time period over which the rated independence, holdover time and performance are assured. | [Please refer to document(s), including user manual, where the warranty details are clearly stated.]
4.9 | Servicing provision
| The product shall be designed, and components selected, with the aim of achieving at least a 10-year life free of repairs or maintenance apart from expected, routine tasks such as de-frosting, cleaning, and replacement of batteries (if any). | [Please state how you ensure a 10-year life span of your appliance.]
4.9.1 | Essential spare parts and user maintenance tools and supplies
| Based on product design and requirements, the type and quantity of spare parts, basic installation tools/supplies and user and technician maintenance manuals (see Section 4.11 Instructions) shall be determined and agreed upon in advance of order placement. As a minimum, each appliance shall be supplied with 10 spare fuses of each fuse size and type used in the appliance. The spare fuses shall be attached within or on the appliance. | [Please refer to the maintenance manual. and provide photos of where spare fuses are stored on your appliance.]
| Legal manufacturers shall publish a list of spare parts recommended for purchases of 10 and 50 appliances and power systems. Legal manufacturers shall ensure supply of spare parts for a minimum of five years from the time of cessation of the last production of equipment. Spare parts (except batteries if present) shall be provided in kit form for storage in appropriate quantities at national or sub-national level in the purchasing country, as agreed with the purchasing agency. | [Please list or refer to a list of spare parts and provide sources for local agents where these will be available.]
4.10 | Disposal and recycling
| The legal manufacturer is to provide information to the buyer on the hazardous materials contained within the system and suggestions for resource recovery/recycling and/or environmentally safe disposal. For legal manufacturers from the European Union WEEE compliance in accordance with European Union Directive 2002/96/EC is mandatory. | [Please refer to the user manual section.]
4.11 | Instructions
Detailed user and maintenance instructions shall be provided in at least one of the UN languages Arabic, English, French, Mandarin Chinese, Russian, or Spanish, or other language by special order. If only one language is included, it must be the language most appropriate to the country of use. Instructions should include easy to understand visuals whenever possible to avoid reliance on text. The instructions must be written for users and repair technicians and need not repeat the information shown on the permanent labels but shall cover the following topics:
| ● description of operations and use; | [Please refer the user manual section.]
| ● instructions for safe transport including the recommended use of tie- downs or any other securing mechanisms; | [Please refer the user manual section.]
| ● temperature adjustments (if applicable); | [Please refer the user manual section.]
| ● detailed guide to all warning lights, other indicators and displays; | [Please refer the user manual section.]
| ● prevention of vaccine freezing (if Grade B or Grade C); | [Please refer the user manual section.]
| ● simple daily, weekly and monthly maintenance tasks; | [Please refer the user manual section.]
| ● periodic preventative maintenance checks; | [Please refer the user manual section.]
| ● diagnostic and repair guidance and procedures for minor damage; | [Please refer the user manual section.]
| ● battery replacement (if applicable); | [Please refer the user manual section.]
| ● itemized list of spare parts including part numbers; | [Please refer the user manual section.]
| ● clear warranty terms and conditions; and | [Please refer the user manual section.]
| ● end-of-life resource recovery and recycling procedures. | [Please refer the user manual section.]
| English versions of all instructions and manuals are required to be supplied at the time of laboratory testing. | [Please refer the user manual (in English).]
| A warning shall be included in the instructions that the device is not approved for use as a permanent replacement for a refrigerator, solar powered or otherwise. | [Please refer the user manual section.]
4.12 | Training
| Training is not required although it is recommended that the legal manufacturer or reseller have the capability to provide remote or in-person training in the countries where their product is deployed for appliances with electronic user interfaces; the ability to operate using multiple power sources; or Grade B or C – UDFP. | [Please explain how training could be possible if required by the client.]
4.13 | Verification
| In accordance with PQS Verification Protocol WHO/PQS/E003/TS01-VP. | [Please refer the test report.]
5. | Packaging
| Legal manufacturers shall be aware that products may be exposed to very high temperatures during shipping and dockside storage and shall take appropriate actions to mitigate associated risks. The packaging is to be a sturdy export quality and of a commercial standard that will provide adequate protection of the goods for carriage by air, sea and/or road to final destinations worldwide, including remote locations under adverse climatic and storage conditions and high humidity. For heavy-weight appliances only, the packaging is to have not less than 17 kN edge crush resistance with minimum 60% remaining with 90% humidity at a temperature of +70°C (tropical conditions). To avoid destructive unpacking prior to arrival at the final destination legal manufacturers are encouraged to add a resealable observation opening in their packaging to access labelling. Instructions on the packaging alerting inspectors to use of the opening and what information will be revealed are also advised. Materials used for packaging the finished appliance shall be free of ozonedepleting compounds as defined in the Montreal Protocol. The general specification of shipping containers will be subject to agreement with the individual procurement agencies. | [Please describe the packaging and refer to the user manual section/page number on how this packaging may be disposed / recycled.]
Each appliance must include a separate user manual and technician installation manual in Arabic, English, French, Mandarin Chinese, Russian and Spanish. An English version of all instructions and manuals are required to be supplied at time of laboratory testing. Instructions must include easy to understand visuals whenever possible to avoid reliance on text. The instructions must be written for users and repair technicians covering the following topics:
• appliance installation procedures
• health and safety procedures, both for the appliance and for the power supply system
• compatible types of voltage stabilizer or equivalent protection systems
• basic operation and temperature adjustment (if applicable)
• vaccine storage instructions
• simple daily, weekly and monthly maintenance tasks
• method for defrosting
• periodic preventative maintenance checks
• diagnostic (trouble shooting) and repair procedures both for the appliance and for the power supply system
• itemized list of spare parts including part numbers
• end-of-life resource recovery and recycling procedures.
Training
Manufacturer and client to agree training for users including:
• operator health and safety and PSE both for the appliance and for the power supply system
• correct installation of the freezer and of the power supply system
• vaccine storage instructions
• thermostat setting (if applicable)
• maintenance of the frozen compartment including daily temperature checks
• cleaning of the appliance
• cleaning of any special components e.g. condenser, solar array
• method for defrosting.
Verification
In accordance with PQS Verification Protocol E003/ULT01-VP.1.