## Sheet1
E003 Vaccine ultra-low temperature freezer compression cycle submission template (V1.1 20 FEB 2025)
| Assessment of appliance: | [Brand & Model here] | For dossier review on [TBC]
| Tested in accordance with PQS verification protocol: | E003/ULT01-VP.1 (22 DEC 2020)
| To verify compliance with PQS specification: | E003/ULT01.1 (22 DEC 2020)
ASSESSOR’S SUMMARY:
Instructions in [square brackets] are for your guidance. Please delete these instructions and replace with evidence or a precise reference (including page and section numbers) showing compliance with that particular PQS specification requirement. In general a "pass" or "complies" statement alone is not acceptable.
One document may provide evidence for several clause requirements. Please carefully reference the sub-sections of any such document so that it is clear that the clause requirement has been met. This will help speed the pre-qualification process.
Greyed-out cells do not require a response.
NOTE: If it is not clear or doubtful whether a requirement has been met, then this is a non-compliance and pre-qualification might be delayed.
|  |  | REVIEWER’S SUMMARY: |  | REVIEWER’S SUMMARY: |  | REVIEWER’S SUMMARY:
Section 7. | Product dossier | Applicant's submission | Reviewer's remarks [Name & Date] | Applicants response [Name & Date] | Reviewer's remarks [Name & Date] | Applicants response [Name & Date] | Reviewer's remarks [Name & Date]
The legal manufacturer or reseller is to provide WHO with a pre-qualification dossier containing the following:
| Dossier examination fee in US dollars. | [Please contact PQS.]
| General information about the legal manufacturer, including name and address. | [Please state or refer to a document.]
| Unique identification reference for the product type. | [Please state.]
| Brand name of the product. | [Please state.]
| Full specifications of the product being offered, covering all the requirements set out in this document, including details of product marking and traceability. | [Please refer to a document.]
| A comprehensive set of photographs showing all external surfaces of the unit, the interior layout, the compressor and a close-up of the thermometer and the control panel. | [Please refer to a document.]
| Certified photocopies of all type-approvals obtained for the product, including CE marking and the like. | [Please refer to a document.]
| An example of a system which could supply continuous power to the appliance marking and the like.(details of this power system may be dependent on location and a site survey. This power supply must be in accordance with PQS specification E003/POW. | [Please refer to a document.]
| Certified photocopies of the legal manufacturer’s ISO 9001 quality system certification. | [Please refer to a document.]
| Where relevant, certified photocopies of the legal manufacturer’s ISO14001 certification, EMAS registration or registration with an equivalent environmental audit scheme. Conformity with an environmental audit scheme is not mandatory however, preference will be given to manufacturers who are able to demonstrate compliance with good environmental practice. | [Please refer to a document.]
| Laboratory test report(s) proving conformity with PQS specifications. | [Please refer to a document.]
| Indicative cost of the product per unit, per 10 units and per 100 units, EXW (Incoterms 2010). | [Please state or refer to a document.]
Section 4. | Requirements
4.1 | General
An ultra-low temperature freezer is set to achieve a specified internal temperature. The precise internal temperature depends on vaccine storage requirements. Manufacturers may offer products which can be set to achieve one or more internal temperatures and which are suitable for one or more climate zones.
4.2 | Performance
4.2.1 | Operating temperature range
| As indicated on the temperature zone rating sticker attached to the product (see Annex 1). | [Please state the temperature zone, and show a photo of this label on the appliance.]
4.2.2 | Refrigeration system
| Cooling-cycle system operating on alternating current or direct current electricity cooling using a direct or indirect system. | [Please specify.]
4.2.3 | Voltage and frequency
| 220-240-volt 50/60 Hz, 100-127-volt 50/60 Hz or low voltage d.c. options must be offered. Performance must be identical for all options, regardless of the nominal voltage and frequency rating of the appliance. | [Please state which voltage and frequency options are available.]
4.2.4 | Power supply
| To ensure a continuous power supply to the appliance and to any air-conditioning essential for good performance of the appliance, a dedicated system in accordance with WHO PQS E003/POW 01.0 must be provided. | [Please specify (with reference to PQS E003/POW 01.0).]
4.2.5 | Space not suitable for vaccine storage
| Internal space outside the correct internal temperatures e.g. space which is too warm, must be excluded from use by design. | [Please indicate with diagrams and/or photos. (Space not suitable for vaccine storage must be clearly indicated.)]
4.2.6 | Temperature control
| The temperature control shall achieve the correct internal temperature for the intended vaccine storage. The vaccine load must remain at, or colder than the freezer temperature +1°C during any continuous ambient temperature test(s) or day/night cycling temperature test(s). Eutectic packs, if supplied, must remain in situ and not be accessible to the user. If more than one setting is possible, the setting tested according to the verification protocol is the freezer temperature that will be prequalified. | [Compliance to this clause must be demonstrated in the laboratory testing. Please make precise reference to these tests.]
4.2.7 | Thermometer
| Externally readable cabinet-mounted electronic thermometer conforming to PQS specification E006/TH06.2 but calibrated at least down to -80°C or the specified freezer temperature. | [Please refer to a photo or diagram of the thermometer and refer to a calibration document of this thermometer.]
4.2.8 | Cool-down time
| No standard set; however, performance data will be published. | [Please state the cool down time and refer to the test that demonstrates this.]
4.2.9 | Holdover time
| No standard set however, performance data will be published. | [Please state the holdover time and refer to the test that demonstrates this.]
4.2.10 | Starting voltage
| 10 out of 10 cold appliance starts at a voltage at least as low as 10% below manufacturers stated voltage must all be successful. | [Please refer to the test evidence that demonstrates compliance with this clause.]
4.2.11 | Power consumption
| No standard set however, results will be reported and listed in the product information sheet in kWh/day. | [Please state typical electricity consumptions and the test references that confirm this.]
4.2.12 | Evaporator configuration
| If the evaporator is mounted in shelves there must be adequate protection to ensure the evaporator cannot be damaged in general use, e.g. from the insertion and removal of storage bins or trays, or from manual defrosting. | [Please state whether this applies and refer to a diagram or photo(s) showing adequate protection to ensure the evaporator cannot be damaged in general use.]
4.2.13 | Corrosion resistance
| The legal manufacturer must declare compliance that all internal and external components are protected against corrosion as appropriate to EN ISO 6270-1 / ASTM D2247 / EN 13523-26, EN ISO 6270-2 / EN 13523-25, ISO 6272 / EN 13523-5 and ISO 2409:2013. It should be expected that the skin and other parts of the appliance will be subject to condensate when operating in a warm humid environment. Critical components such as thermostatic controls and displays must not be vulnerable to any moisture. Evidence for compliance to this clause shall be demonstrated. | [Please refer to a document that shows test evidence for compliance to this clause. i.e. precisely which components have which proven level of corrosion protection.]
4.2.14 | Electrical safety rating
| Electrical safety: The legal manufacturer must declare compliance with IEC 60335-1, IEC 60335-2-24 or equivalent standard as appropriate. Evidence for compliance to this clause shall be demonstrated. | [Please refer to a document that shows evidence for electrical safety compliance.]
4.2.15 | EMC compliance
| The legal manufacturer must declare EMC compliance with the latest edition of IEC 61000-6-1 and IEC 61000-6-3. Evidence for compliance to this clause shall be demonstrated. | [Please refer to a document that shows evidence for EMC compliance.]
4.2.16 | Markings
| Compressors must be marked with the blue identifying symbol shown in Annex 2. In addition, the freezer cabinet must be permanently marked, near the compressor position, with the chemical name of the refrigerant, or with the refrigerant number, amount of refrigerant formula or proportion (for blended refrigerants). | [Please refer to a photo of the warning label on the appliance. Note this is required for all flammable refrigerants.]
4.2.17 | Rating plate
There should be an external rating plate showing the following information:
| • model number or ID | [Please refer to copy of the rating plate showing all the required information and also refer to a photo showing the position of the rating plate on the appliance.]
• serial number
• required voltage and frequency
• minimum VA rating required
• compatible types of voltage stabilizer
• climate zone(s)
• chemical name of the refrigerant, or the refrigerant number
• amount of refrigerant formula or proportion (for blended refrigerants)
• type of insulation
• vaccine storage volume.
4.2.18 | Vaccine storage advice
All units must carry a factory-fitted non-removable label, designed to last the lifetime of the appliance, carrying the following information:
| • vaccine freezers: vaccine storage instructions and the appropriate temperature zone symbol as Annex 1 | [Please refer to copy of these instructions showing all the required information.]
• health and safety procedures
• a notice that PCMs or water packs should not be stored in the appliance at the same time.
| The instructions should be fixed to the lid of chest freezers and near the top of the door on upright freezers. Instructions should be in one of the languages specified in Clause 4.12, as indicated by the purchaser at the time of ordering. | [Please refer to a photo showing the position of the instructions on the appliance.]
4.3 | Environmental requirements
Ambient temperature range during transport, storage and use:
| • -30°C to +70°C when the product is inactivated. | [Please provide evidence of compliance.]
| • -30°C to +50°C when the product is activated in use. | [Please provide evidence of compliance.]
Ambient humidity range during transport, storage and use:
| • 5% to 95% RH, non-condensing when the product is inactivated. | [Please provide evidence of compliance.]
| • 5% to 85% RH, non-condensing when the product is activated. | [Please provide evidence of compliance.]
4.4 | Physical characteristics
4.4.1 | Overall dimensions
| To allow for maneuvring around corners and through corridors and doorways, the minimum dimension of the product (either length, width or height) should not exceed 710 mm; exceptionally a minimum dimension up to 830 mm can be accepted, but this may restrict the number of sites where the appliance can be installed. The maximum dimension should not exceed 1700 mm and the maximum diagonal (corner to corner) dimension must not exceed 1850 mm. All dimensions must be clearly stated on the PQS information sheet. | [Please state the external dimensions or refer to specifications.]
4.4.2 | Weight
| Mechanical lifting equipment will often not be available at the installation sites. It is recommended that the refrigerator and any associated components should be designed for lifting in such a way that no single worker is required to carry more than 25 kg while working on their own, or in a group. Mechanical lifting equipment at the installation sites is necessary when component weight exceed practical weight limits of the workforce. | [Please state how the appliance may be easily lifted or manoeuvred into position on site.]
4.5 | Interface requirements
4.5.1 | Voltage stabilizer compatibility
| The appliance must be supplied with an appropriate device of equivalent performance to those voltage stabilizers that are currently prequalified according to PQS Specification E007/VS01.5. In some cases, this requirement may be met by sub-components installed as part of the continuous power system compliant with PQS Specification E003/POW 01.0. A warning must be affixed to the unit stating the type(s) of voltage stabilizer with the minimum VA rating that may be used and the user’s manual and spare parts list must clearly record this warning (see Clause 4.2.17). The connection between the freezer, the voltage stabilizer and the power supply system must be compatible. | [Please state which voltage stabilizers (prequalified according to PQS Specification E007/VS01.5) will be provided. Please give an example of a continuous power system compliant with PQS Specification E003/POW 01.0.]
4.5.2 | EMS facility
| The appliance must be supplied with a facility to log internal temperatures in accordance with PQS Specification E006/TR03.2 or similar PQS specification so that temperature performance at least for the preceding month can be speedily downloadable via a direct connection or wireless connection. Downloading information via a cloud-based system alone is not acceptable in case an internet connection is not available. | [Please state how temperature performance at least for the preceding month can be downloadable. Please state the applicable PQS specification.]
4.5.3 | Temperature display and monitoring
| It is recommended that the external digital display of the warmest freezer temperature has a resolution of 0.1°C so that any drift away from the correct vaccine storage temperature can be obvious. There shall be acoustic and visual indications for temperature deviations (high and low), door open, condenser high temperature, power failure other failures, and low internal battery. | [Please state the resolution of the display and refer to a document showing experimental evidence for this resolution. Note that resolution is the smallest different in temperature that can be distinguished, not the smallest increment in the display.]
| Alarm system shall be adjustable for warm and cold alarms and powered by a rechargeable battery with a minimum of 48 hours battery back-up for audio and visual alarm system and temperature recording device. There should be a provision to send alarms via SMS. | [Please state the specification and full capabilities of the alarm system.]
| Other optional alarms: power failure, abnormal voltage, sensor error, filter screen check, thermostat failure, condenser clean, high ambient temperature. | [Please state which alarm options are available.]
4.5.4 | Power lead
| The appliance must be supplied with a power lead with a sealed-on plug compatible with the electrical socket standard in the country where the equipment must be installed. The power lead must be at least 1.5 metre and not more than 2.5 metre in length. There should be provision to secure extraneous lengths of the power lead so that it does not cause a trip hazard. | [Please state the length of the power cord.]
4.5.5 | Defrosting
| Freezer may be defrosted manually or have automatic defrosting. | [Please state which type of defrosting and refer to the user manual section for defrosting instructions.]
4.6 | Human factors
4.6.1 | Generally
| The product must be usable by the widest practicable range of active health workers, regardless of age, gender, size or minor disability, including colour blind users and long-sighted people without glasses, in accordance with the general principles laid out in ISO 20282-1:2006. | [Please state how your appliance complies with this clause.]
4.6.2 | Control panel and thermometer
| Controls, thermometer and other visual displays may be positioned on the front of the unit preferably as close to eye level as possible. Alternatively, they may be mounted on top of the unit at a height not exceeding 1.3 metres. If a low-level position is essential, the display must be aligned so that it can easily be read without the user having to squat or kneel down. The on-off and/or defrost switch, if present, should be recessed or otherwise protected so that it is not possible inadvertently to activate it. | [Please refer to photo(s) showing the position of the visual display and the on/off and defrost switches on the appliance.]
4.6.3 | Thermostat settings
| Thermostat settings may be factory set or in the case of multi-temperature appliances, they may be adjustable. A setting may only be adjusted via an appliance code or similar security to ensure that no casual adjustment can be made. | [Please state how the thermostat setting(s) may be adjusted or refer to a section in the user manual.]
4.6.4 | Door opening
| The freezer should have a facility to allow the door to be quickly & effortlessly opened at any time. This could include a heated air vent with ice-cleaning plunger to prevent vacuum formation. Note this applies both to external and internal doors if fitted. It is recommended that an alarm should sound if the door is left open for a longer than usual time. | [Please state how your appliance complies with this clause.]
4.6.5 | Personal protection equipment
| Protective gloves and similar equipment to prevent freezer burn must be provided. There must a warning sign on the door or lid of the freezer to remind users to handle very cold items, internal trays, bins etc. with care. | [Please state what personal protection equipment will be provided with the appliance. Please refer to a photo of the warning sign on the appliance.]
4.7 | Materials and construction
4.7.1 | Refrigerant
| Hydrocarbon (HC) refrigerants such as R600a or other refrigerants with a global warming potential (GWP) ≤ 11 and zero ozone-depletion potential (ODP) are required. The suitability of alternative refrigerant gases will continue to be assessed and preference will be given to appliances that use gases with low GWP. | [Please state type of refrigerant(s) and their GWP.]
4.7.2 | Thermal insulation foaming agents
| Any gas complying with the limitations and deadlines set by the Montreal/Kigali Protocols on the elimination of ozone-depleting chemicals. | [Please state type of foaming agent, it’s compliance with the Montreal Protocol and its GWP.]
4.7.3 | Other restricted materials
| The product and its constituent components, including batteries, must not contain lead, mercury, cadmium, hexavalent chromium, polybrominated biphenyls (PBB) or polybrominated biphenyl ethers (PBDE). | [Please refer to a declaration of compliance with this clause.]
4.7.4 | Freezer body
| The body should be made from plastic-coated steel or stainless-steel (minimum grade: Steel 18/10 AISI 304) or pre-painted galvanized steel or pre-painted galvanized aluminum. | [Please state the level of protection of the freezer body material and refer to the test evidence which demonstrates this.]
4.7.5 | Door and locking
| The door or lid must have heavy duty hinges and must be fitted with a lock. At least two keys must be supplied with every unit. An inner door is optional. | [Please state how your appliance complies with this clause.]
4.7.6 | Door seal
| The door seal should be a silicone-based high-performance gasket with optional frozen relief door frame heating. | [Please state how your appliance complies with this clause.]
4.7.7 | Shelves
| Stainless steel or plastic-coated steel shelves (adjustable) to be designed with a minimum load of 15kg (35 lbs) per shelf. | [Please state how your appliance complies with this clause.]
4.8 | Warranty
| The appliance to be covered by a three-year replacement warranty from receipt of appliance in the event of any component failure arising from defective design, materials or workmanship. | [Please refer to a warranty statement.]
4.9 | Servicing provision
| The product must be designed to achieve a maintenance-free life of not less than 10 years, apart from routine defrosting and cleaning and replacement of batteries (if applicable). | [Please state how the appliance is designed to last at least 10 years.]
4.10 | Spare parts and accessories
| Legal manufacturer must ensure supply of spare parts for a minimum of 10 years from the time of cessation of the last production of appliance. | [Please declare compliance with this clause and how spare parts may be obtained e.g. via local agents.]
The following to be offered as optional accessories:
| • cryo gloves in small, medium and large sizes | [Please state whether cryo gloves can be available.]
| • ice scraper | [Please state whether an ice scraper can be available.]
| • ice clearing brush | [Please state whether an ice cleaning brush can be available.]
| • seismic restraint kit | [Please state whether a seismic restraint kit can be available.]
4.11 | Disposal and recycling
| The manufacturer must provide information to the buyer on the hazardous materials contained within the system and suggestions for resource recovery/recycling and/or environmentally safe disposal. For the European Union WEEE compliance in accordance with European Union Directive 2002/96/EC is mandatory. | [Please refer to a document giving details of any hazardous materials.] [Please refer to information explaining resource recovery/recycling and/or environmentally safe disposal. e.g. a section in the user manual.]
4.12 | Instructions
| Each appliance must include a separate user manual and technician installation manual in Arabic, English, French, Mandarin Chinese, Russian and Spanish. An English version of all instructions and manuals are required to be supplied at time of laboratory testing. Instructions must include easy to understand visuals whenever possible to avoid reliance on text. The instructions must be written for users and repair technicians covering the following topics: | [Please confirm manuals can be provided in requested languages.]
| • appliance installation procedures | [Please refer to a particular section in the installation manual.]
| • health and safety procedures, both for the appliance and for the power supply system | [Please refer to a particular section in the installation manual.]
| • compatible types of voltage stabilizer or equivalent protection systems | [Please refer to a particular section in the installation manual.]
| • basic operation and temperature adjustment (if applicable) | [Please refer to a particular section in the installation manual.]
| • vaccine storage instructions | [Please refer to a particular section in the installation manual.]
| • simple daily, weekly and monthly maintenance tasks | [Please refer to a particular section in the installation manual.]
| • method for defrosting | [Please refer to a particular section in the installation manual.]
| • periodic preventative maintenance checks | [Please refer to a particular section in the installation manual.]
| • diagnostic (trouble shooting) and repair procedures both for the appliance and for the power supply system | [Please refer to a particular section in the installation manual.]
| • itemized list of spare parts including part numbers | [Please refer to a particular section in the installation manual.]
| • end-of-life resource recovery and recycling procedures. | [Please refer to a particular section in the installation manual.]
4.13 | Training
Manufacturer and client to agree training for users including:
| • operator health and safety and PSE both for the appliance and for the power supply system | [Please explain the options for training.]
• correct installation of the freezer and of the power supply system
• vaccine storage instructions
• thermostat setting (if applicable)
• maintenance of the frozen compartment including daily temperature checks
• cleaning of the appliance
• cleaning of any special components e.g. condenser, solar array
• method for defrosting.
4.14 | Verification
| In accordance with PQS Verification Protocol E003/ULT01-VP.1. | [Please refer to PQS Verification Protocol E003/ULT01-VP.1.]
5. | Packaging
| Materials used for packaging the finished product must be free of ozone depleting compounds as defined in the Montreal Protocol. The general specification of shipping containers will be subject to agreement with the individual procurement agencies. | [Please describe the packaging and refer to the user manual section/page number on how this packaging may be disposed / recycled.]