## Sheet1
E004 Submission to PQS for approval: Vaccine cold box |  |  | Brand & Model:
| Tested in accordance with PQS verification protocol: | E004/CB01-VP.3 | Company name:
| To verify compliance with PQS specification: | E004/CB01.3
Main contact:
Instructions in [square brackets] are for your guidance. Please delete these instructions and replace with evidence or a precise reference (including page and section numbers) showing compliance with that particular PQS specification requirement. In general a "pass" or "complies" statement alone is not acceptable.
One document may provide evidence for several clause requirements. Please carefully reference the sub-sections of any such document so that it is clear that the clause requirement has been met. This will help speed the pre-qualification process.
Greyed-out cells do not require a response.
NOTE: If it is not clear or doubtful whether a requirement has been met, then this is a non-compliance and pre-qualification might be delayed.
|  |  | REVIEWER’S SUMMARY: |  | REVIEWER’S SUMMARY: |  | REVIEWER’S SUMMARY:
7 | Product dossier | Applicant's response | Reviewer's remarks [Name & Date] | Applicant's response | Reviewer's remarks [Name & Date] | Applicant's response | Reviewer's remarks [Name & Date]
The legal manufacturer or reseller is to provide WHO with a pre-qualification dossier containing the following:
| Dossier examination fee in US dollars. | [Please arrange payment with PQS.]
| General information about the legal manufacturer, including name and address. | [Please state.]
| Unique identification reference for the product type. | [Please state.]
| Brand name of the product. | [Please state.]
| Full specifications of the product being offered, covering all the requirements set out in this document, including details of product marking and traceability. | [Please refer to documents which show the product specifications and how the PQS specifications are met.]
| A set of digital photographs showing a three-quarter view of the unit with the lid closed, and three-quarter view with the lid fully opened and a top view of the interior with the full compliment of water-packs in place. | [Please refer to a document showing the photos.]
| Certified photocopies of all type-approvals obtained for the product. | [Please refer to certified copies of any type approvals.]
| Certified photocopies of the legal manufacturer’s current ISO 9001 quality system certification. | [Please refer to certified copies of the ISO 9001 certificate.]
| Where relevant, certified photocopies of the legal manufacturer’s ISO 14001:2004 certification, EMAS registration or registration with an equivalent environmental audit scheme. Conformity with an environmental audit scheme is not mandatory; however preference will be given to manufacturers who are able to demonstrate compliance with good environmental practice. | [Please refer to certified copies of the ISO 14001 certificate or some other environmental conformity scheme.]
| Where available, laboratory test report(s) proving conformity with the product specifications. | [Please refer to the test report carried out in accordance with E004/CB01-VP.3.]
| Indicative cost of the product per unit, per 10 units and per 100 units, including water-packs, EXW (Incoterms 2000). | [Please state or refer to a document.]
4 | Requirements
4.1 | General: Short range or long range insulated cold box, with hinged and insulated lid, designed for transporting vaccines.
4.2 | Performance
4.2.1 | Vaccine storage capacity:
| • Short range: 5.0 to 25.0 litres. • Long range: 5.0 to 25.0 litres | [Please state whether your CB is short or long range.]
4.2.2 | Cold life
| • Short range: minimum 48 hours. • Long range: minimum 96 hours. | [Please refer to the test report section that shows the cold life test result.]
4.2.3 | Cool life
| No standard set; however, performance data will be published. | [Please refer to the test report section that shows the cool life test result.]
4.2.4 | Warm life
| No standard set; however, performance data will be published. | [Please refer to the test report section that shows the warm life test result.]
4.2.5 | Shape
| Containers must be substantially square or rectangular in plan and section. Rounded corners are permitted. | [Please refer to photos or drawings showing compliance with this requirement.]
4.2.6 | Design principles:
| The design of the container, including the placement of the packs and of the load, must promote the free circulation of air within the container to ensure minimum temperature stratification. Container design should seek to minimize the weight of water-packs required to meet the cold life requirement | [Please describe / show how the design of the CB meets these requirements.]
4.2.7 | Lid seal
| The lid of the container must be fitted with a captive labyrinth seal which engages with the container walls when the lid is closed so as to minimize cold bridging and maximize structural strength. | [Please describe / show how your product is designed to meet this requirement.]
4.2.8 | Hinges
| Hinges, where fitted, must allow the lid to open beyond 90° to give full access to the interior of the vaccine carrier. Preferably the hinges should be recessed so that they are fully protected against damage during transport and storage. Hinges must be maintenance-free, without need for lubrication and must be secured to the container in a manner which prevents loosening due to vibration. | [Please describe / show how your product is designed to meet this requirement.]
4.2.9 | Lid stay
| The container must be fitted with a stay device designed to prevent damage caused by over-stressing the lid when it is open. The stay device must be designed to prevent it becoming trapped between the lid seal and the body of the container at all points between fully closed and fully opened. The stay must be secured to the container in a manner that prevents loosening due to vibration | [Please provide photos of the lid stay with the lid in a fully-open and in a fully-closed position (if possible).]
4.2.10 | Catches
| The lid must be secured in the closed position by a catch or catches. It must not be possible for a catch to open accidentally once engaged. Catches must be recessed so that they are fully protected against damage during transport and storage. Catches must be maintenance-free, without need for lubrication and must be secured to the container and to the lid in a manner that prevents loosening due to vibration | [Please provide photos of a catch with the lid fully-open and fully-closed position. Describe how the catches are protected against corrosion.]
4.2.11 | Carrying handles
| The body of the container must be fitted with at least two moulded-in or hinged handles designed so that they can be used to lift and carry the container comfortably when it is fully loaded. Hinge mechanisms must restrain the handle in a near-horizontal position when the container is lifted and the handles must automatically drop back into the vertical position when they are released. Preferably the handles should be recessed so that they are fully protected against damage during transport and storage. Handles must be maintenance-free, without need for lubrication and must be secured to the container in a manner which prevents loosening due to vibration. The handle arrangement must not prevent stable stacking of the boxes. | [Please refer to photos and evidence that shows carry-handle quality and compliance with this requirement.]
4.2.12 | Vaccine storage advice and load restraint instructions
| Cold boxes must carry factory-fitted non-removable labels designed to last the lifetime of the appliance. Labels should be in the UN language most appropriate to the country of use (Arabic, English, French, Mandarin Chinese, Russian or Spanish, or other language, by special order) and should carry the following information. • On the outside of the lid, and/or on the front face of the cold box: As Annex 1. • On the inside of the lid: As Annex 2. | [Please refer to photos of the 2 labels to show compliance with Annex 1 and with Annex 2. (For the purpose of this submission, please show labels in English.)]
4.2.13 | Stacking and handling
| The design of the base and lid of the container should include moulded features that allow multiple units of the same model to be stacked on top of one another in a safe and stable manner. The lid of the container should be designed to withstand the simultaneously weight of three loaded containers of the same model (as containers are often stacked when transported). The base of the container must be designed to withstand repeated dragging across hard rough floor surfaces. | [Please show photos (e.g. of the top and bottom of the CB) and describe how your product is designed to meet this requirement.]
4.2.14 | Corrosion resistance
| All metallic components and their fixings must be constructed in stainless steel or a suitable non-ferrous metal. | [Please describe materials of the carrier.]
4.2.15 | Chemical resistance
| The external and internal surfaces of the container must be resistant to chemicals used for disinfecting (e.g. sodium hypochlorate, 5.25% in water). | [Please describe how the surfaces are resistant to disinfecting chemicals as stated in the requirement. NOTE: This could be test evidence or certification from your material supplier.]
4.2.16 | IP rating
| Protection of the container with lid closed and latched not less than IEC 60529: IP55. | [Please refer to the test report(s) showing compliance with this IP requirement.]
4.2.17 | Robustness
| The container must withstand a one meter drop onto each face, edge, and corner at its rated fully-loaded weight. At the end of the test there must be no damage that affects the performance of the product and the lid must still close and latch correctly. | [Please refer to the test report section showing compliance with this requirement.]
4.3 | Environmental requirements:
4.3.1 | Ambient temperature range during transport, storage and use: -30°C to +55°C | [Please refer to evidence that show the materials used are good within this temperature range.]
4.4 | Physical characteristics
4.4.1 | Overall dimensions: In conformity with clause 4.5.1 to 4.5.3. | [Please state the external dimensions.]
4.4.2 | Weight: The weight of the container must not exceed the following figures: • Short range: Maximum loaded weight, inclusive of the recommended number of water filled water-packs: 35.0 kg • Long range: Maximum loaded weight, inclusive of the recommended number of water filled water-packs 50.0 kg. | [Please state the weight of the container.]
4.5 | Interface requirements
4.5.1 | Dimensional compatibility with packs: The internal dimensions of the container must be compatible with any of the three standard types of water-pack specified in E005/IP01. However, it is acceptable for the product to achieve its designated cold life at its designated vaccine storage capacity using only one of these three types. | [Please make reference to photos or drawings showing compliance with this requirement.]
4.5.2 | Dimensional compatibility with vaccine packaging: Cold boxes are generally used to transport vaccine in secondary packaging. The net dimensions of the storage compartment (length, breadth and height, with water-packs in place) should accommodate the widest possible range of standard immunization vaccines in this form of packaging, with the vials arranged upright in the container. | [Please state the dimensions of the vaccine storage compartment make reference to photos or drawings showing example vaccine packaging.]
4.5.3 | Dimensional compatibility with transport mode: Short range cold boxes must be designed so that they can safely be carried upright inside cars, light vans or 4WD vehicles, strapped to the luggage rack of a small motorcycle or moped or carried by a single porter or a pack animal. Long range cold boxes must be designed so that they can safely be carried upright inside cars, light vans, 4WD vehicles or boats, or carried by two porters or a pack animal. | [Please state how the vaccine carrier is designed to be easily strapped, upright, to the luggage rack of a vehicle. NOTE: Photos or drawings can be useful.]
4.6 | Human factors
4.6.1 | Generally: The product must be designed in accordance with the general usability principles laid out in ISO 20282-1: 2006. | [Please describe how the vaccine carrier complies with this requirement.]
4.7 | Materials:
4.7.1 | Casing material selection: Internal and external casing materials and all joints between the moulded components must be water and vapour proof, must resist UV degradation, must be easy to clean and must be selected with environmentally safe end-of-life disposal in mind. Chlorinated plastics and composites containing epoxy resins are not permitted. | [Please describe how the product is designed to meet this requirement. i.e. how the internal and external casing materials and all joints between the moulded components are: - waterproof, - vapour proof, - resist UV degradation, - must be easy to clean, - environmentally safe end-of-life disposal.]
4.7.2 | Thermal insulation foaming agents: Any gas complying with the limitations and deadlines set by the Montreal Protocol on the elimination of ozone-depleting chemicals. Cyclopentane and similar foaming agents with a low global warming potential (GWP) are preferred. | [Please state the thermal insulation foaming agent. its GWP and how it complies with the limitations and deadlines set by the Montreal Protocol on the elimination of ozone-depleting chemicals.]
4.7.3 | PCM: Integrated thermal buffer materials may be used to prevent freezing temperatures from propagating to the vaccine storage compartment or for other thermal purposes. The buffer material may be PCM-based but if so, must comply with WHO/PQS/ E005/PCMC0.1– PCM specification for phase change material containers. | [Please state the PCM used, if any, and refer to a test report showing compliance with WHO/PQS/ E005/PCMC0.1– PCM specification for phase change material containers.]
4.8 | Warranty
| The product is to be covered by a two year replacement warranty in the event of any component failure arising from defective design, materials or workmanship. The manufacturer should state the time period over which the rated cold life is assured. | [Please refer to a document showing the replacement warranty including the time period over which the rated cold life is assured. This document must also explain the terms and conditions.]
4.9 | Servicing
| The product is to be designed to achieve a maintenance-free life of not less than 5 years, apart from routine cleaning. | [Please describe how your product is designed to meet this requirement.]
4.10 | Disposal and recycling
| The manufacturer is to provide information to the buyer on any hazardous materials contained within the product and suggestions for resource recovery/recycling and/or environmentally safe disposal. | [Please refer to the information on any hazardous materials contained within the product and suggestions for resource recovery/recycling and/or environmentally safe disposal. NOTE: this should be a section in the user manual.]
4.11 | Instructions
| User and maintenance instructions in Arabic, English, French, Mandarin Chinese, Russian and Spanish. The instructions need not repeat the information shown on the permanent labels, but must include guidance on how to repair minor damage. | [Please refer to the user manual in English including instructions how to repair minor damage.]
4.12 | Training
| No requirement. Training on prevention of vaccine freeze damage and correct use of ice-packs, cool-packs and warm-packs is the responsibility of the purchaser. | [Please state what training could be available if requested by the purchaser.]
4.13 | Verification
| In accordance with PQS Verification Protocol E004/CB01-VP.3. | [Please refer to the test report carried out in accordance with E004/CB01-VP.3.]
5 | Packaging
| Materials used for packaging the finished product are to be free of ozone-depleting compounds as defined in the Montreal Protocol. The general specification of shipping containers will be subject to agreement with the individual procurement agencies. | [Please describe how the packing meets this requirement and refer to recommendations for correct disposal in the user manual.]