## Sheet1
E004 Submission to PQS for approval: Vaccine cold box |  |  | Brand & Model:
| Tested in accordance with PQS verification protocol: | E004/CB02-VP.1 | Company name:
| To verify compliance with PQS specification: | E004/CB02.1
Main contact:
Instructions in [square brackets] are for your guidance. Please delete these instructions and replace with evidence or a precise reference (including page and section numbers) showing compliance with that particular PQS specification requirement. In general a "pass" or "complies" statement alone is not acceptable.
One document may provide evidence for several clause requirements. Please carefully reference the sub-sections of any such document so that it is clear that the clause requirement has been met. This will help speed the pre-qualification process.
Greyed-out cells do not require a response.
NOTE: If it is not clear or doubtful whether a requirement has been met, then this is a non-compliance and pre-qualification might be delayed.
|  |  | REVIEWER’S SUMMARY: |  | REVIEWER’S SUMMARY: |  | REVIEWER’S SUMMARY:
7 | Product dossier | Applicant's response | Reviewer's remarks [Name & Date] | Applicant's response | Reviewer's remarks [Name & Date] | Applicant's response | Reviewer's remarks [Name & Date]
The legal manufacturer or reseller is to provide WHO with a pre-qualification dossier containing the following:
| Dossier examination fee in US dollars. | [Please arrange payment with PQS.]
| General information about the legal manufacturer, including name and address. | [Please state.]
| Unique identification reference for the product type. | [Please state.]
| Brand name of the product. | [Please state.]
| Full specifications of the product being offered, covering all the requirements set out in this document, including details of product marking and traceability. | [Please refer to documents which show the product specifications and how the PQS specifications are met.]
| A set of digital photographs showing a three-quarter view of the unit with the lid closed, and three-quarter view with the lid fully opened and a top view of the interior with the full compliment of coolant-packs in place. | [Please refer to a document showing the photos.]
| Certified photocopies of all type-approvals obtained for the product. | [Please refer to certified copies of any type approvals.]
| Certified photocopies of the legal manufacturer’s current ISO 9001 quality system certification. | [Please refer to certified copies of the ISO 9001 certificate.]
| Where relevant, certified photocopies of the legal manufacturer’s ISO 14001 certification, EMAS registration or registration with an equivalent environmental audit scheme. Conformity with an environmental audit scheme is not mandatory; however preference will be given to manufacturers who are able to demonstrate compliance with good environmental practice. | [Please refer to certified copies of the ISO 14001 certificate or some other environmental conformity scheme.]
| Where available, laboratory test report(s) proving conformity with the product specifications. | [Please refer to the test report carried out in accordance with E004/CB02-VP.1.]
| Indicative cost of the product per unit, per 10 units and per 100 units, including coolant-packs, EXW (Incoterms 2010). | [Please state or refer to a document.]
4 | Requirements
4.1 | General: Short range or long range insulated large capacity cold box, with an operable door or lid designed for transporting vaccines in bulk.
4.2 | Performance
4.2.1 | Vaccine storage capacity:
| 100 litres or greater. | [Please state storage capacity of your CB.]
4.2.2 | Minimum vaccine storage compartment temperature
| The temperature inside the vaccine storage compartment must remain above 0°C during the cold life test conducted at an ambient temperature of +32°C, and optionally during the day/night +43°C/+32°C cold life test. The temperature must also remain above 0°C during the minimum rated ambient temperature test, when conducted at +15°C (or a lower temperature specified by the manufacturer). | [Please refer to the test report section that shows the minimum vaccine storage compartment temperature test result.]
4.2.3 | Cold life
| • Short range: minimum 24 consecutive hours within the range 0.0°C to +10.0°C. • Long range: minimum 48 consecutive hours within the range 0.0°C to +10.0°C. | [Please refer to the test report section that shows the cold life test result.]
4.2.4 | Minimum rated ambient temperature:
| The container will be challenged to determine an operational minimum rated ambient temperature by testing at an ambient of +15°C, or at a lower temperature if so requested by the manufacturer. | [Please refer to the test report section that shows the minimum rated ambient temperature test result.]
4.2.5 | Cold climate freeze protection life:
| No standard set; however performance data will be published. | [Please refer to the test report section that shows the cold climate freeze protection life test result.]
4.2.6 | Conditions of use
| Large capacity cold boxes may typically be transported on open, curtain-sided or fully enclosed light vans or trucks over road surfaces, including smooth tarmac, heavily perforated tarmac, rough gravel with strong corrugations, wet mud and dirt. They may also be carried in boats. Products must be designed to accommodate these conditions and the design must take account of possible long-term exposure to rainfall and to high levels of direct sunlight. | [Please describe / show how the design of the CB meets these requirements.]
4.2.7 | Design principles:
The design of the container should respect the following principles:
| • Weight of coolant-packs must be the minimum needed to achieve the cold life required. | [Please describe / show how the design of the CB meets these requirements.]
| • Placement of coolant-packs must minimize temperature stratification within the load. | [Please describe / show how the design of the CB meets these requirements.]
| • Cooling of the container may be achieved by one of the following methods: - Fully frozen ice-packs taken directly from a freezer room or freezer at -25°C. Buffer materials integrated into the container may be used to prevent freezing temperatures from propagating to the vaccine storage compartment. | [Please describe the cooling method of the CB.]
| • Coolant packs must not require any form of secondary conditioning – preparation of the coolant must be a one-step operation. | [Please describe / show how the design of the CB meets these requirements.]
| • The design must eliminate the possibility of vaccine packages being exposed to air temperatures reaching 0.0°C, or contacting surfaces with a temperature of 0.0°C or less under any of the specified test conditions. | [Please describe / show how the design of the CB meets these requirements.]
4.2.8 | Reuse cycle with buffering technology
| If a temperature buffering technology is used to prevent freezing, it must be possible to reuse the container safely within 12 hours of removal of the previous charge of coolant-packs at the minimum rated ambient temperature. The design of the container and instructions for use must ensure that the unit cannot inadvertently be re-used until the buffering system is in the correct thermodynamic state to receive a new coolant charge | [Please state if a temperature buffering technology is used, and is so that it meets these requirements.]
4.2.9 | Shape
| The outside of the container and the vaccine storage compartment must be substantially square or rectangular in plan and section. Rounded external corners are preferred. | [Please refer to photos or drawings showing compliance with this requirement.]
4.2.10 | Door or lid seal
| The door or lid of the container must be must be fitted with an effective seal which engages with the container walls when the door or lid is closed so as to minimize cold bridging and maximize structural strength and resistance to transport vibrations and rough handling. | [Please describe / show how your product is designed to meet this requirement.]
4.2.11 | Hinges
| The door or lid must be fitted with a robustly constructed hinge mechanism which can open beyond 90° to give full access to the interior of the cold box. Hinges must be recessed so that they are fully protected against damage during transport and storage. Hinges must be maintenance-free, without need for lubrication, must be recessed so that they are fully protected against damage during transport and storage and must be secured to the container in a manner which prevents loosening due to vibration. | [Please describe / show how your product is designed to meet this requirement.]
4.2.12 | Door or lid stay
| The container must be fitted with a stay device designed to prevent damage caused by over-stressing the door or lid when it is open. The stay device must be designed to prevent it becoming trapped between the lid or door seal and the body of the container at all points between fully closed and fully opened. The stay must be secured to the container in a manner which prevents loosening due to vibration. | [Please provide photos of the lid stay with the lid in a fully-open and in a fully-closed position (if possible).]
4.2.13 | Catches
| The door or lid must be secured in the closed position by a catch or catches and it must not be possible for a catch to open accidentally once engaged. Catches must be recessed so that they are fully protected against damage during transport and storage. Catches must be maintenance-free, without need for lubrication and must be secured to the container and to the door or lid in a manner which prevents loosening due to vibration. | [Please provide photos of a catch with the lid fully-open and fully-closed position. Describe how the catches are protected against corrosion.]
4.2.14 | Manoeuvring handles
| The body of the container must be fitted with at least two moulded-in or hinged handles designed so that they can be used to move the container when it is empty or to guide the container comfortably when it is fully loaded and mounted on castors or wheels. Hinge mechanisms must restrain the handle in a near-horizontal position when the container is being moved or guided and the handles must automatically drop back into the vertical position when they are released. The handles should be recessed so that they are fully protected against damage during transport and storage. Handles must be maintenance-free, without need for lubrication and must be secured to the container in a manner which prevents loosening due to vibration. The handle arrangement must not prevent stable stacking of the containers or prevent close packing when the containers are placed side by side. | [Please refer to photos and evidence that shows handle quality and compliance with this requirement.]
4.2.15 | Vaccine storage advice and load restraint instructions
| Vaccine storage advice and load restraint instructions Cold boxes must carry factory-fitted non-removable labels designed to last the lifetime of the container. • On the outside of the door or lid, and/or on the front face of the cold box: ‘Stop!’ label and the appropriate temperature zone symbol as Annex 1. In addition there must be clear instructions on the correct use of the load restraint equipment supplied with the product. These instructions should be predominately graphical. • On the inside of the door or lid: Coolant and vaccine loading instructions as Annex 2. This label should be in the UN language most appropriate to the country of use (Arabic, English, French, Mandarin Chinese, Russian or Spanish, or other language, by special order). | [Please refer to photos of the 2 labels to show compliance with Annex 1 and with Annex 2. (For the purpose of this submission, please show labels in English.)]
4.2.16 | Stacking and handling
| The container should include moulded features that allow for safe lifting and moving by pallet trucks and pallet lifting machinery. Non-wheeled or removable wheel containers of the same model must be designed to allow for stacking on top of one another in a safe and stable manner. Integrated pallet bases must be designed without bottom stringers to allow for four sided lifting. The base of non-wheeled containers must be designed to withstand repeated dragging across hard rough floor surfaces. | [Please show photos (e.g. of the top and bottom of the CB) and describe how your product is designed to meet this requirement.]
4.2.17 | Optional castors or wheels
| The product may be supplied with integrated castors or wheels, not less that 100mm diameter, complying with ISO 22883, that are suitable for moving a fully loaded container over rough concrete or other hard horizontal or near-horizontal surfaces. The wheels must all be fitted with a wheel locking/braking device. Preferably there should be a central locking/braking device. Alternatively an optional wheeled chassis may be offered, fitted with pneumatic or puncture-resistant tyres not less than 200mm diameter, and suitable for moving a loaded container over gravel or dirt surfaces with inclines up to five degrees. The wheels must be fitted with a hand operated braking mechanism that enables the operator to achieve full control over a container at its maximum loaded weight. | [Please state if castors or wheels are provided, and if so, how they meet these requirements.]
4.2.18 | Corrosion resistance
| Hinges, stays, catches or handles and fixings, if metallic, must be constructed in stainless steel, aluminium, or other suitable non-ferrous metal. Casings and linings, if metallic, must be constructed in stainless steel, aluminium, or using zinc coated steel sheet to EN 10152 with a corrosion-resistant plastics coating to EN 10169-1. | [Please describe materials of the carrier.]
4.2.19 | Chemical resistance
| The external and internal surfaces of the container must be resistant to chemicals used for disinfecting (e.g. sodium hypochlorate, 5.25% in water). | [Please describe how the surfaces are resistant to disinfecting chemicals as stated in the requirement. NOTE: This could be test evidence or certification from your material supplier.]
4.2.20 | IP rating
| Protection of the container with lid closed and latched not less than IEC 60529: IP55. | [Please refer to the test report(s) showing compliance with this IP requirement.]
4.2.21 | Robustness
| The container, fully loaded with vaccine, must pass ISTA Procedure 3H and the ASTM D4169-09: Schedule F Loose Load Vibration test without physical damage to the contents. At the end of these tests there must be no damage that affects the performance of the container and the lid or door must still close and latch correctly. | [Please refer to the test report section showing compliance with this requirement.]
4.2.22 | Load restraint equipment
| The container must be provided with a minimum of one load restraint attachment point on each exterior side wall and must be supplied with a full set of compatible ratchet tie-down straps to EN 12195-2. The complete assembly must be designed to prevent the container at its maximum loaded weight from breaking free when attached to the specified number of vehicle lashing points tested in accordance with ISO 27956. | [Please describe / show how your product is designed to meet this requirement.]
4.2.23 | Coolant-packs
| The container must be provided with two full sets of coolant-packs with complete installation and preparation instructions. Ice-packs must comply with E005/IP01, except for their physical dimensions. The container must achieve its designated cold life at its designated vaccine storage capacity using ice-packs. | [Please describe how your product meets this requirement.]
4.2.24 | Coolant-pack restraint system
| The container must be designed so that coolant-packs are physically held in position without need for support by the load. If coolant-packs are required on top of the load, these packs must be located in a tray or some other support arrangement which is hinged or otherwise attached to the container body. | [Please describe / show how your product is designed to meet this requirement.]
4.2.25 | Lock
| The door or lid must be fitted with a lock. Two keys are to be supplied with every unit. | [Please describe / show how your product is designed to meet this requirement.]
4.2.26 | Thermometer (optional)
| Integrated externally readable photovoltaic-powered digital thermometer complying with PQS specification E006/TH06.2 Type C. The instrument must be recessed into the casing or otherwise protected against mechanical damage. The sensor must be located in the coldest part of the container. | [Please state if product has an optional thermometer, and if it meets these requirements.]
4.3 | Environmental requirements:
4.3.1 | Ambient temperature range during transport, storage and use: -30°C to +55°C | [Please refer to evidence that show the materials used are good within this temperature range.]
4.4 | Physical characteristics
4.4.1 | Overall dimensions: In conformity with clause 4.5.1 to 4.5.3. | [Please state the external dimensions.]
4.4.2 | Weight: No specific standard is set for the weight of different sizes of container. However the maximum loaded weight of both short and long range containers must not exceed 1,000 kg. | [Please state the weight of the container.]
4.5 | Interface requirements
4.5.1 | Dimensional compatibility with vaccine packaging: Containers will be used to transport vaccine in secondary or tertiary cartons. The net dimensions of the storage compartment (length, breadth and height, with coolant-packs in place) should accommodate the widest possible range of routine vaccines supplied in this form of packaging, with the cartons arranged correct side upwards in the container. Refer to Annex 3 for data on carton sizes. | [Please state the dimensions of the vaccine storage compartment make reference to photos or drawings showing example vaccine packaging.]
4.5.2 | Dimensional compatibility with pallets: Container plan dimensions must follow one of the following standard pallet formats: • ISO: 1.219 x 1.016 metres • ISO, EUR 2 or 3: 1.2 x 1.0 metres • EUR pool: 1.2 x 0.8 metres • EUR 6: 0.8 metres x 0.6 metres Front opening containers should not exceed 1.8 metres in height. Top opening containers should not exceed 0.9 metres in height. | [Please state the container plan dimensions, make reference to photos or drawings showing example container packaging.]
4.6 | Human factors
4.6.1 | Generally: The product must be designed in accordance with the general usability principles laid out in ISO 20282-1. | [Please describe how the vaccine carrier complies with this requirement.]
4.7 | Materials:
4.7.1 | Casing material selection: Internal and external casing materials must resist UV degradation caused by long-term exposure to sunlight. All joints or liners between the moulded components must be water and vapour proof, must be easy to clean and must be selected with environmentally safe end-of-life disposal in mind. Chlorinated plastics and composites containing epoxy, polyurethane, phenol formaldehyde and urea formaldehyde resins are not permitted. | [Please describe how the product is designed to meet this requirement. i.e. how the internal and external casing materials and all joints between the moulded components are: - waterproof, - vapour proof, - resist UV degradation, - must be easy to clean, - environmentally safe end-of-life disposal.]
4.7.2 | Thermal insulation foaming agents: Any gas complying with the limitations and deadlines set by the Montreal Protocol on the elimination of ozone-depleting chemicals. Insulation materials with a low global warming potential (GWP) are preferred. Foam insulation materials must prevent any loss of foaming agent during the design life of the product in all cases where the foaming agent contributes to thermal performance. | [Please state the thermal insulation foaming agent. its GWP and how it complies with the limitations and deadlines set by the Montreal Protocol on the elimination of ozone-depleting chemicals.]
4.7.3 | PCM: Integrated thermal buffer materials may be used to prevent freezing temperatures from propagating to the vaccine storage compartment or for other thermal purposes. The buffer material may be PCM-based but if so, must comply with WHO/PQS/E005/PCMC0.1– PCM specification for phase-change material containers. | [Please state the PCM used, if any, and refer to a test report showing compliance with WHO/PQS/ E005/PCMC0.1– PCM specification for phase change material containers.]
4.8 | Warranty
| The product is to be covered by a two year replacement warranty in the event of any component failure arising from defective design, materials or workmanship. The manufacturer should state the time period over which the rated cold life is assured. | [Please refer to a document showing the replacement warranty including the time period over which the rated cold life is assured. This document must also explain the terms and conditions.]
4.9 | Servicing
| The product is to be designed to achieve a maintenance-free life of not less than 5 years, apart from routine cleaning. | [Please describe how your product is designed to meet this requirement.]
4.10 | Disposal and recycling
| The manufacturer is to provide information to the buyer on any hazardous materials contained within the product and suggestions for resource recovery/recycling and/or environmentally safe disposal. | [Please refer to the information on any hazardous materials contained within the product and suggestions for resource recovery/recycling and/or environmentally safe disposal. NOTE: this should be a section in the user manual.]
4.11 | Instructions
| User and maintenance instructions in Arabic, English, French, Mandarin Chinese, Russian and Spanish. The instructions need not repeat the information shown on the permanent labels, but must include guidance on how to repair minor damage. Instructions must include standard operating procedures for coolant recharging including ice-pack preparation and placement instructions and conditioning of buffering materials before loading. Instructions must be complete and clear, preferably using illustrations only. The instructions may optionally include a section describing alternative coolant-pack combinations as indicated in clause 4.2.7. | [Please refer to the user manual in English including instructions how to repair minor damage, procedures for coolant recharging, and other instructions.]
4.12 | Training
| No requirement. Training on prevention of vaccine freeze damage and correct use of ice-packs, cool-packs and warm-packs is the responsibility of the purchaser. | [Please state what training could be available if requested by the purchaser.]
4.13 | Verification
| In accordance with PQS Verification Protocol E004/CB02-VP.1. | [Please refer to the test report carried out in accordance with E004/CB02-VP.1.]
5 | Packaging
| Materials used for packaging the finished product are to be free of ozone-depleting compounds as defined in the Montreal Protocol. The general specification of shipping containers will be subject to agreement with the individual procurement agencies. | [Please describe how the packing meets this requirement and refer to recommendations for correct disposal in the user manual.]