## Sheet1
E004 Submission to PQS for approval: Vaccine Cold Box - long-term storage (35 days) |  |  | Brand & Model:
| Tested in accordance with PQS verification protocol: | E004/CB03-VP.1 | Company name:
| To verify compliance with PQS specification: | E004/CB03.1
Main contact:
Instructions in [square brackets] are for your guidance. Please delete these instructions and replace with evidence or a precise reference (including page and section numbers) showing compliance with that particular PQS specification requirement. In general a "pass" or "complies" statement alone is not acceptable.
One document may provide evidence for several clause requirements. Please carefully reference the sub-sections of any such document so that it is clear that the clause requirement has been met. This will help speed the pre-qualification process.
Greyed-out cells do not require a response.
NOTE: If it is not clear or doubtful whether a requirement has been met, then this is a non-compliance and pre-qualification might be delayed.
|  |  | REVIEWER’S SUMMARY: |  | REVIEWER’S SUMMARY: |  | REVIEWER’S SUMMARY:
7. | Product dossier | Applicant's submission [Name & Date] | Reviewer's remarks [Name & Date] | Applicant's response [Name & Date] | Reviewer's remarks [Name & Date] | Applicant's response [Name & Date] | Reviewer's remarks [Name & Date]
The legal manufacturer or reseller is to provide WHO with a pre-qualification dossier containing the following:
| Dossier examination fee in US dollars. | [Please arrange payment with PQS.]
| General information about the legal manufacturer, including name and address. | [Please state.]
| Unique identification reference for the product type. | [Please state.]
| Brand name of the product. | [Please state.]
| Full specifications of the product being offered, covering all the requirements set out in this document, including details of product marking and traceability. | [Please refer to documents which show the product specifications and how the PQS specifications are met.]
| A set of digital photographs showing a three-quarter view of the unit with the entry closed, and three-quarter view with the door or lid fully opened and a top view of the interior with the full complement of coolant-packs in place. | [Please refer to a document showing the photos.]
| Certified photocopies of all type-approvals obtained for the product. | [Please refer to certified copies of all type approvals.]
| Certified photocopies of the legal manufacturer’s current ISO 9001 quality system certification. | [Please refer to certified copies of the ISO 9001 certificate.]
| Where relevant, certified photocopies of the legal manufacturer’s ISO 14001 certification, EMAS registration or registration with an equivalent environmental audit scheme. Conformity with an environmental audit scheme is not mandatory; however preference will be given to manufacturers who are able to demonstrate compliance with good environmental practice. | [Please refer to certified copies of the ISO 14001 certificate or some other environmental conformity scheme.]
| Where available, laboratory test report(s) proving conformity with the product specifications. | [Please refer to laboratory test report(s).]
| Indicative cost of the product per unit, per 10 units and per 100 units, including coolant-packs, EXW (Incoterms 2010). | [Please state or refer to a document.]
1. | Scope
| These containers are typically used to maintain the cold chain when vaccines are stored for extended periods at a fixed location, such as a health facility, which provides immunization services and outreach efforts. Optionally, the units may also be used for transport. | (For information.)
| Acceptable cold life must be provided for a minimum of 10 days with one charge of coolant at cycling day/night temperatures of either hot zone temperatures (+43°C/+25°C) and/or at cycling temperate zone temperatures (+32°C/+15°C). | [Please state whether the CB is for hot zone or temperate zone pre-qualification.]
| Long-term storage cold box products may prequalify as either: Type 1: A stationary cold box; or Type 2: A combination transportable and stationary cold box. | [Please state whether the CB is for Type 1 or Type 2 pre-qualification.]
4. | Requirements
4.1 | General
This equipment standard is intended for long-term storage cold boxes that are designed for storing vaccine in a fixed, stationary location and which provide acceptable cold life and temperature control for a minimum period of 35 days. Optionally these cold boxes may be designed for mobile use as a combined transportable and stationary cold box.
4.2 | Performance
4.2.1 | Vaccine storage capacity:
| Minimum 5 litres net vaccine storage capacity when configured in such a way to achieve the cold life required by this specification. Manufacturers of products in which vaccine is designed to be stored in primary containers (vials, ampoules and pre-filled devices), must demonstrate that the container can store the equivalent of 2.5 litres of mixed vaccines, where: ∑ ((PI) x primary container radius squared x primary container height) >= 2.5 litres. | [Please state the vaccine storage capacity. Please show (photos or diagrams) how the container can store at least 2.5 litres of mixed vaccines.]
4.2.2 | Minimum vaccine storage compartment temperature
| The temperature inside the vaccine storage compartment must remain above 0°C during the cold life test conducted at an ambient temperature of +43°C, or +32°C, and during the minimum rated ambient temperature test at conducted at +15°C (or a lower temperature specified by the manufacturer). | [Please refer to the test report section that shows the minimum rated ambient temperature result.]
4.2.3 | Cold life
| Minimum 35 days over a repeating day/night ambient temperature cycle of +43°C/+25°C (Hot zone) or +32°C/+15°C (Temperate zone). | [Please refer to the test report section that shows the cold life test result.]
4.2.4 | Minimum rated ambient temperature
| The containers will be challenged to determine an operational minimum rated ambient temperature by testing at a constant ambient temperature of +15°C, or at a lower temperature if so requested by the manufacturer. | [Please state the minimum rated ambient temperature and refer to the test report section that shows this result.]
4.2.5 | Cold climate freeze protection
| No standard set; however performance data will be published. | [Please state the cold life time and refer to the test report section that shows this result.]
4.2.6 | Design principles. The design of the container should respect the following principles:
| • Weight of coolant-packs must be the minimum needed to achieve the cold life required. | [Please state the amount of coolant packs required.]
| • Products must be designed to be resupplied with coolant at a maximum interval of not less than 35 days. The manufacturer must specify the standard operating procedures and equipment requirements for coolant recharging. The coolant-pack design must take account of the need to transport coolant from a remote location and may need to include a suitable transport container for carrying the coolant so that cold life is not adversely affected. | [Please specify the standard operating procedures and equipment requirements for coolant recharging.]
The required cold life must be achieved under the following use cases:
| • The container is periodically re-stocked with cool vaccine at +2°C to +8°C. Typically this will be at one week intervals. Re-stocking may or may not coincide with coolant recharging. • Pre-filled devices, single-dose vials and ampoules are removed for use during daily immunisation sessions. • Multi-dose vials are removed for use during daily immunisation sessions and partially used vials are returned to the container at the end of each session in accordance with the WHO Multi-dose Vial Policy. • During daily immunization sessions, the container must be able to supply sufficient coolant capacity to maintain reconstituted and other vaccines at a temperature of greater than 0°C and equal or less than +8°C for a period of up to six hours at the point of use. Alternatively the container must be designed for deployment at the point of use so that vaccine opened for an immunization session can be kept at a temperature of greater than 0°C and equal or less than +8°C at the point of use without need for additional coolant. | (For information.)
| • Placement of coolant-packs must minimize temperature stratification within the load. | [Please state or show how coolant-packs placement minimizes temperature stratification.]
| • Cooling of the container may be achieved by one of the following methods: - Fully frozen ice and/or ice-packs taken directly from a freezer room or freezer at a minimum of -25°C. - Conditioned ice-packs equipped with indicators to alert health workers when proper conditioning has occurred. | [Please state the recommended ice-pack temperature to achieve test performance results.]
| • The design must eliminate the possibility of vaccine packages being exposed to air temperatures 0°C or less or contacting surfaces with a temperature of 0°C or less under any of the specified test conditions. | [Please describe or show how the design eliminates vaccine packages being exposed to air temperatures or surface temperatures at or colder than 0°C ].
| • Temperature buffering materials integrated into the container may be used to prevent freezing temperatures from propagating to the vaccine storage compartment. | [Please refer to a photo or diagram which shows any temperature buffering materials.]
| • All pre-qualification testing is carried out with a full complement of ice-packs or cool-packs. However, where such arrangements have the potential to extend the ambient operating temperature range of the product, manufacturers are free to describe alternative arrangements which comprise a mix of frozen and unfrozen coolant-packs of the same type. | [Please refer to a diagram or a photo which shows whether an arrangement of a mix of frozen and unfrozen coolant-packs (of the same type) extends the operating temperature range.]
| • Generic PQS prequalified water-packs must comply with E005/IP01. Other coolant-packs may be either plastic or corrosion-resistant metal and must comply with the standards in this document, including materials, filling requirements, deformation, robustness and leakage. | [Please show the PQS pre-qualification reference number for the water packs used in the carry box.]
| • All plastic coolant-packs, water-packs and pack filling-cap materials must resist UV degradation and must be selected with environmentally safe end-of-life disposal in mind. The exterior must be easy to clean. Manufacturers must use materials that are known to be non-toxic when incinerated at any temperature between 650°C and 1,200°C. Chlorinated plastics and composites containing epoxy resins are not permitted. | [Please show evidence for how the water pack materials: 1. resist UV degradation 2. can be disposed safely 3. can be cleaned 4. and their properties when incinerated at any temperature between 650°C and 1,200°C.]
4.2.7 | Coolant-pack filling requirements
| Coolant-packs that require the user to add coolant must be delivered empty and supplied with a removable filling cap. It must be possible to check the coolant level inside with the cap in place. Filling must be designed: • With the recommended level for filling the coolant-pack clearly visible on the outside of the coolant-pack; or • In such a way that the coolant-pack cannot be over-filled. | [Please describe or show how the coolant packs can be filled with the correct amount of coolant.]
4.2.8 | Coolant-pack deformation
| Deformation caused by expansion must not impact placement in the container and must be reversible. | [Please refer to photos showing expansion does not impeded placement or removal of coolant packs.]
4.2.9 | Coolant-pack robustness
| Coolant-packs must be able to withstand a 2 metre drop onto every face, edge, and corner when frozen to -25°C. After thawing they must then pass a leakage test. Coolant-packs must also be able to withstand a two metre drop onto every face, edge, and corner with the contents in the liquid state, at +10°C. They must then pass a leakage test. | [Please refer to a test report section that shows compliance to this requirement.]
4.2.10 | Coolant-pack leakage
| Unfrozen filled coolant-packs, including the cap must be able to resist a lateral force of 80 kg applied to top, bottom and sides without leaking. | [Please refer to a test report section that shows compliance to this requirement.]
4.2.11 | Door or lid seal
| The door or lid of Type 1 containers must be designed to minimize cold bridging. The door or lid of Type 2 containers must be fitted with an effective seal which engages with the container walls when the door or lid is closed so as to minimize cold bridging and maximize structural strength and resistance to transport vibrations and rough handling. | [Type 1: Please refer to photos or diagrams which show minimal thermal bridging in the design of the cold box. Type 2: Please refer to photos or diagrams which shows an effective door/lid seal so as to minimize cold bridging and maximize structural strength with resistance to transport vibrations and rough handling]
4.2.12 | Hinges
| If hinges are used, the door or lid of the container must be fitted with a robustly constructed hinge mechanism. The lid must open beyond 90° to give full access to the interior of the cold box. Hinges must be maintenance-free, without need for lubrication. The hinges of Type 2 containers must be recessed so that they are fully protected against damage during transport and storage and must be secured to the container in a manner which prevents loosening due to vibration. | [Please refer to evidence that shows hinge quality and compliance with this requirement.]
4.2.13 | Lid stay
| If hinged lids are used, the container must be fitted with a stay device designed to prevent damage caused by over-stressing the lid when open. The stay device must be designed to prevent it from becoming trapped between the lid seal and the body of the container when the lid is closed. Type 2 lid stays must be secured to the container in a manner which prevents loosening due to vibration. | [Please refer to evidence that shows lid-stay quality and compliance with this requirement.]
4.2.14 | Catches
| The door or lid of the container must be secured in the closed position by a catch or catches and it must not be possible for a catch to open accidentally once engaged. Catches must be maintenance-free, without need for lubrication. Type 2 catches must be recessed so that they are fully protected against damage during transport and storage and must be secured to the container in a manner which prevents loosening due to vibration. | [Please refer to evidence that shows catch quality and compliance with this requirement.]
4.2.15 | Carrying handles (Type 2 only)
| The body of the container must be fitted with at least two moulded-in or hinged handles designed so that they can be used to lift and carry the container comfortably when it is fully loaded. Hinged handles must have a mechanism that restrains the handles in a near-horizontal position when the container is lifted and the handles must automatically drop back into the vertical position when they are released. Preferably handles should be recessed or so positioned that they are fully protected against damage during transport and storage. Handles must be maintenance-free, without need for lubrication and must be secured to the container in a manner which prevents loosening due to vibration. The handle arrangement must not prevent stable stacking of the boxes. | [Please refer to evidence that shows carry-handle quality and compliance with this requirement.]
4.2.16 | Vaccine storage advice
| Cold boxes must carry factory-fitted non-removable labels designed to last the lifetime of the product. Labels should be in the UN language most appropriate to the country of use (Arabic, English, French, Mandarin Chinese, Russian or Spanish, or other language, by special order) and should carry the following information: • On the outside of the lid or door, and/or on the front face of the cold box: ‘Stop!’ label and the appropriate temperature zone symbol as Annex 1. • Visible on the outside of the cold box or on the inside of the lid or door: As Annex 2. | [Please refer to photos of both labels (in compliance with both Annexes).]
4.2.17 | Stacking and handling (Type 2 only)
| The design of the base and lid of the container should include moulded features that allow multiple units of the same model to be stacked on top of one another in a safe and stable manner. The lid of the container should be designed to withstand the simultaneously weight of three loaded containers of the same model (as containers are often stacked when transported). The base of the container must be designed to withstand repeated dragging across hard rough floor surfaces. | [Please refer to photos showing moulded features on the top and on the bottom of the CB allowing multiple units to be stacked on top of one another in a safe and stable manner. Also show how the base of the container is designed to withstand repeated dragging across hard rough floor surfaces.]
4.2.18 | Corrosion resistance
| Hinges, stays, catches or handles and fixings, if metallic, must be constructed in stainless steel, aluminium, or other suitable non-ferrous metal. Casings and linings, if metallic, must be constructed in stainless steel, aluminium, or using zinc coated steel sheet to EN 10152 with a corrosion-resistant plastics coating to EN 10169-1. | [Please refer to documents or test reports that show evidence for compliance to these corrosion requirements. NOTE: Compliance to equivalent or better international standards can be acceptable.]
4.2.19 | Chemical resistance
| The external and internal surfaces of the container must be resistant to chemicals used for disinfecting (e.g. sodium hypochlorate, 5.25% in water). | [Please show how the surfaces are resistant to disinfecting chemicals as stated in the requirement. NOTE: This could be test evidence or certification from your material supplier.]
4.2.20 | IP rating (Type 2 only)
| Protection of the container with lid closed and latched not less than IEC 60529: IP55. | [Please refer to test report(s) showing compliance with this IP requirement.]
4.2.21 | Robustness and vibration (Type 2 only)
| The container with a full coolant load and a full load of vaccine vials and diluent ampoules must withstand a 1 metre drop onto each face, edge, corner, or chime when tested in accordance with ASTM D5276-09. At the end of the test there must be no damage that affects the performance of the product and the lid or door must still close and latch correctly. The container with a full coolant load and a full load of vaccine vials and diluent ampoules must withstand random vibration testing typical of unpaved road conditions in accordance with ASTM D4169-09, with no damage to either the container or contents. | [Please refer to test report(s) showing compliance with this drop test and vibration requirement.]
4.2.22 | Load restraint equipment (Type 2 only)
| The container must be provided with a minimum of two load restraint attachment points on exterior side walls and must be supplied with a full set of compatible ratchet tie-down straps. The complete assembly must be designed to prevent the container at its maximum loaded weight from breaking free when attached to the specified number of vehicle lashing points tested in accordance with ISO 27956 Road vehicles – Securing of cargo in delivery vans –Requirements and test methods. | [Please refer to photos or drawings showing compliance with this requirement.]
4.2.23 | Lock
| The door or lid must be fitted with a lock. Two keys are to be supplied with every unit. | [Please refer to photo(s) showing compliance with this requirement.]
4.2.24 | Thermometer
| Thermometers that use battery or photovoltaic (light) are permitted. Thermometers powered by a battery must use batteries with 5 year service life at +32°C. Option A: Externally readable cabinet-mounted gas or vapour pressure dial thermometer complying with PQS specification E06/TH02. Option B: Externally readable cabinet-mounted electronic thermometer conforming to PQS specification E06/TH06. Thermometer and other visual displays may be positioned on the front of the unit; preferably as close to eye level as possible. Alternatively they may be mounted on top of the unit at a height not exceeding 1.3 metres. If a low level position is essential, the display should be aligned so that it can easily be read without the user having to squat or kneel down. | [Please state and show the thermometer option selected for this CB. The photo should show the positioning of the Thermometer and if at low level, display alignment.]
4.2.25 | Coolant for immunization sessions
| It is mandatory that the product provide a means for cooling reconstituted and other vaccines that are removed for daily immunization sessions. Vaccine is to be maintained at a temperature of greater than 0°C and less than +8°C at the point of use. Coolant for this purpose may be stored in a separate container, but this container must be supplied, and will be tested, as a component part of the product. The ability also to supply removable coolant for outreach sessions is desirable but not mandatory. | [Does the product have a separate container to store coolant supplied, and will be tested, as a component part of the product? If so, please refer to a separate specification sheet for this container.]
4.2.26 | Areas not suitable for vaccine storage
| Areas of an otherwise acceptable product which are too warm or too cold must be excluded from use by design. | [Please refer to test report sections that show all vaccine storage temperatures are within the acceptable temperature range or if outside the range, show how these regions are excluded from use.]
4.3 | Environmental requirements:
4.3.1 | Ambient temperature range during transport, storage and use: -30°C to +55°C. | [Please refer to evidence that show each material of construction performs well in this temperature range.]
4.3.2 | Ambient humidity range during transport, storage and use 5% to 95% relative humidity, non-condensing. | [Please refer to evidence that show each material of construction performs well in this humidity range.]
4.4 | Physical characteristics
4.4.1 | Overall dimensions: In conformity with clause 4.5.1 to 4.5.4. | [See Clauses 4.5.1 to 4.5.4.]
4.4.2 | The weight of the container must not exceed the following figures: Type 1: Mechanical lifting equipment will typically not be available at the installation sites. It is recommended that the long-term storage cold box and any associated components should be designed for lifting in such a way that no single worker is required to carry more than 25 kg whilst working on their own, or in a group. Maximum weight: 100 kg per individual component in its packaging. Type 2: Maximum loaded weight, inclusive of the recommended number of coolant-packs: 50 kg. | [Please state the weight of the container and if more than 25 kg, show how it is designed for lifting by more than one person.]
4.5 | Interface requirements
4.5.1 | Dimensional compatibility with water-packs
| If water-packs are used, the internal dimensions of the container must be compatible with at least one of the three standard types of water-pack specified in E05/IP01. However, it is acceptable for the product to achieve its designated cold life at its designated vaccine storage capacity using only one of these three types. | [Please refer to photos or diagrams which show optional water-pack arrangement(s). (As stated, the water-packs may be of different size but they must not be of mixed size within one arrangement.)]
4.5.2 | Dimensional compatibility with vaccine packaging
| The net dimensions of the vaccine storage compartment (with coolant-packs in place) preferably should accommodate the widest possible range of standard immunization vaccines in secondary packaging or as primary containers, with the primary containers arranged upright in the container in a holding device. The holding device in Type 2 units must be designed to prevent dislodgement of the primary containers during transport. See Annex 3 and 4. | [Please state the dimensions or refer to the test report section of the vaccine storage compartment refer to photos or drawings showing example vaccine packaging (See Annex 3).]
4.5.3 | Dimensional compatibility with transport mode
| Type 1 containers, or their component parts, must be designed so that they can safely be carried inside cars, light vans, 4WD vehicles, boats, or carried by four porters. Type 2 containers and/or a single set of coolant container(s) or coolant-packs sufficient for one full recharge must be designed so that they can safely be carried upright inside cars, light vans or 4WD vehicles, strapped to the luggage rack of a small motorcycle or moped or carried by two porters or a pack animal. | [Please refer to large clear photos that show how the vaccine carrier is designed to be easily carried inside cars, light vans, 4WD vehicles, boats, or carried by four porters, strapped to the luggage rack of a small motorcycle or moped or carried by two porters or a pack animal.]
4.5.4 | Overall dimensions
| Type 1: To allow for manoeuvring through corners, corridors and doorways, the minimum dimension of the product (length, width or height) should not exceed 710 mm; exceptionally a minimum dimension up to 830 mm can be accepted, but this will restrict the number of sites where the product can be installed. The maximum dimension must not exceed 1700 mm and the maximum diagonal (corner to corner) dimension must not exceed 1850 mm. Type 2: To allow for transport by car the maximum width or diameter should be 700 mm or less and the maximum height must be 1000 mm or less. | [Please state the dimensions or refer to the test report section that shows the external dimensions.]
4.6 | Human factors
| The product must be designed in accordance with the general usability principles laid out in ISO 20282-1: 2006. | [Please explain how your product complies with the requirements of ISO 20282-1 : 2006.]
4.7 | Materials:
4.7.1 | Casing material selection
| Internal and external casing materials and all joints between the moulded components must be water and vapour proof, must resist UV degradation, must be easy to clean and must be selected with environmentally safe end-of life disposal in mind. Chlorinated plastics and composites containing epoxy resins are not permitted. | [Please refer to evidence that shows: 1. Internal and external casing materials and all joints between the moulded components are water and vapour proof, 2. resistance to UV degradation, 3. are easy to clean 4. are consistent with environmentally safe end-of life disposal.]
4.7.2 | Thermal insulation foaming agents
| Any gas complying with the limitations and deadlines set by the Montreal Protocol on the elimination of ozone-depleting chemicals. Insulation materials with a low global warming potential (GWP) are preferred. Foam insulation materials must prevent any loss of foaming agent during the design life of the product in all cases where the foaming agent contributes to thermal performance. | [Please state the thermal insulation foaming agent and how it complies with the limitations and deadlines set by the Montreal Protocol on the elimination of ozone-depleting chemicals.]
4.7.3 | PCM
| Integrated thermal buffer materials may be used to prevent freezing temperatures from propagating to the vaccine storage compartment or for other thermal purposes. The buffer material may be PCM-based but if so, must comply with WHO/PQS/ E005/PCMC0.1– PCM specification for phase-change material containers. | [Please state the PCM used, if any, and refer to a test report showing compliance with WHO/PQS/ E005/PCMC0.1– PCM specification for phase change material containers.]
4.8 | Warranty
| The product is to be covered by a two year replacement warranty in the event of any component failure arising from defective design, materials or workmanship. The manufacturer should state the time period over which the rated cold life is assured. | [Please refer to the warranty including the time period over which the rated cold life is assured. This document should also explain the terms and conditions.]
4.9 | Servicing
| The product is to be designed to achieve a life of not less than 10 years, apart from routine cleaning and routine maintenance as defined by the manufacturer. | [Please show evidence that the design life is at least 10 years. Also specify the routine cleaning and routine maintenance as detailed in the user manual to ensure the design-life of not less than 10 years.]
4.10 | Disposal and recycling
| The manufacturer is to provide information to the buyer on any hazardous materials contained within the product and suggestions for resource recovery/recycling and/or environmentally safe disposal. | [Please refer to information on any hazardous materials contained within the product and suggestions for resource recovery/recycling and/or environmentally safe disposal. NOTE: this should be a section in the user manual.]
4.11 | Instructions
| User and maintenance instructions in Arabic, English, French, Mandarin Chinese, Russian and Spanish. The instructions need not repeat the information shown on the permanent labels, but must include guidance on how to repair minor damage. Instructions must include standard operating procedures for coolant recharging that thoroughly address coolant-pack preparation including freezing appliance capacity and dimensions, transport requirements, energy consumption, maintenance and operations. Instructions must clearly indicate how to manage and transport coolant, specifying the type of coolant transport container to be used and restrictions on transport time and temperature. If removable baskets or other vaccine compartment accessories are supplied, fix a multi-lingual warning within the long-term storage cold box instructing users to Store vaccine in baskets only, or other appropriate instruction. | [Please provide a clear copy of the user and maintenance instructions English in the first instance. These instructions must include clear guidance on: - repair of minor damage - standard operating procedures for coolant recharging that thoroughly address coolant-pack preparation including freezing appliance capacity and dimensions, transport requirements, energy consumption, maintenance and operations. - Instructions must clearly indicate how to manage and transport coolant, specifying the type of coolant transport container to be used and restrictions on transport time and temperature. - use of removable baskets or other vaccine compartment accessories if provided and show a photo of the multi-lingual warning within the long-term storage cold box instructing users to "Store vaccine in baskets only", or other appropriate instruction.]
4.12 | Training
| No requirement. Training on prevention of vaccine freeze damage and correct use of coolant-packs and cool-packs is the responsibility of the purchaser. | (For information.)
4.13 | Verification
| In accordance with PQS Verification Protocol E004/CB03 VP.1. | [Please refer to the laboratory report.]
5. | Packaging
| Materials used for packaging the finished product are to be free of ozone-depleting compounds as defined in the Montreal Protocol. The general specification of shipping containers will be subject to agreement with the individual procurement agencies. | [Please state how the packing meets this requirement and refer to recommendations for correct disposal in the user manual.]