## Sheet1
E004 Submission to PQS for approval: Vaccine cold box with freeze-prevention technology |  |  | Brand & Model:
| Tested in accordance with PQS verification protocol: | E004/CB05-VP.4 | Company name:
| To verify compliance with PQS specification: | E004/CB05.4
Main contact:
Instructions in [square brackets] are for your guidance. Please delete these instructions and replace with evidence or a precise reference (including page and section numbers) showing compliance with that particular PQS specification requirement. In general a "pass" or "complies" statement alone is not acceptable.
One document may provide evidence for several clause requirements. Please carefully reference the sub-sections of any such document so that it is clear that the clause requirement has been met. This will help speed the pre-qualification process.
Greyed-out cells do not require a response.
NOTE: If it is not clear or doubtful whether a requirement has been met, then this is a non-compliance and pre-qualification might be delayed.
|  |  | Reviewer's summary: |  | Reviewer's summary: |  | Reviewer's summary:
7. | Product dossier | Applicant's submission | Reviewer's remarks [Reviewer's name & date] | Applicant's response | Reviewer's remarks [Reviewer's name & date] | Applicant's response | Reviewer's remarks [Reviewer's name & date] | Applicant's response
The legal manufacturer or reseller is to provide WHO with a pre-qualification dossier containing the following:
| Dossier examination fee in US dollars. | [Please arrange payment with PQS.]
| General information about the legal manufacturer, including name and address. | [Please state.]
| Unique identification reference for the product type. | [Please state.]
| Brand name of the product. | [Please state.]
| Full specifications of the product being offered, covering all the requirements set out in this document, including details of product marking and traceability. | [Please list the documents which show the product specifications.]
| A set of digital photographs showing a three-quarter view of the unit with the entry closed, and three-quarter view with the door or lid fully opened and a top view of the interior with the full complement of coolant-packs in place. | [Please refer to a document showing good quality photos. NOTE that photos are useful to help show compliance to particular clauses in the PQS Specification.]
| Certified photocopies of all type-approvals obtained for the product. | [Please refer to certified copies of all type approvals.]
| Certified photocopies of the legal manufacturer’s current ISO 9001 quality system certification. | [Please refer to a certified copy of the ISO 9001 certificate.]
| Where relevant, certified photocopies of the legal manufacturer’s ISO 14001 certification, EMAS registration or registration with an equivalent environmental audit scheme. Conformity with an environmental audit scheme is not mandatory; however preference will be given to manufacturers who are able to demonstrate compliance with good environmental practice. | [Please refer to certified copies of the ISO 14001 certificate and/or some other environmental conformity scheme.]
| Where available, laboratory test report(s) proving conformity with the product specifications. | [Please refer to the laboratory report.]
| Indicative cost of the product per unit, per 10 units and per 100 units, including coolant-packs, EXW (Incoterms 2010). | [Please state or refer to a document.]
Section 4. | Requirements
4.1 | General
Short range or long range thermally insulated cold box, with an operable lid designed for transporting vaccines in secondary cartons and capable of being carried by no more than two people when fully loaded (i.e., not exceeding 50kg in weight when fully loaded). The accuracy necessary for all temperature measurements and meeting of all temperature requirements is ±0.5°C.
4.2 | Performance
4.2.1 | Vaccine storage capacity
| 5.0 litres or greater. | [Please state the vaccine storage capacity.]
4.2.2 | Minimum vaccine storage compartment temperature
| The temperature inside the vaccine storage compartment must remain above 0°C during the cold life test conducted at an ambient temperature of +43°C. The temperature must also remain above 0°C during the minimum rated ambient temperature test when conducted at +15°C (or at a lower temperature specified by the manufacturer). | [Please refer to the test report section that shows compliance with this requirement.]
4.2.3 | Cold life
| • Short range: minimum 48 consecutive hours within the range 0°C to +10.0°C when tested at constant ambient temperatures of +43°C. • Long range: minimum 96 consecutive hours within the range 0°C to +10.0°C when tested at constant ambient temperatures of +43°C. | [Please state whether short or long range and refer to the test report section that shows compliance with the cold life requirement.]
4.2.4 | Cold climate freeze protection life
| No standard set; however, performance data will be published. | [Please refer to the test report section that shows the cold climate freeze protection life result.]
4.2.5 | Conditions of use
| Cold boxes will typically be transported on open, curtain-sided or fully enclosed light vans or trucks, or on the luggage rack of motorcycles, over road surfaces, including smooth tarmac, heavily perforated tarmac, rough gravel with deep corrugations, wet mud and dirt. They may also be carried by porters or pack animals or transported in boats. Products must be designed to accommodate these conditions and the design must take account of possible long-term exposure to dust, rainfall, high humidity, and to high levels of direct sunlight. | [Please state the features about the vaccine carrier design make it good for: - being transported by foot, - being transported by bicycle, - being transported by light motorcycle - being transported by a road vehicle over road surfaces with rough gravel with deep corrugations, wet mud and dirt. - carried by pack animals, - transported in boats, - long-term exposure to dust, - long-term exposure to rainfall, - high levels of direct sunlight.]
4.2.6 | Design principles
| The design of the container should respect the following principles: | [Please state the features about the vaccine carrier design make it good for:
| • Consistent with achieving the specified performance, the ratio between external volume and vaccine storage capacity must be kept to a maximum of 13:1. However, if for any reason, such as a unique design or geometric configuration, the device’s ratio is higher than 13:1, it will be considered by PQS on a case-by-case basis. | - the ratio between external volume and vaccine storage capacity,
| • The weight of coolant-packs must be the minimum needed to achieve the cold life required. | - weight of coolant-packs minimum needed to achieve the required cold life,
| • Placement of coolant-packs must minimize temperature stratification within the load. However, if coolant-packs are located above the vaccine storage compartment they must not close off the vaccine storage compartment and prevent access during immunization sessions. | - minimize temperature stratification within the load,
| • Cooling of the container must be achieved using fully frozen ice-packs taken directly from a freezer room or freezer at -25°C. | - cooling of the container by ice-packs,
| • Integrated buffer materials and/or container configuration may be used to prevent freezing temperatures from propagating to the vaccine storage compartment. Water-based buffers as well as PCM-based buffers are allowed as of the published date of this document. If PCM is used, the material and design must comply with WHO/PQS/ E005/PCMC0.1– PCM specification for phase-change material containers.. | - prevention of freezing temperatures from propagating to the vaccine storage compartment,
| • The design must eliminate the possibility of vaccine packages being exposed to air temperatures reaching 0°C, or contacting surfaces with a temperature of 0°C or less under any of the specified test conditions. | - no possibility of vaccine packages being exposed to air temperatures reaching 0°C, or contacting surfaces with a temperature of 0°C or less under any of the specified test conditions,
| • All pre-qualification testing must be carried out with a full complement of ice-packs. | - testing has been performed with all ice-packs,
| • The minimum possible number of user-interventions should be used to eliminate the possibility of vaccine packages or containers being exposed to freezing temperatures. The carrier design will be given a freeze protection classification based on the number of user-interventions. All pre-qualification testing must be carried out with a full complement of icepacks, or warm water-packs. However, where such arrangements have the potential to extend the ambient operating temperature range of the product, manufacturers are free to describe alternative arrangements which comprise a mix of frozen and unfrozen coolant-packs of the same type. See clause 4.11. | - minimum possible number of user-interventions to eliminate the possibility of vaccine packages or containers being exposed to freezing temperatures.]
4.2.7 | Reuse cycle
| It must be possible to reuse the container safely within 12 hours of removal of the previous charge of coolant-packs when the container is stored at the minimum rated ambient temperature | [Please state compliance or otherwise to this requirement.]
4.2.8 | Shape
| The recommended shape of the outside of the container and the vaccine storage compartment is substantially square or rectangular in plan and section. Rounded external corners are preferred. The reason for this is to ensure greater stability during transport, and to permit easy stacking of multiple cold boxes in a spatially-effective manner. However, if for any reason, such as a unique design or geometric configuration, the device is designed to have a different shape, it will be considered by PQS on a case-by-case basis. | [Please refer to photos or drawings showing compliance with this requirement.]
4.2.9 | Lid seal
| The lid of the container must be fitted with an effective seal which engages with the container walls when the lid is closed. This is to minimize cold bridging and maximize structural strength and resistance to transport vibrations and rough handling. | [Please refer to photos and describe how your product is designed to meet this requirement.]
4.2.10 | Hinges
| The lid must be fitted with a robustly constructed, non-detachable hinge mechanism that can open beyond 90° to give full access to the interior of the cold box. Hinges must be recessed so that they are fully protected against damage during transport and storage, and must be secured to the container in a manner that prevents loosening due to vibration. | [Please refer to photos and describe how your product is designed to meet this requirement.]
4.2.11 | Lid stay
| The container must be fitted with a stay device designed to prevent damage caused by over-stressing the lid when it is open. The stay device must be designed to prevent it becoming trapped between the lid seal and the body of the container at all points between fully closed and fully opened. The stay must be secured to the container in a manner that prevents loosening due to vibration. | [Please refer to photos and describe how your product is designed to meet this requirement.]
4.2.12 | Catches
| The lid must be secured in the closed position by a catch or catches. It must not be possible for a catch to open accidentally once engaged. Catches must be recessed so that they are fully protected against damage during transport and storage. Catches must be maintenance-free, without need for lubrication and must be secured to the container and to the lid in a manner that prevents loosening due to vibration. | [Please refer to photos and describe how your product is designed to meet this requirement.]
4.2.13 | Carrying handles
| The body of the container must be fitted with at least two moulded-in or hinged handles designed so that they can be used to lift and carry the container comfortably when it is fully loaded. Hinge mechanisms must restrain the handle in a near-horizontal position when the container is being moved or guided and the handles must automatically drop back into the vertical position when they are released. Preferably the handles should be recessed so that they are fully protected against damage during transport and storage. Handles must be maintenance-free, without need for lubrication and must be secured to the container in a manner that prevents loosening due to vibration. The handle arrangement must not prevent stable stacking of the containers or prevent close packing when the containers are placed side by side. | [Please refer to photos and describe how your product is designed to meet this requirement.]
4.2.14 | Vaccine storage advice and load restraint instructions
| Cold boxes must carry factory-fitted non-removable labels designed to last the lifetime of the container. • On the outside of the door or lid, and/or on the front face of the cold box: “Stop!” label and the appropriate temperature zone symbol as Annex 1. These instructions should be predominately graphical/visual. • On the inside of the door or lid: Coolant-pack and vaccine loading instructions as Annex 2. These instructions should be predominately graphical/visual. This label must be in the UN language most appropriate to the country of use (Arabic, English, French, Mandarin Chinese, Russian, or Spanish, or other language, by special order). | [Please refer to photos of the 2 labels.]
4.2.15 | Stacking and handling
| The design of the base and lid of the container should include moulded features that allow multiple units of the same model to be stacked on top of one another in a safe and stable manner. The lid of the container should be designed to withstand the simultaneously weight of three loaded containers of the same model (as containers are often stacked when transported). The base of the container must be designed to withstand repeated dragging across hard rough floor surfaces. | [Please refer to photos and describe how your product is designed to meet this requirement including how any strap or handle does not hinder stacking. NOTE: Photos of the top and bottom of the carrier are required.]
4.2.16 | Ventilation of stored containers
| The design of the lid hinges and catches must enable the container lid to open minimally for ventilation purposes during long-term storage in order to prevent mould growth on the interior lining. The mechanism and geometry for achieving this must not prevent stable stacking as described in clause 4.2.15. In addition, there must be no risk that the ventilation mechanism can become engaged when the container is lined with coolant-packs and the lid is closed for transport. The lid propping mechanism must be fully integrated into the container so that it cannot be separated and lost. | [Please refer to photos and describe how your product is designed to meet this requirement.]
4.2.17 | Corrosion resistance
| Hinges, stays, catches or handles and fixings, if metallic, must be constructed in stainless steel, aluminium, or other suitable non-ferrous metal. Casings and linings, if metallic, must be constructed in stainless steel, aluminium, or using zinc coated steel sheet. Metal components must be protected against corrosion as appropriate to EN ISO 6270-1 / ASTM D2247 / EN 13523-26 Determination of resistance to humidity – Part 1: Continuous condensation, EN ISO 6270-2 / EN 13523-25 Determination of resistance to humidity - Part 2: Procedure for exposing test specimens in condensation-water atmospheres, ISO 6272 / EN 13523-5 Impact resistance – external cabinet, and ISO 2409: 2013: Paints and varnishes – cross cut test (external cabinet). | [Please refer to evidence that shows any hinges, stays, catches or handles and fixings (if metallic), are sufficiently corrosion protected. NOTE: Compliance to equivalent or better international standards can be acceptable.]
4.2.19 | Chemical resistance
| The external and internal surfaces of the container must be resistant to chemicals used for disinfecting (e.g. sodium hypochlorate, 5.25% in water). | [Please explain how materials used make the surfaces resistant to disinfecting chemicals as stated in the requirement. NOTE: This could be test evidence or certification from your material supplier.]
4.2.19 | IP rating
| Protection of the container with lid closed and latched must not be less than IEC 60529: IP55. | [Please refer to a particular section in the test report or another report.]
4.2.20 | Robustness and vibration
| The container with a full coolant-pack load and a full load of vaccine vials and diluent ampoules must withstand a one metre drop onto each face, edge, corner, or chime when tested in accordance with ASTM D5276-09. At the end of the test there must be no damage that affects the performance of the product and the lid or door must still close and latch correctly. The container with a full coolant load and a full load of vaccine vials and diluent ampoules must withstand random vibration testing typical of unpaved road conditions in accordance with ASTM D4169-09, with no damage to either the container or contents. | [Please refer to a particular section in the test report or another report.]
4.2.21 | Coolant-packs
| The container must be provided with one full set of coolant-packs with complete installation and preparation instructions. Coolant-packs must comply with E005/IP01, Type 1 (0.3 litre), Type 2 (0.4 litre), and Type 3 (0.6 litre). Any of these three types may be used, but a mix of the three sizes is not acceptable. In the event the container uses unique shaped coolant-packs (i.e., not complying with Type 1, 2, or 3), two full sets of coolant-packs must be provided. Such unique sized coolant packs must be approved by PQS. | [Please state how many of a particular size coolant pack are provided with the carrier. Also refer to the PQS product reference number for these packs. Please also see Clause 4.7.4.]
4.2.22 | Coolant-packs restraint system
| The container must be designed so that coolant-packs are physically held in position without need for support by the load. The positioning of coolant-packs must not require user-adjustable fittings (e.g., removable barriers, wrappings needed for freeze protection for vaccines), since these would be regarded as user-interventions for freeze protection. | [Please refer to photos and describe how your product is designed to meet this requirement.]
4.3 | Environmental requirements
4.3.1 | Ambient temperature range during transport, storage and use:
| -30°C to +70°C. | [Please provide evidence that the materials used in construction of the vaccine carrier maintain their structural integrity during the quoted ambient range.]
4.4 | Physical characteristics
4.4.1 | Overall dimensions
| In conformity with clauses 4.5.1 and 4.5.2. | [Please state external dimensions.]
4.4.2 | Weight
| The maximum loaded weight of the container must not exceed 50 kg, inclusive of the recommended number of coolant-packs. | [Please state the maximum loaded weight ]
4.5 | Interface requirements
4.5.1 | Dimensional compatibility with vaccine packaging
| Containers will generally be used to transport vaccine in secondary cartons. The net dimensions of the storage compartment (length, breadth and height, with coolant-packs in place) should accommodate the widest possible range of routine vaccines supplied in this form of packaging. Refer to Annex 3. Cartons may be oriented in all three planes – for transport purposes it is not essential that vial caps face upwards. | [Please state the dimensions of the vaccine storage compartment provide photos or drawings showing example vaccine packaging with reference to Annex 3 "Table of secondary carton sizes".]
4.5.2 | Compatibility with distribution method
| Containers must be designed so that they can be securely strapped, upright, to the bed of a truck or the luggage rack of a motorcycle. | [Please refer to photos and describe how the vaccine carrier is designed to be easily strapped, upright, to a truck bed or luggage rack of a motorcycle.]
4.6 | Human factors
4.6.1 | Generally
| The product must be designed in accordance with the general usability principles laid out in ISO 20282-1. | [Please describe how the vaccine carrier complies with this standard.]
4.7 | Materials:
4.7.1 | Casing material selection
| Internal and external casing materials must resist UV degradation caused by long-term exposure to sunlight. All joints or liners between the moulded components must be water and vapour proof, must be easy to clean and must be selected with environmentally safe end-of-life disposal in mind. Chlorinated plastics and composites containing epoxy, polyurethane, phenol formaldehyde and urea formaldehyde resins are not permitted. | [Please provide evidence on how the properties the materials used in construction make the carrier - water and vapour proof, - resist UV degradation, - easy to clean and - environmentally safe end-of-life disposal.]
4.7.2 | Thermal insulation foaming agents
| Any gas complying with the limitations and deadlines set by the Montreal Protocol on the elimination of ozone-depleting chemicals. Insulation materials with a low global warming potential (GWP) are preferred. The chosen foam insulation material must be warranted to have a Long-Term Thermal Resistance (LTTR) no worse than [7%] lower than its rated ‘R’ value at the time of foaming, when subjected to an accelerated ageing test in accordance with ASTM C1303 / C1303M-12. | [Please state the thermal insulation foaming agent and how it complies with the limitations and deadlines set by the Montreal Protocol on the elimination of ozone-depleting chemicals.]
4.7.3 | Vacuum panels
| Vacuum panel construction is an acceptable alternative to foam insulation provided the panels are warrantied to retain their vacuum for a minimum period of five years under the intended conditions of use. | [Please state whether vacuum panels are used in the construction plus the warranty with terms and conditions.]
4.7.4 | PCM
| Integrated thermal buffer materials may be used to prevent freezing temperatures from propagating to the vaccine storage compartment or for other thermal purposes. The buffer material may be PCM-based but if so, must comply with WHO/PQS/ E005/PCMC0.1– PCM specification for phase change material containers. | [Please state the PCM used, if any, and make reference to a test report showing compliance with WHO/PQS/ E005/PCMC0.1– PCM specification for phase change material containers.]
4.8 | Warranty
| The product is to be covered by a two year replacement warranty in the event of any component failure arising from defective design, materials or workmanship. The manufacturer should state the time period over which the rated cold life is assured. | [Please refer to a document showing the replacement warranty including the time period over which the rated cold life is assured. This document should also explain the terms and conditions.]
4.9 | Servicing
| The product is to be designed to achieve a maintenance-free life of not less than 5 years, apart from routine cleaning. | [Please provide evidence to show how the properties of the materials used in construction help to achieve a maintenance-free life of not less than 5 years.]
4.10 | Disposal and recycling
| The manufacturer is to provide information to the buyer on any hazardous materials contained within the product and suggestions for resource recovery/recycling and/or environmentally safe disposal. | [Please refer to information on any hazardous materials contained within the product and suggestions for resource recovery/recycling and/or environmentally safe disposal. NOTE: this should be a section in the user manual.]
4.11 | Instructions
| User and maintenance instructions must be provided in at least one of the UN languages Arabic, English, French, Mandarin Chinese, Russian, or Spanish, or other language by special order. If only one language is included it must be the language most appropriate to the country of use. The instructions need not repeat the information shown on the permanent labels, but must include guidance on how to repair minor damage. | [Please refer to the user manual in English including instructions how to repair minor damage.]
4.12 | Training
| No requirement. Training on prevention of vaccine freeze damage and correct use of coolant-packs and warm-packs is the responsibility of the purchaser. | [Please state whether training could be available if requested by the purchaser.]
4.13 | Verification
| In accordance with PQS Verification Protocol E004/CB05-VP.4. | [Please refer to the laboratory report.]
5. | Packaging
| Materials used for packaging the finished product are to be free of ozone-depleting compounds as defined in the Montreal Protocol. The general specification of shipping containers will be subject to agreement with the individual procurement agencies. | [Please state how the packing meets this requirement and refer to recommendations for correct disposal in the user manual.]