## Sheet1
E004 Submission to PQS for approval: Vaccine carrier (standard) |  |  | Brand & Model:
| Tested in accordance with PQS verification protocol: | E004/VC01-VP.2 | Company name:
| To verify compliance with PQS specification: | E004/VC01.2
Main contact:
Instructions in [square brackets] are for your guidance. Please delete these instructions and replace with evidence or a precise reference (including page and section numbers) showing compliance with that particular PQS specification requirement. In general a "pass" or "complies" statement alone is not acceptable.
One document may provide evidence for several clause requirements. Please carefully reference the sub-sections of any such document so that it is clear that the clause requirement has been met. This will help speed the pre-qualification process.
Greyed-out cells do not require a response.
NOTE: If it is not clear or doubtful whether a requirement has been met, then this is a non-compliance and pre-qualification might be delayed.
|  |  | Reviewer's summary: |  | Reviewer's summary: |  | Reviewer's summary:
7. | Product dossier | Applicant's submission | Reviewer's remarks [Reviewer's name & date] | Applicant's response | Reviewer's remarks [Reviewer's name & date] | Applicant's response | Reviewer's remarks [Reviewer's name & date] | Applicant's response
The legal manufacturer or reseller is to provide WHO with a pre-qualification dossier containing the following:
| Dossier examination fee in US dollars. | [Please arrange payment with PQS.]
| General information about the legal manufacturer, including name and address. | [Please state.]
| Unique identification reference for the product type. | [Please state.]
| Brand name of the product. | [Please state.]
| Full specifications of the product being offered, covering all the requirements set out in this document, including details of product marking and traceability. | [Please list the documents which show the product specifications.]
| A set of digital photographs showing a three-quarter view of the unit with the lid closed, and three-quarter view with the lid fully opened and a top view of the interior with the full compliment of water-packs in place. | [Please refer to a document showing good quality photos. NOTE that photos are useful to help show compliance to particular clauses in the PQS Specification.]
| Certified photocopies of all type-approvals obtained for the product. | [Please refer to certified copies of all type approvals.]
| Certified photocopies of the legal manufacturer’s current ISO 9001 quality system certification. | [Please refer to a certified copy of the ISO 9001 certificate.]
| Where relevant, certified photocopies of the legal manufacturer’s ISO 14001 certification, EMAS registration or registration with an equivalent environmental audit scheme. Conformity with an environmental audit scheme is not mandatory; however preference will be given to manufacturers who are able to demonstrate compliance with good environmental practice. | [Please refer to certified copies of the ISO 14001 certificate and/or some other environmental conformity scheme.]
| Where available, laboratory test report(s) proving conformity with the product specifications. | [Please refer to the laboratory report.]
| Indicative cost of the product per unit, per 10 units and per 100 units, including coolant-packs, EXW (Incoterms 2000). | [Please state or refer to a document.]
Section 4. | Requirements
4.1 | General
Short range or long range insulated vaccine carrier, with insulated lid, designed for transporting vaccines. Both rigid non-rigid and semi-rigid designs are acceptable.
4.2 | Performance
4.2.1 | Vaccine storage capacity
| • Short range: 0.5 to 5.0 litres. • Long range: 1.0 to 5.0 litres. | [Please state the vaccine storage capacity and whether short or long range.]
4.2.2 | Cold life
| • Short range: minimum 15 hours. • Long range: minimum 30 hours. | [Please state whether short or long range and refer to the test report section that shows compliance with the cold life requirement.]
4.2.3 | Cool life
| No standard set; however, performance data will be published. | [Please refer to the test report section that shows the cool life result.]
4.2.4 | Warm life
| No standard set; however, performance data will be published. | [Please refer to the test report section that shows the warm life result.]
4.2.5 | Shape
| Vaccine carriers should be substantially square or rectangular in plan and section. Rounded corners are preferred. | [Please refer to photos or drawings showing compliance with this requirement.]
4.2.6 | Design principles
The design of the container, including the placement of the packs and of the load, must promote the free circulation of air within the container to ensure minimum temperature stratification. Container design should seek to minimize the weight of ice-packs required to meet the cold life requirement. Non-rigid and semi-rigid designs are acceptable provided they meet the required criteria for robustness.
4.2.7 | Lid
| Vaccine carriers must be fitted with an insulated lid which fits securely to the body of the container when closed so as to minimize thermal bridging and maximize structural strength and resistance to transport vibrations and rough handling. Hinged lids are acceptable, but are not mandatory | [Please refer to photos and describe how your product is designed to meet this requirement.]
4.2.8 | Hinges
| Hinges, where fitted, must allow the lid to open beyond 90° to give full access to the interior of the vaccine carrier. Preferably the hinges should be recessed so that they are fully protected against damage during transport and storage. Hinges must be maintenance-free, without need for lubrication and must be secured to the container in a manner which prevents loosening due to vibration. | [Please refer to photos and describe how your product is designed to meet this requirement.]
4.2.9 | Closure device
| The lid should be fitted with a mechanism to secure it in place so that the vaccine carrier does not open if it is dropped onto its side or onto its lid when full. Acceptable closure devices include, but are not confined to, magnetic or mechanical catches. It must not be possible for a catch to open accidentally once engaged. Mechanical catches must be recessed so that they are fully protected against damage during transport and storage. Catches must be maintenance-free, without need for lubrication and must be secured to the container in a manner which prevents loosening due to vibration. | [Please refer to photos and describe how your product is designed to meet this requirement.]
4.2.10 | Carrying device
| The body of the container must be fitted with one or more of the following carrying devices arranged so that the vaccine carrier can be comfortably carried in a substantially upright position: • Carrying handle: A hinged, sliding or moulded-in handle, attached to, or forming an integral part of the container body or lid. When folded away, moveable handles must not extend beyond the maximum length, width or height of the container. The handle arrangement must not prevent stable stacking of the carriers. • Shoulder strap: An adjustable strap arrangement which allows the vaccine carrier to be carried over the shoulder. • Backpack: An adjustable padded strap arrangement which allows the vaccine carrier to be carried as a backpack. All carrying devices must be robustly constructed and firmly attached in order to survive rough handling | [Please refer to photos and describe how your product is designed to meet at least one of the three options: - Carrying handle, - Shoulder strap,. - Backpack.]
4.2.11 | Vial holder (Optional):
| The vial holders, if offered, must be located immediately below the lid and above the vaccine storage compartment. It should be designed to hold a maximum of five opened vaccine vials for use during immunization sessions when the lid of the vaccine carrier is open. The vial holder must fully close off the vaccine storage compartment so that there is minimal loss of cold life. Design solutions may include, but are not limited to, the following: • Soft foam plastic pad, moulded to receive vaccine vials. The indentations must be designed so that vials are held from below and cannot be pushed through the pad. • Moulded plastic tray with indentations to receive vials. • Moulded plastic tray as previously described, but forming a separately hinged inner lid to the vaccine carrier. The vial holder must be able to hold vials, with capacities ranging from one to fifty doses, in a stable position and without risk of overturning the vial. | [Please state if a vial holding device is offer and whether this is - a soft foam plastic pad design, or - Moulded plastic tray with indentations to receive primary containers, or - Moulded plastic tray as previously described, but forming a separately hinged inner lid to the vaccine carrier, or - Another design (describe and show). (Photos and/or drawings can be useful.)]
4.2.12 | Vaccine storage advice
| Vaccine carriers must carry factory-fitted non-removable labels designed to last the lifetime of the appliance. Labels should be in the UN language most appropriate to the country of use (Arabic, English, French, Mandarin Chinese, Russian or Spanish, or other language, by special order) and should carry the following information. • On the outside of the lid, and/or on the front face of the vaccine carrier: As Annex 1. • On the inside of the lid: As Annex 2 | [Please refer to photos of the 2 labels.]
4.2.13 | Stacking
| The design of the base and lid of the container should allow multiple units of the same model to be stacked on top of one another in a safe and stable manner. The base of the container must be designed to withstand repeated dragging across hard, rough floor surfaces | [Please refer to photos and describe how your product is designed to meet this requirement including how any strap or handle does not hinder stacking. NOTE: Photos of the top and bottom of the carrier are required.]
4.2.14 | Corrosion resistance
| All metallic components and their fixings must be constructed in stainless steel or a suitable non-ferrous metal. | [Please describe the materials that make up all metallic components of the carrier.]
4.2.15 | Chemical resistance
| The external and internal surfaces of the container must be resistant to chemicals used for disinfecting (e.g. sodium hypochlorate, 5.25% in water). | [Please explain how materials used make the surfaces resistant to disinfecting chemicals as stated in the requirement. NOTE: This could be test evidence or certification from your material supplier.]
4.2.16 | IP rating
| Protection of the container with lid closed and latched must not be less than IEC 60529: IP55. | [Please refer to a particular section in the test report or another report.]
4.2.17 | Robustness
| The container must withstand a one meter drop onto each face, edge, and corner at its rated fully-loaded weight. At the end of the test there must be no damage that affects the performance of the product and the lid must still close and latch correctly. | [Please refer to a particular section in the test report or another report.]
4.3 | Environmental requirements
4.3.1 | Ambient temperature range during transport, storage and use:
| -30°C to +55°C. | [Please provide evidence that the materials used in construction of the vaccine carrier maintain their structural integrity during the quoted ambient range.]
4.4 | Physical characteristics
4.4.1 | Overall dimensions
| In conformity with clause 4.5.1 to 4.5.3. | [Please state external dimensions.]
4.4.2 | Weight
| The weight of the container must not exceed the following figures: • Short range: Maximum loaded weight, inclusive of the recommended number of water filled packs: 7.0 kg. • Long range: Maximum loaded weight, inclusive of the recommended number of water filled water-packs: 8.0 kg. | [Please state whether short or long range, the maximum loaded weight and the recommended number of water-filled packs.]
4.5 | Interface requirements
4.5.1 | Dimensional compatibility with packs
| The internal dimensions of the container must be compatible with any of the three standard types of water-pack specified in E005/IP01. However, it is acceptable for the product to achieve its designated cold life at its designated vaccine storage capacity using only one of these three types. | [Please state how many of a particular size coolant pack are provided and show photos of the coolant pack arrangement(s). Please also see Clause 4.7.3.]
4.5.2 | Dimensional compatibility with vaccine packaging
| Vaccine carriers are generally used to transport vaccine in individual vials. The net dimensions of the storage compartment (length, breadth and height, with water-packs in place) should accommodate all types of pre-filled vaccine presentation and the complete range of standard vaccine vials and ampoules up to 50 dose size. | [Please state the dimensions of the vaccine storage compartment provide photos or drawings showing example vaccine packaging with reference to PQS E004 VC02.1 Annex 3 "Table of secondary carton sizes".]
4.5.3 | Compatibility with distribution vehicles
| Vaccine carriers must be designed so that they can easily be strapped, upright, to the luggage rack of a bicycle or light motorcycle. | [Please refer to photos and describe how the vaccine carrier is designed to be easily strapped, upright, to the luggage rack of a bicycle or light motorcycle.]
4.6 | Human factors
4.6.1 | Generally
| The product must be designed in accordance with the general usability principles laid out in ISO 20282-1: 2006. | [Please describe how the vaccine carrier complies with this standard.]
4.6.2 | Portability
| All vaccine carriers must be designed so that they can comfortably be carried, when fully loaded, for periods of several hours by a male or female health worker wearing traditional dress. Backpack units may be incompatible with some types of clothing. | [Please refer to photos and describe how the vaccine carrier is designed to be comfortably carried, when fully loaded, for periods of several hours by a male or female health worker wearing traditional dress.]
4.7 | Materials:
4.7.1 | Casing material selection
| Internal and external casing materials and all joints between the moulded components must be water and vapour proof, must resist UV degradation, must be easy to clean and must be selected with environmentally safe end-of-life disposal in mind. Chlorinated plastics and composites containing epoxy resins are not permitted. | [Please provide evidence on how the properties the materials used in construction make the carrier - water and vapour proof, - resist UV degradation, - easy to clean and - environmentally safe end-of-life disposal.]
4.7.2 | Thermal insulation foaming agents
| Any gas complying with the limitations and deadlines set by the Montreal Protocol on the elimination of ozone-depleting chemicals. Cyclopentane and similar foaming agents with a low global warming potential (GWP) are preferred. | [Please state the thermal insulation foaming agent and how it complies with the limitations and deadlines set by the Montreal Protocol on the elimination of ozone-depleting chemicals.]
4.7.3 | PCM
| Integrated thermal buffer materials may be used to prevent freezing temperatures from propagating to the vaccine storage compartment or for other thermal purposes. The buffer material may be PCM-based but if so, must comply with WHO/PQS/ E005/PCMC0.1– PCM specification for phase change material containers. | [Please state the PCM used, if any, and make reference to a test report showing compliance with WHO/PQS/ E005/PCMC0.1– PCM specification for phase change material containers.]
4.8 | Warranty
| The product is to be covered by a two year replacement warranty in the event of any component failure arising from defective design, materials or workmanship. The manufacturer should state the time period over which the rated cold life is assured. | [Please refer to a document showing the replacement warranty including the time period over which the rated cold life is assured. This document should also explain the terms and conditions.]
4.9 | Servicing
| The product is to be designed to achieve a maintenance-free life of not less than 5 years, apart from routine cleaning. | [Please provide evidence to show how the properties of the materials used in construction help to achieve a maintenance-free life of not less than 5 years.]
4.10 | Disposal and recycling
| The manufacturer is to provide information to the buyer on any hazardous materials contained within the product and suggestions for resource recovery/recycling and/or environmentally safe disposal. | [Please refer to information on any hazardous materials contained within the product and suggestions for resource recovery/recycling and/or environmentally safe disposal. NOTE: this should be a section in the user manual.]
4.11 | Instructions
| User and maintenance instructions in Arabic, English, French, Mandarin Chinese, Russian and Spanish. The instructions need not repeat the information shown on the permanent labels, but must include guidance on how to repair minor damage. | [Please refer to the user manual in English including instructions how to repair minor damage.]
4.12 | Training
| No requirement. Training on prevention of vaccine freeze damage and correct use of ice-packs, cool-packs and warm-packs is the responsibility of the purchaser. | [Please state whether training could be available if requested by the purchaser.]
4.13 | Verification
| In accordance with PQS Verification Protocol E004/VC01-VP.2. | [Please refer to the laboratory report.]
5. | Packaging
| Materials used for packaging the finished product are to be free of ozone-depleting compounds as defined in the Montreal Protocol. The general specification of shipping containers will be subject to agreement with the individual procurement agencies. | [Please state how the packing meets this requirement and refer to recommendations for correct disposal in the user manual.]