## Sheet1
E004 Submission to PQS for approval: Vaccine carrier with freeze-prevention technology |  |  | Brand & Model:
| Tested in accordance with PQS verification protocol: | E004/VC02-VP.1 | Company name:
| To verify compliance with PQS specification: | E004/VC02.1
Main contact:
Instructions in [square brackets] are for your guidance. Please delete these instructions and replace with evidence or a precise reference (including page and section numbers) showing compliance with that particular PQS specification requirement. In general a "pass" or "complies" statement alone is not acceptable.
One document may provide evidence for several clause requirements. Please carefully reference the sub-sections of any such document so that it is clear that the clause requirement has been met. This will help speed the pre-qualification process.
Greyed-out cells do not require a response.
NOTE: If it is not clear or doubtful whether a requirement has been met, then this is a non-compliance and pre-qualification might be delayed.
|  |  | Reviewer's summary: |  | Reviewer's summary: |  | Reviewer's summary:
7. | Product dossier | Applicant's submission | Reviewer's remarks [Reviewer's name & date] | Applicant's response | Reviewer's remarks [Reviewer's name & date] | Applicant's response | Reviewer's remarks [Reviewer's name & date] | Applicant's response
The legal manufacturer or reseller is to provide WHO with a pre-qualification dossier containing the following:
| Dossier examination fee in US dollars. | [Please arrange payment with PQS.]
| General information about the legal manufacturer, including name and address. | [Please state.]
| Unique identification reference for the product type. | [Please state.]
| Brand name of the product. | [Please state.]
| Full specifications of the product being offered, covering all the requirements set out in this document, including details of product marking and traceability. | [Please list the documents which show the product specifications.]
| A set of digital photographs showing a three-quarter view of the unit with the entry closed, and three-quarter view with the door or lid fully opened and a top view of the interior with the full complement of coolant-packs in place. | [Please refer to a document showing good quality photos. NOTE that photos are useful to help show compliance to particular clauses in the PQS Specification.]
| Certified photocopies of all type-approvals obtained for the product. | [Please refer to certified copies of all type approvals.]
| Certified photocopies of the legal manufacturer’s current ISO 9001 quality system certification. | [Please refer to a certified copy of the ISO 9001 certificate.]
| Where relevant, certified photocopies of the legal manufacturer’s ISO 14001 certification, EMAS registration or registration with an equivalent environmental audit scheme. Conformity with an environmental audit scheme is not mandatory; however preference will be given to manufacturers who are able to demonstrate compliance with good environmental practice. | [Please refer to certified copies of the ISO 14001 certificate and/or some other environmental conformity scheme.]
| Where available, laboratory test report(s) proving conformity with the product specifications. | [Please refer to the laboratory report.]
| Indicative cost of the product per unit, per 10 units and per 100 units, including coolant-packs, EXW (Incoterms 2010). | [Please state or refer to a document.]
Section 4. | Requirements
4.1 | General
Short range or long range thermally insulated vaccine carrier, with insulated lid, designed for transporting vaccines. Both rigid non-rigid and semi-rigid designs are acceptable. Typically vaccines will be loaded as loose primary containers. Larger vaccine carriers may also be used to transport secondary cartons. The accuracy necessary for all temperature measurements and meeting of all temperature requirements is ±0.5°C.
4.2 | Performance
4.2.1 | Vaccine storage capacity
| • Short range: 0.5 to 5.0 litres. • Long range: 1.0 to 5.0 litres. Provided the weight limits in clause 4.4.2 are not exceeded, no maximum capacity is set. | [Please state the vaccine storage capacity and whether short or long range.]
4.2.2 | Minimum vaccine storage compartment temperature
| The temperature inside the vaccine storage compartment must remain above 0°C during the cold life test conducted at an ambient temperature of +43°C. The temperature must also remain above 0°C during the minimum rated ambient temperature test when conducted at +15°C (or at a lower temperature specified by the manufacturer). | [Please refer to the test report section that shows compliance with this requirement.]
4.2.3 | Cold life
| • Short range: minimum 15 hours when tested at constant ambient temperatures of +43°C and +15°C (or at a lower test temperature specified by the manufacturer). • Long range: minimum 30 hours when tested at constant ambient temperatures of +43°C and +15°C (or at a lower test temperature specified by the manufacturer). For both short and long range carriers, the cooldown must not exceed 8 hours. | [Please state whether short or long range and refer to the test report section that shows compliance with the cold life requirement.]
4.2.4 | Cold climate freeze protection life
| No standard set; however, performance data will be published. | [Please refer to the test report section that shows the cold climate freeze protection life result.]
4.2.5 | Conditions of use
| Vaccine carriers will typically be transported by foot, bicycle, light motorcycle or other road vehicle over road surfaces, including smooth tarmac, heavily perforated tarmac, rough gravel with deep corrugations, wet mud and dirt. They may also be carried by pack animals or transported in boats. Products must be designed to accommodate these conditions and the design must take account of possible long-term exposure to dust, rainfall and to high levels of direct sunlight. | [Please state the features about the vaccine carrier design make it good for: - being transported by foot, - being transported by bicycle, - being transported by light motorcycle - being transported by a road vehicle over road surfaces with rough gravel with deep corrugations, wet mud and dirt. - carried by pack animals, - transported in boats, - long-term exposure to dust, - long-term exposure to rainfall, - high levels of direct sunlight.]
4.2.6 | Design principles
| The design of the container should respect the following principles: | [Please state the features about the vaccine carrier design make it good for:
| • Non-rigid and semi-rigid designs are acceptable provided they meet the required criteria for robustness. | - robustness,
| • Consistent with achieving the specified performance, the ratio between external volume and vaccine storage capacity should be kept as low as possible. It is recommended that the ratio be kept to an absolute maximum of [22:1]. | - the ratio between external volume and vaccine storage capacity,
| • The weight of coolant-packs must be the minimum needed to achieve the cold life required. | - weight of coolant-packs minimum needed to achieve the required cold life,
| • Placement of coolant-packs must minimize temperature stratification within the load. However, if coolant-packs are located above the vaccine storage compartment they must not close off the vaccine storage compartment and prevent access during immunization sessions. | - minimize temperature stratification within the load,
| • Cooling of the container must be achieved using fully frozen ice-packs taken directly from a freezer room or freezer at -25°C. Integrated buffer materials and/or container configuration may be used to prevent freezing temperatures from propagating to the vaccine storage compartment. Water-based buffers as well as PCM-based buffers are allowed as of the published date of this document. If PCM is used, the material and design must comply with WHO/PQS/ E005/PCMC0.1– PCM specification for phase-change material containers.. | - prevention of freezing temperatures from propagating to the vaccine storage compartment,
| • The design must eliminate the possibility of vaccine packages being exposed to air temperatures reaching 0°C, or contacting surfaces with a temperature of 0°C or less under any of the specified test conditions. | - no possibility of vaccine packages being exposed to air temperatures reaching 0°C, or contacting surfaces with a temperature of 0°C or less under any of the specified test conditions,
| • The minimum possible number of user-interventions should be used to eliminate the possibility of vaccine packages or containers being exposed to freezing temperatures. The carrier design will be given a freeze protection classification based on the number of user-interventions. All pre-qualification testing must be carried out with a full complement of icepacks, or warm water-packs. However, where such arrangements have the potential to extend the ambient operating temperature range of the product, manufacturers are free to describe alternative arrangements which comprise a mix of frozen and unfrozen coolant-packs of the same type. See clause 4.11. | - minimum possible number of user-interventions to eliminate the possibility of vaccine packages or containers being exposed to freezing temperatures.]
4.2.7 | Reuse cycle
| It must be possible to reuse the container safely within 12 hours of removal of the previous charge of coolant-packs when the container is stored at the minimum rated ambient temperature | [Please state compliance or otherwise to this requirement.]
4.2.8 | Shape
| The outside of the container may be any shape consistent with comfortable carrying. Rounded external corners are preferred. The vaccine storage compartment may be substantially square or rectangular, however, other shapes with horizontal, planar top and bottom are acceptable. | [Please refer to photos or drawings showing compliance with this requirement.]
4.2.9 | Lid
| Vaccine carriers must be fitted with an insulated lid which fits securely to the body of the container when closed so as to minimize thermal bridging and maximize structural strength and resistance to transport vibrations and rough handling. Hinged lids are acceptable, but are not mandatory | [Please refer to photos and describe how your product is designed to meet this requirement.]
4.2.10 | Hinges
| Hinges, where fitted, must allow the lid to open beyond 90° to give full access to the interior of the vaccine carrier. Preferably the hinges should be recessed so that they are fully protected against damage during transport and storage. Hinges must be maintenance-free, without need for lubrication and must be secured to the container in a manner which prevents loosening due to vibration. | [Please refer to photos and describe how your product is designed to meet this requirement.]
4.2.11 | Closure device
| The lid should be fitted with a mechanism to secure it in place so that the vaccine carrier does not open if it is dropped onto its side or onto its lid when full. Acceptable closure devices include, but are not confined to, magnetic or mechanical catches. It must not be possible for a catch to open accidentally once engaged. Mechanical catches must be recessed so that they are fully protected against damage during transport and storage. Catches must be maintenance-free, without need for lubrication and must be secured to the container in a manner which prevents loosening due to vibration. | [Please refer to photos and describe how your product is designed to meet this requirement.]
4.2.12 | Carrying device
| The body of the container must be fitted with one or more of the following carrying devices arranged so that the vaccine carrier can be comfortably carried in a substantially upright position: • Carrying handle: A hinged, sliding or moulded-in handle, attached to, or forming an integral part of the container body or lid. When folded away, moveable handles must not extend beyond the maximum length, width or height of the container. The handle arrangement must not prevent stable stacking of the carriers. • Shoulder strap: An adjustable strap arrangement which allows the vaccine carrier to be carried over the shoulder. • Backpack: An adjustable padded strap arrangement which allows the vaccine carrier to be carried as a backpack. All carrying devices must be robustly constructed and firmly attached in order to survive rough handling | [Please refer to photos and describe how your product is designed to meet at least one of the three options: - Carrying handle, - Shoulder strap,. - Backpack.]
4.2.13 | Primary container holding device
| A holding device, if offered, should be designed to retain opened primary containers in a stable upright position during the course of an immunization session. The device must fully close off the vaccine storage compartment in order to limit air convection, limit loss of cold life, and keep the opened primary containers as cool as possible when the lid of the vaccine carrier is open. It must also be able to hold containers, with capacities ranging from one to fifty doses. Design solutions may include, but are not limited to, the following: • Soft foam plastic pad designed to receive primary containers. If indentations or slits are used, they must be designed so that containers are held from below and cannot be pushed through the pad. • Moulded plastic tray with indentations to receive primary containers. • Moulded plastic tray as previously described, but forming a separately hinged inner lid to the vaccine carrier. Because of the risk of loss or damage, the holding device must be available to purchasers as a spare part if it is separate or easily removed from the carrier itself. | [Please state if a primary container holding device is offer and whether this is - a soft foam plastic pad design, or - Moulded plastic tray with indentations to receive primary containers, or - Moulded plastic tray as previously described, but forming a separately hinged inner lid to the vaccine carrier, or - Another design (describe and show). (Photos and/or drawings can be useful.)]
4.2.14 | Vaccine storage advice
| Vaccine carriers must have factory-fitted non-removable labels designed to last the lifetime of the container. • On the outside of the door or lid, and/or on the front face of the carrier: ‘Stop!’ label as shown in Annex 1. These instructions should be predominately graphical. • On the inside of the door or lid: Coolant and vaccine loading instructions and carrier preparation instructions as required as shown in Annex 2. This label must be in the UN language most appropriate to the country of use (Arabic, English, French, Mandarin Chinese, Russian or Spanish, or other language, by special order). | [Please refer to photos of the 2 labels.]
4.2.15 | Stacking and handling
| The design of the base and lid of the container should allow multiple units of the same model to be stacked on top of one another in a safe and stable manner. The base of the container must be designed to withstand repeated dragging across hard, rough floor surfaces | [Please refer to photos and describe how your product is designed to meet this requirement including how any strap or handle does not hinder stacking. NOTE: Photos of the top and bottom of the carrier are required.]
4.2.16 | Corrosion resistance
| Hinges, stays, catches or handles and fixings, if metallic, must be constructed in stainless steel, aluminium, or suitable corrosion-resistant metal. Casings and linings, if metallic, must be constructed in stainless steel, aluminium, or using zinc coated steel sheet to EN 10152 with a corrosion-resistant plastics coating to EN 10169-1. | [Please refer to evidence that shows any hinges, stays, catches or handles and fixings (if metallic), are sufficiently corrosion protected. NOTE: Compliance to equivalent or better international standards can be acceptable.]
4.2.17 | Chemical resistance
| The external and internal surfaces of the container must be resistant to chemicals used for disinfecting (e.g. sodium hypochlorate, 5.25% in water). | [Please explain how materials used make the surfaces resistant to disinfecting chemicals as stated in the requirement. NOTE: This could be test evidence or certification from your material supplier.]
4.2.18 | IP rating
| Protection of the container with lid closed and latched must not be less than IEC 60529: IP55. | [Please refer to a particular section in the test report or another report.]
4.2.19 | Robustness and vibration
| The container with a full coolant-pack load and a full load of vaccine vials and diluent ampoules must withstand a one metre drop onto each face, edge, corner, or chime (circumferential edge) when tested in accordance with ASTM D5276-09. At the end of the test there must be no damage that affects the performance of the carrier and the lid or door must still close and latch correctly. The container with a full coolant load and a full load of vaccine vials and diluent ampoules must withstand random vibration testing typical of unpaved road conditions in accordance with ASTM D4169-09, with no damage to either the container or contents. | [Please refer to a particular section in the test report or another report.]
4.2.20 | Coolant-packs
| The container must be provided with one full set of coolant-packs with complete installation and preparation instructions. Coolant-packs must comply with PQS/E005/IP01, Type 2 (0.3 litre), Type 3 (0.4 litre), or Type 4 (0.6 litre). Only one of these types may be used – a mix of the sizes is not acceptable. | [Please state how many of a particular size coolant pack are provided with the carrier. Also refer to the PQS product reference number for these packs. Please also see Clause 4.7.3.]
4.3 | Environmental requirements
4.3.1 | Material properties at ambient temperatures:
| Carrier materials must be able to withstand temperatures during transport, storage and use of -30°C to +70°C. | [Please provide evidence that the materials used in construction of the vaccine carrier maintain their structural integrity during the quoted ambient range.]
4.4 | Physical characteristics
4.4.1 | Overall dimensions
| In conformity with clause 4.5.1 to 4.5.3. | [Please state external dimensions.]
4.4.2 | Weight
| The maximum loaded weight of the container, inclusive of the recommended number of coolant-packs, must not exceed the following figures: • Short range: 7.0 kg. • Long range: 8.0 kg | [Please state whether short or long range, the maximum loaded weight and the recommended number of water-filled packs.]
4.5 | Interface requirements
4.5.1 | Dimensional compatibility with coolant-packs
| The dimensions of the coolant-pack locations must be compatible with coolant-packs complying with PQS/E005/IP01 Type 2, 3, or 4. It is acceptable for the product to achieve its designated cold life at its designated vaccine storage capacity using only one of these types. | [Please state how many of a particular size coolant pack are provided and show photos of the coolant pack arrangement(s). Please also see Clause 4.7.3.]
4.5.2 | Dimensional compatibility with vaccine packaging
| Vaccine carriers are generally used to transport mixed loads of vaccine in individual primary containers. The net dimensions of the storage compartment (length, breadth and height, with coolant-packs in place) should accommodate all types of pre-filled vaccine presentation and the complete range of standard vaccine vials and ampoules up to 50 dose size as well as tubes and pre-filled devices. Larger vaccine carriers may also be used to transport secondary cartons. The net dimensions of the storage compartment (length, breadth and height, with coolant-packs in place) should accommodate the widest possible range of routine vaccines supplied in these forms of packaging. Secondary cartons may be oriented in all three planes – for transport purposes it is not essential that vial caps face upwards. | [Please state the dimensions of the vaccine storage compartment provide photos or drawings showing example vaccine packaging with reference to PQS E004 VC02.1 Annex 3 "Table of secondary carton sizes".]
4.5.3 | Compatibility with distribution method
| Containers must be designed so that they can be securely strapped, upright, to the luggage rack of a bicycle or light motorcycle. | [Please refer to photos and describe how the vaccine carrier is designed to be easily strapped, upright, to the luggage rack of a bicycle or light motorcycle.]
4.6 | Human factors
4.6.1 | Generally
| The product must be designed in accordance with the general usability principles laid out in ISO 20282-1: 2006. | [Please describe how the vaccine carrier complies with this standard.]
4.6.2 | Portability
| All vaccine carriers must be designed so that they can comfortably be carried, when fully loaded, for periods of several hours by a male or female health worker wearing traditional dress. Backpack units may be incompatible with some types of clothing. | [Please refer to photos and describe how the vaccine carrier is designed to be comfortably carried, when fully loaded, for periods of several hours by a male or female health worker wearing traditional dress.]
4.7 | Materials:
4.7.1 | Casing material selection
| Internal and external casing materials and all joints between the moulded components must be water and vapour proof, must resist UV degradation, must be easy to clean and must be selected with environmentally safe end-of-life disposal in mind. Chlorinated plastics and composites containing epoxy resins are not permitted. | [Please provide evidence on how the properties the materials used in construction make the carrier - water and vapour proof, - resist UV degradation, - easy to clean and - environmentally safe end-of-life disposal.]
4.7.2 | Thermal insulation foaming agents
| Any gas complying with the limitations and deadlines set by the Montreal Protocol on the elimination of ozone-depleting chemicals may be used. Insulation materials with a low global warming potential (GWP) are preferred. The chosen foam insulation material must be warrantied to have a Long-Term Thermal Resistance (LTTR) no worse than 7% lower than its rated ‘R’ value at the time of foaming, when subjected to an accelerated ageing test in accordance with ASTM C1303 / C1303M-12. | [Please state the thermal insulation foaming agent and how it complies with the limitations and deadlines set by the Montreal Protocol on the elimination of ozone-depleting chemicals.]
4.7.3 | PCM
| Integrated thermal buffer materials may be used to prevent freezing temperatures from propagating to the vaccine storage compartment or for other thermal purposes. The buffer material may be PCM-based but if so, must comply with WHO/PQS/ E005/PCMC0.1– PCM specification for phase change material containers. | [Please state the PCM used, if any, and make reference to a test report showing compliance with WHO/PQS/ E005/PCMC0.1– PCM specification for phase change material containers.]
4.7.4 | Vacuum panels
| Vacuum panel construction is an acceptable alternative to foam insulation provided the panels are warrantied to retain their vacuum for a minimum period of five years under the intended conditions of use. | [Please state whether vacuum panels are used in the construction plus the warranty with terms and conditions.]
4.8 | Warranty
| The product is to be covered by a two year replacement warranty in the event of any component failure arising from defective design, materials or workmanship. The manufacturer should state the time period over which the rated cold life is assured. | [Please refer to a document showing the replacement warranty including the time period over which the rated cold life is assured. This document should also explain the terms and conditions.]
4.9 | Servicing
| The product is to be designed to achieve a maintenance-free life of not less than 5 years, apart from routine cleaning. | [Please provide evidence to show how the properties of the materials used in construction help to achieve a maintenance-free life of not less than 5 years.]
4.10 | Disposal and recycling
| The manufacturer is to provide information to the buyer on any hazardous materials contained within the product and suggestions for resource recovery/recycling and/or environmentally safe disposal. | [Please refer to information on any hazardous materials contained within the product and suggestions for resource recovery/recycling and/or environmentally safe disposal. NOTE: this should be a section in the user manual.]
4.11 | Instructions
| User and maintenance instructions must be provided in at least one of the UN languages Arabic, English, French, Mandarin Chinese, Russian, or Spanish, or other language by special order. If only one language is included it must be the language most appropriate to the country of use. The instructions need not repeat the information shown on the permanent labels, but must include guidance on how to repair minor damage. | [Please refer to the user manual in English including instructions how to repair minor damage.]
4.12 | Training
| No requirement. Training on prevention of vaccine freeze damage and correct use of coolant-packs and warm-packs is the responsibility of the purchaser. | [Please state whether training could be available if requested by the purchaser.]
4.13 | Verification
| In accordance with PQS Verification Protocol E004/VC02-VP.1. | [Please refer to the laboratory report.]
5. | Packaging
| Materials used for packaging the finished product are to be free of ozone-depleting compounds as defined in the Montreal Protocol. The general specification of shipping containers will be subject to agreement with the individual procurement agencies. | [Please state how the packing meets this requirement and refer to recommendations for correct disposal in the user manual.]