## Sheet1
E005 Submission to PQS for approval: Phase-change material containers |  |  | Brand & Model:
| Tested in accordance with PQS verification protocol: | E005/PCMC01-VP0.1 | Company name:
| To verify compliance with PQS specification: | E005/PCMC0.1
Main contact:
Instructions in [square brackets] are for your guidance. Please delete these instructions and replace with evidence or a precise reference (including page and section numbers) showing compliance with that particular PQS specification requirement. In general a "pass" or "complies" statement alone is not acceptable.
One document may provide evidence for several clause requirements. Please carefully reference the sub-sections of any such document so that it is clear that the clause requirement has been met. This will help speed the pre-qualification process.
Greyed-out cells do not require a response.
NOTE: If it is not clear or doubtful whether a requirement has been met, then this is a non-compliance and pre-qualification might be delayed.
|  |  | Reviewer's summary: |  | Reviewer's summary: |  | Reviewer's summary:
7. | Product dossier | Applicant's submission | Reviewer's remarks [Reviewer's name & date] | Applicant's response | Reviewer's remarks [Reviewer's name & date] | Applicant's response | Reviewer's remarks [Reviewer's name & date] | Applicant's response
If the container is not a modified version of a PQS approved container where only the shape of the container has changed, then the legal manufacturer or reseller must provide WHO with a prequalification dossier containing the following:
| Dossier examination fee in US dollars. | [Please arrange payment with PQS.]
| General information about the legal manufacturer, including name and address. | [Please state.]
| Unique identification reference for the product type. | [Please state.]
| Full specifications of the product being offered, covering all the requirements set out in this document, including details of product marking and traceability. | [Please make reference to documents which show the product specifications and how the PQS specifications are met.]
| A description of the PCM with which the container is filled. This must include at a minimum: the type of PCM or primary component of the PCM if it is a blend of chemicals; the chemical or trade name of the PCM if different from the type identified; the name of the producer of the PCM; contact information for the producer of the PCM if different from the legal manufacturer of the product. | [Please make reference to documents which state the PCM with which the container is filled.]
| Certified photocopies of any applicable type-approvals obtained for the product (e.g. CE Mark, UL, etc.). | [Please make reference to certified copies of all type approvals.]
| Certified photocopies of the legal manufacturer’s current ISO 9001 quality system certification. | [Please make reference to certified copies of the ISO 9001 certificate.]
| Where relevant, certified photocopies of the legal manufacturer’s ISO 14001 certification, EMAS registration or registration with an equivalent environmental audit scheme. Conformity with an environmental audit scheme is not mandatory; however preference will be given to manufacturers who are able to demonstrate compliance with good environmental practice. | [Please make reference to certified copies of the ISO 14001 certificate or some other environmental conformity scheme.]
| Where available, laboratory test report(s) proving conformity with the product specifications. | [Please make reference to the laboratory report.]
| Provide one sample for preliminary inspection. | [Please contact PQS for address.]
| Indicative cost of the product per 100 units, per 1,000 units, and per 10,000 units EXW (Incoterms 2000). | [Please state or refer to a document.]
4. | Requirements
General
A robust container designed to store PCM which, when conditioned to an appropriate temperature, provides the thermal inertia needed to maintain safe storage conditions for vaccines and biological specimens. The container must be a fixed container. The fixed container must be produced as a separable component of the CCE that is later integrated or affixed into the CCE to avoid removal by end users. Removable containers are not permitted under this specification.
4.1 | Performance
4.1.1 | Filling
| Containers must be filled with PCM by the CCE manufacturer or obtained pre-filled from another entity. Containers must be permanently and robustly sealed by the manufacturer or this other entity. | [Please state how product meets this requirement.]
4.1.2 | Migration
| Containers holding PCM must resist migration of PCM through container walls. PCM migration through container walls must be on average less than 0.1% per month by mass as measured over a 60-day period. | [Please show compliance with this requirement by referencing appropriate test report section.]
4.1.3 | Compatibility
| Container materials must be compatible with the PCM used. Compatible means that the PCM will not significantly weaken the material, causing the container to fail over the life of the container and PCM. The PCM must also be compatible with the materials in the CCE in which the PCM container is used. The CCE manufacturer may cite existing literature to establish the compatibility of the PCM container and CCE materials with the PCM in the container. Acceptable literature is articles from refereed scientific or engineering journals, test data from the PCM manufacturer or test data from the PCM container manufacturer if the PCM container manufacturer is not the CCE manufacturer. | [Please refer to literature, documents, or test reports showing compliance with this requirement.]
4.1.4 | Robustness
| Containers must be able to withstand the following stress tests: • Freeze/thaw cycling – Containers shall not exhibit more than a 1% decrease in filled mass due to lost PCM after being subjected to 100 freeze/thaw cycles. • Compression test – No filled containers exhibit visual leakage after being subjected to an 800 N compression load on a minimum of three combinations of container sides. | [Please show compliance with this requirement by referencing appropriate test report section.]
4.2 | Environmental requirements
| Ambient temperature range during transport, storage and use: −30°C to +55°C. | [Please state the temperature range for the product.]
4.3 | Phase-change material hazards
4.3.1 | Physical hazards
| The PCM must not have any of the GHS physical hazard codes (H2xx). | [Please provide documentation showing compliance with the requirement.]
4.3.2 | Health hazards
| The PCM must not have any of the GHS health hazard codes (H3xx) with the exception of the following: • H304 − May be fatal if swallowed and enters airways. • H305 − May be harmful if swallowed and enters airways. • H316 − Causes mild skin irritation. • H317 − May cause an allergic skin reaction. • H320 − Causes eye irritation. | [Please provide documentation showing compliance with the requirement.]
4.3.3 | Environmental hazards
| The PCM must not have any of the GHS environmental hazard codes (H4xx). PCM and container must not generate toxic substances when incinerated between 650°C and 1,200°C. | [Please provide documentation showing compliance with the requirement.]
4.4 | Container labelling
| The outside of the PCM container must have labels indicating hazards of the PCM. Labels must conform to the Globally Harmonized System for Classification and Labelling of Chemicals and any additional labelling standards of the country in which the container is to be used. Labels must be in the United Nations (UN) language most appropriate for the country of use. Labels must be obvious, permanent and remain fixed and legible if it comes into contact with water or the PCM in the container. | [Please provide photos of the container showing appropriate labels.]
4.5 | Materials
| Containers must be made from materials that are known to be nontoxic when incinerated at any temperature between 650°C and 1,200°C. | [Please state how the materials used are to be non-toxic when incinerated at any temperature between 650°C and 1,200°C.]
4.6 | Warranty
| Fixed containers will have the same warranty as the CCE in which the PCM container is installed. | [Please make reference to a document showing the replacement warranty including terms and conditions.]
4.7 | Servicing provision
| Fixed containers should have a maintenance-free life at least as long as the CCE in which the PCM container is installed. | [Please state how the product is designed to meet this requirement.]
4.8 | Instructions
| User instructions must be provided in at least one of the UN languages: Arabic, English, French, Mandarin Chinese, Russian, or Spanish or other language by special order. If only one language is included, it must be the language most appropriate for the country of use. The instructions must state the volume and general type of PCM with which the container is filled. Instructions for clean-up of PCM and disposal in case of container failure must also be included. | [Please confirm the manual is available in the requested languages. Make reference to the user manual in English including the volume and general type of PCM with which the container is filled and on clean-up of PCM and disposal in case of container failure.]
4.9 | Training
No requirement. Training on prevention of vaccine freeze damage and correct use of PCM is the responsibility of the purchaser.
4.10. | Disposal and recycling
| The manufacturer is to provide information to the buyer on any hazardous materials contained within the PCM or PCM container and suggestions for resource recovery/recycling and/or environmentally-safe disposal. Instructions for product recovery/recycling and/or environmentally-safe disposal must be permanently affixed to the outside of the equipment in which the container is installed and permanently affixed to the outside of the container. | [Please make reference to the information on any hazardous materials contained within the product and suggestions for resource recovery/recycling and/or environmentally safe disposal. NOTE: this could be a section in the user manual.]
4.11 | Verification
| In accordance with World Health Organization (WHO) Performance, Quality and Safety (PQS) Verification Protocol E005 PCMC VP0.1. | [Please make reference to the laboratory report.]
5. | Packaging
| Materials used for packaging the finished PCM container or CCE are to be free of ozone-depleting compounds as defined in the Montreal Protocol. The general specification of shipping containers will be subject to agreement with the individual procurement agencies. | [Please state how the packing meets this requirement and refer to recommendations for correct disposal in the user manual.]