## Spec
E006 Submission to PQS for approval: Data logger and machine-to-machine interface for Equipment Monitoring Systems
| Compliance with PQS specification: | E006/DL01.3 | Brand & Model:
| As shown by testing in accordance with PQS verification protocol: | E006/DL01-VP.3
Instructions in [square brackets] are for your guidance. Please delete these instructions and replace with evidence e.g. a precise reference (including page and section numbers) showing compliance with that particular PQS specification requirement. In general a "pass" or "complies" statement alone is not acceptable - evidence for compliance must be provided.
One document may provide evidence for several clause requirements. Please carefully reference the sub-sections of any such document so that it is clear that the clause requirement has been met. This will help speed the pre-qualification process.
Greyed-out cells are for information and do not require a response.
NOTE: If it is not clear or doubtful whether a requirement has been met, then this is a non-compliance and pre-qualification might be delayed.
| Please also complete the VP sheet. |  | REVIEWER’S SUMMARY: |  | REVIEWER’S SUMMARY: |  | REVIEWER’S SUMMARY:
Section 7. | Product dossier | Applicant's submission | Reviewer's remarks [Name & Date] | Applicants response [Name & Date] | Reviewer's remarks [Name & Date] | Applicants response [Name & Date] | Reviewer's remarks [Name & Date]
The legal manufacturer or reseller is to provide WHO with a pre-qualification dossier containing the following:
| Dossier examination fee in US dollars. | [Please arrange payment with WHO PQS.]
| General information about the legal manufacturer, including name and address. | [Please state.]
| Unique identification reference for the logger. | [Please state.]
| Brand name of the logger. | [Please state.]
| Full specifications of the logger being offered, covering all the requirements set out in this document, including details of logger marking and traceability. | [Please list the document(s provided in the submission for pre-qualification).]
| Certified copy of the legal manufacturer’s/reseller's current ISO 9001 quality system certification. | [Please refer to a document.]
| Where relevant, certified photocopies of the legal manufacturer’s ISO 14001 certification, EMAS registration or registration with an equivalent environmental audit scheme. Conformity with an environmental audit scheme is not mandatory; however preference will be given to manufacturers who are able to demonstrate compliance with good environmental practice. | [Please refer to a document.]
| Laboratory test report(s) proving conformity with the product specifications. | [Please refer to a test report where the logger has been tested in accordance with PQS verification protocol WHO/PQS/E006/DL01-VP.3.]
Section 4. | Requirements
4.1 | General
| The logger is primarily tasked with maintaining relative time, recording appliance data objects, generating and recording logger data objects, and making that data available in a standardized way to other equipment monitoring devices and systems. | NOTE: To show compliance with many different Clause 4 requirements, it may be possible to refer to only one or a small number of downloaded data-files but precise referencing to, and annotated highlighting, will help speed the PQS review process.
| Data objects are listed as either “required” or “optional”. The logger shall have a means of receiving data inputs from the appliance’s integrated sensors and storing that data in its onboard data storage, with an associated relative timestamp. The logger shall apply its relative timestamp to data received from the appliance, as the data will not contain timestamps. | [Please list required data objects (for cross-checking purposes) and please list optional data objects for this logger.] [Please state the amount of onboard data storage (number of data objects for a number of days), and confirm that the logger will apply a relative timestamp to the data as received from the appliance.]
4.2 | Performance
4.2.1 | Relative time keeping
| As the logger is not expected to have access to absolute time, its timestamps shall be considered relative to a constant arbitrary reference. If a logger has absolute time (e.g. logging functions are integrated with remote EMD functions), it may record the absolute time in the appropriate data element, but must still satisfy the requirement for logging relative time and including it in the filenames as defined in Clause 4.5.13. | [Please state whether the logger has access to absolute time, as well as relative time (including the filenames as defined in Clause 4.5.13).]
| Relative time requirements are specified in the Data Standard specification: WHO/PQS/E006/DS01. The logger shall begin keeping relative time at the point of manufacture such that relative time is a measure of the total life of the appliance from time of manufacture. | [Please demonstrate this by reference to relative timestamps on a downloaded data-file.]
4.2.2 | Continuous time keeping
| The logger shall continue time keeping functions throughout power outages and for the life of the appliance within which it is integrated. There shall be no gap in time keeping during extended appliance supply power outages. Some recording functions of the logger may become deactivated after a period without supply power (see Clause 4.2.7), but relative time keeping shall never be disrupted by lack of power supply. | [Please state whether any recording functions of the logger are deactivated after a period without supply power. (See Clause 4.2.7).]
4.2.3 | Disruption to time keeping
| If for any reason there is a disruption to timekeeping (e.g. timekeeping battery failure), the logger shall record that a disruption occurred when normal operation is restored and continue relative timekeeping immediately after the point it left off prior to the disruption (i.e. do not reset to zero, and do not repeat timestamps). This disruption shall be recorded by updating the RTC Wakeup value (RTCW), which contains the relative timestamp of the start of the most recent continuous timekeeping operation. When logger timekeeping is resumed after any disruption, the RTCW value shall be updated with the current relative time value. See Annex 2 for more information about RTCW. | [Please show that when there is a disruption to timekeeping, the logger records this disruption and continues relative timekeeping immediately after this disruption. (See E006/DL01-VP.3 Test 8.)]
4.2.4 | Timestamp drift
| Maximum allowable time ‘drift’ for devices that are not able to sync to global clocks (i.e. local logging only) is 60 minutes per year at ambient temperatures ranging from +10°C to +43°C. | [Please show the time drift is within 60 minutes per year (at ambient temperatures ranging from +10°C to +43°C). (See E006/DL01-VP.3 Test 4.)]
4.2.5 | Required data objects
| The logger shall record all of the following data objects unless they are not applicable to the appliance (e.g. DC voltage for a mains-powered ILR). If an appliance generates and communicates to the logger data objects not included within this list, the logger shall record any manufacturer-specific data objects based on the requirements in Clause 4.2.9. | NOTE: Logging of Clause 4.2.5 required data objects can all be demonstrated in ONE referenced document or by a referenced example of a downloaded data data-file or pdf.
Administration:
| Appliance administrative information: data related to the appliance shall include manufacturer, model, manufacturer serial number, date of production, and appliance WHO PQS code and category. | [See above note. Please show that this required information is logged.]
| Compressor electronic control unit administrative information: if available, data related to the compressor(s) or other cooling control unit(s) should include manufacturer, product code, production date, and software version. | [See above note. Please show that this required information is logged.]
| Logger administrative data: data related to the logger shall include manufacturer, model, software version, manufacturer serial number, date of production, and WHO PQS code. | [See above note. Please show that this required information is logged.]
| Relative time: Strictly increasing timestamp with an arbitrary but constant reference point. (e.g. device commissioning is t=0), standardized according to ISO 8601 Durations profile, represented by the format PnDTnHnMnS, where the [n] is replaced by the value for each of the day and time elements that follow the [n]. | [See above note. Please show that this required information is logged.]
| RTC Wakeup (RTCW): Relative timestamp of the last time the logger resumed normal operation after a disruption to timekeeping (e.g. extended power failure) as described in Clause 4.2.3. | [See above note. Please show that this required information is logged.]
Environment
| Ambient temperature: temperature of the immediate surroundings within which the appliance is operating, measured in degrees Celsius (°C) at least once every 15 minutes. Care shall be taken integrating ambient temperature sensors within the appliance to minimize the effects of sensor positioning in a location where ambient temperature sensor readings could be affected by the heat generated by the appliance or other environmental factors (e.g. positioning of the fridge within a facility). Appropriate positioning of sensor will be validated through performance testing. | [See above note. Please show that this required information is logged.]
| AC supply voltage availability (AC supply appliances): data indicating the number of seconds within each 15-minute period when the AC voltage is within acceptable bounds. | [See above note. Please show that this required information is logged.]
| DC supply voltage to appliance (DC supply appliances): average DC supply voltage to appliance within each 15-minute period. Average value of raw samples collected at intervals not longer than 10 seconds. | [See above note. Please show that this required information is logged.]
Performance
| DC current drawn by the appliance (DC supply appliances): average DC current being drawn from supply by the operation of the appliance within each 15-minute period. Average value of raw samples collected at intervals not longer than 10 seconds, measured in Amps. This is intended to measure current for all appliance cooling, control and monitoring functions, and at supplier discretion may also include current for Energy Harvest Control outputs. | [See above note. Please show that this required information is logged.]
| Compressor runtime: Total duration compressor or other cooling system operated within each 15-minute period, measured in seconds. | [See above note. Please show that this required information is logged.]
| Error codes: any error codes related to appliance sensor or performance issues shall be generated and recorded by the logger when an error event occurs. Additionally, the logger shall generate and record error codes for any condition that may impair its normal operation (e.g. broken or disconnected sensors, self-test failure). | [See above note. Please show that this required information is logged.]
| Logger battery remaining: estimated number of hours of battery remaining to operate the logger normally, measured in days (DDDD). | [See above note. Please show that this required information is logged.]
| Vaccine compartment temperature: temperature of the location where vaccines may be stored, measured in degrees Celsius (°C) at least once every 15 minutes. For appliances storing vaccine above 0°C where freezing is a risk, the sensor shall be in the coldest location of the compartment (based on appliance testing). For appliances storing vaccines in frozen form, this sensor shall be at the warmest location of the compartment (based on appliance testing). | [See above note. Please show that this required information is logged.]
Use
| Door/lid open (cumulative): total duration door was opened within each 15-minute period, measured in seconds; for combined refrigerator-freezer units with separate compartments and doors/lids, each compartment shall generate distinct open duration data. For appliances with inner and outer doors/lids, only the outer door/lid needs to be monitored. | [See above note. Please show that this required information is logged. (See E006/DL01-VP.3 Test 5.)]
| Door/lid open (instantaneous): continuous duration door has been open at the time of USB mount by an EMD or other USB host, measured in seconds. If the door is closed at the time of USB mount, this value will be zero. For combined refrigerator-freezer units with separate compartments and doors/lids, each compartment shall generate distinct open duration data. For appliances with inner and outer doors/lids, only the outer door/lid needs to be monitored. This data object may be used by E-EMDs to inform door alarming. | [See above note. Please show that this required information is logged. (See E006/DL01-VP.3 Test 5.)]
| Main ON/OFF switch: indication of the status of the supply power switch of the appliance, measured when an event occurs in terms of absolute ON/OFF status. This field is required for refrigerated vehicles and transportable powered appliances that are likely to sit unused for periods of time, but optional for all others. | [See above note. Please show that this required information is logged.]
4.2.6 | Continuous recording without appliance supply power
| The logger shall continue to record data objects noted in Clause 4.2.7 during appliance supply power outages. The minimum period of uninterrupted recording without appliance supply power assumes any power storage components are fully charged at the time power is disconnected: | [Please reference the appropriate section in the test report. (See E006/DL01-VP.3 Test 4.)]
| Stationary solar-powered appliances including cold/freezer rooms: 1.5x rated autonomy time at +43°C or 96 hours at +43°C, whichever is greater. | [Please state if applicable and show evidence for the duration of continuous recording without appliance power.]
| Stationary mains-powered appliances including cold/freezer rooms: 2.0x rated holdover time at +43°C or 96 hours at +43°C, whichever is greater. | [Please state if applicable and show evidence for the duration of continuous recording without appliance power.]
| Transportable powered appliances: 1.5x rated independence time at +43°C or 18 hours at +43°C, whichever is greater. It is acceptable for continuous recording to stop when the appliance is intentionally switched off, but the continuous timekeeping requirement of Clause 4.2.2 must still be satisfied. | [Please state if applicable and show evidence for the duration of continuous recording without appliance power.]
| Refrigerated vehicles: 1.5x rated non-idle run time of the refrigeration unit. It is acceptable for continuous recording to stop when the unit is intentionally switched off, but the continuous timekeeping requirement of Clause 4.2.2 must still be satisfied. | [Please state if applicable and show evidence for the duration of continuous recording without appliance power.]
4.2.7 | Recording essential data objects during power outage
| During periods when there is insufficient power to operate the appliance, the logger shall continue to record the following essential data objects: • Vaccine compartment temperature • Ambient temperature • Door openings | [Please confirm recording of these essential data objects by referring to your evidence in Clause 4.2.6.]
4.2.8 | Generating data objects
| The following data objects are optional for appliance generation and logger recording. If any optional data object is generated by the appliance, it shall be communicated to the logger: | NOTE: Logging of Clause 4.2.8 optional data objects can all be demonstrated in ONE referenced document if or by a referenced example of a downloaded data data-file or pdf.
Administration:
| Data integrity: File data integrity data such as cyclic redundancy check (CRC), hash, or cryptographic file signature that can be used for data validation. | [See above note. Please whether this optional information is logged.]
Environment
| AC supply voltage to the appliance (AC supply appliances): Appliance input voltage from power supply (mains/generator), measured in volts (V) at least once every 15 minutes. | [See above note. Please whether this optional information is logged.]
Performance
| AC current drawn by the appliance (AC supply appliances): Current being drawn from supply by the operation of the appliance within each 15-minute period. Average value of raw samples collected at intervals not longer than 10 seconds. Measured in amps (A). | [See above note. Please whether this optional information is logged.]
| Compartment relative humidity: Internal vaccine compartment relative humidity, measured in RH% at least once every 15 minutes. | [See above note. Please whether this optional information is logged.]
| Compressor electronic unit temperature: Unit temperature as measured by the compressor electronic unit(s) or other cooling controller(s), measured in degrees celsius (°C) at least once every 15 minutes. | [See above note. Please whether this optional information is logged.]
| Compressor speed: Maximum operating speed of the compressor, measured in revolutions per minute (rpm), within each 15-minute period. | [See above note. Please whether this optional information is logged.]
| Condenser temperature: Surface temperature of the appliance’s condenser coil/circuit or other heat rejection surface measured in degrees celsius (°C) at the warmest location at least once every 15 minutes. | [See above note. Please whether this optional information is logged.]
| Fan speed: Operating speed of compressor cooling fan, measured in % of max speed at least once every 15 minutes. | [See above note. Please whether this optional information is logged.]
| Holdover, autonomy, or independence time remaining: Estimated number of days of holdover/autonomy/independence time remaining (DD.D). | [See above note. Please whether this optional information is logged.]
| Thermostat/control temperature sensor - freezer compartment: Applicable for appliances with a freezer compartment; sensor that is responsible for controlling freezer compartment refrigeration cycle based on setpoints, measured in degrees Celsius (°C) at least once every 15 minutes. | [See above note. Please whether this optional information is logged.]
| Thermostat/control temperature sensor - vaccine compartment: Sensor that is responsible for controlling vaccine compartment refrigeration cycle based on setpoints, measured in degrees Celsius (°C) at least once every 15 minutes. | [See above note. Please whether this optional information is logged.]
Use
| Door openings: Number of door openings within each 15-minute sample period. | [See above note. Please whether this optional information is logged.]
4.2.9 | Manufacturer-specific data objects
| Legal manufacturers may define and include data objects in the JSON raw data files that are not in the EMS data objects list, provided compliance with the conditions defined in the latest version of WHO/PQS/E006/DS01 and the following: • Additional temperature measurements beyond those defined in the data objects list shall be communicated as manufacturer-specific data objects. • While it is recommended that custom objects are embedded along with required objects in the 15-minute data records, if custom objects are saved as separate data records, each record must contain a relative timestamp (RELT) and RTC wakeup value (RTCW). | [Please demonstrate any other data objects (e.g. in the form of a downloaded data-file) not defined in the above clauses. (See E006/DL01-VP.3 Test 1.)]
4.2.10 | Data standard
| All required and optional data objects shall be recorded and formatted in compliance with the latest published version of the Data Standard specification: WHO/PQS/E006/DS01. | NOTE: The formatting of data objects in accordance with this data standard is a requirement.]
4.2.11 | Data recording interval
| For all data objects sampled on a regular, time-basis (e.g. temperature), data recording shall occur at least every 15 minutes and include a relative timestamp for the recording. | [For all the sub-clauses in Clause 4.2.11, show either one example (e.g. a downloaded data-file) for all the required items or five examples for each item individually.]
| For all data objects that are sampled on an event-basis (e.g. door opening), the logger shall save a record of those events in the associated time-based data object. | [For all the sub-clauses in Clause 4.2.11, show either one example (e.g. a downloaded data-file) for all the required items or five examples for each item individually.]
For example, if at minute 10 of a 15-minute sampling period the vaccine compartment door is left open for 15 minutes, the data object after opening will reflect an open duration of 5 minutes (e.g. “DORV”: 300), and the following data object will reflect an open duration of 10 minutes (e.g. “DORV”: 600).
| For the special present door/lid status data objects, e.g. IDRV & IDRF, the logger shall save a record of these events at the time of USB mount. | [For all the sub-clauses in Clause 4.2.11, show either one example (e.g. a downloaded data-file) for all the required items or five examples for each item individually.]
In the above example, if an E-EMD mounts the logger USB mass storage one minute after the logger’s sampling period, the present door opening data object will reflect an open duration of six minutes (e.g. “IDRV”: 360). When the E-EMD mounts 15 minutes later, the door is closed, so the continuous time of door opening is zero (e.g. “IDRV”: 0).
4.2.12 | Period of data recording & storage
| The logger shall be able to record and store all required, optional, and manufacturer- specific parameters for a period of at least one year, and available for download via the M2M data interface. The data storage system shall be configured such that at least one year of most recent data is preserved at all times and the oldest data is overwritten on a first in, first out basis. | [Please show that there is sufficient storage capacity for 12 months by taking the storage for 2 months data and multiplying by 6. This should already be logged before sending to lab.]
4.2.13 | Data storage
| Non-volatile, solid-state memory shall be used to avoid data loss in the event of power supply failure. | [Please state the type and tested life of the solid state memory. (See E006/DL01-VP.3 Test 1.)]
4.2.14 | Data validation
| Legal manufacturers or resellers are not required to employ methods for validating the integrity of data unless so requested or required by the employer. If implementing any form of data validation, such methods shall not interfere with the ability of external devices to access and download data on the logger. | [Please describe your method of data validation.]
4.2.15 | Sensors
| Temperature sensors: • Operating range: upper limit: +50°C, lower limit: -30°C • Accuracy: ±0.5°C or better, within the range -30°C to +20°C. • Resolution: 0.1°C in the range -30°C to +20°C. • Response time: T90 within 20 minutes maximum. | [The following can be all included in the calibration report or a separate document: • Show the measured operating range. • Measured accuracy within the range. • Measured resolution. (Note that resolution must be measured. This should not be assumed to be the "flicker" of the last digit. It is the difference in RH that can be resolved. A laboratory can measure resolution at the same time as calibration.) • Measured response time. (See Specification Clause 4.2.17.)]
| Door/lid opening sensors: • Operating range: upper limit: +50°C, lower limit: -30°C. • Accuracy: binary, open/closed. • Response time: an “open” event shall be identified whenever the door panel is not fully seated in the closed position for proper compartment sealing. | [The following can be all included in the calibration report or a separate document: • Opening sensor measured operation range. • Show that the sensor is binary - completely open or completely closed. • Measured response time.]
| Relative humidity sensors: • Operating range: upper limit: +50°C, lower limit: -30°C. • Accuracy: ±5% or better in the range 5% at 95% RH at a temperature of +5°C. • Resolution: 1%. • Response time: T63 within 60 seconds maximum. | [The following can be all included in the calibration report or a separate document: • Measured operating range of +50°C to -30°C. • Measured accuracy in the range 5% at 95% RH at a temperature of +5°C. • Measured resolution. (Note that resolution must be measured. This should not be assumed to be the "flicker" of the last digit. It is the difference in RH that can be resolved. A laboratory can measure resolution at the same time as calibration.) • Measured response time.]
| Current monitoring sensors: • Operating range: upper limit: +50°C, lower limit: -30°C. • Accuracy: ±5% in the range 0-10 A DC or ±5% in the range 0-30 A AC • Resolution: 0.1 A in the range 0-10 A DC or 0.1 A in the range 0-30 A AC | [The following can be all included in the calibration report or a separate document: • Measured operating range. • Measured accuracy in the range 0-30 A AC • Measured resolution in the range 0-30 A AC.]
| Voltage monitoring sensors: • Operating range: upper limit: +50°C, lower limit: -30°C. • Accuracy: ±2% in the range 0-50 V DC (unless the appliance is designed for higher voltages, in which case the maximum input voltage applies), or ±2% in the range 0-600 V AC. • Resolution: 0.1 V in the range 0-50 V DC or 0.1 V in the range 0-600 V AC. | [The following can be all included in the calibration report or a separate document: • Measured operating range. • Measured accuracy. • Measured resolution in the range 0-600 V AC.]
| Other sensors: Other sensors used to monitor optional data objects are permitted. The legal manufacturer shall ensure that these sensors are robust and appropriately designed for the lifetime of the appliance. | [Please explain how the design or quality procedures ensure the sensors are robust and appropriately designed for the lifetime of the appliance.]
4.2.16 | Sensor & cable fixings and replacement
| Fixings: all sensors and cables shall be fixed in place and integrated within the cabinet of the appliance to avoid inadvertent repositioning, damage, or monitoring disruption. On-site installation or positioning of sensors and cables shall not be required. | [Please show (e.g. diagram) how the site within the appliance minimizes disruption.]
| Replacement: replacement of integrated sensors shall be possible by a trained technician. Designs shall allow for sensor replacement without damaging the appliance or any sensor fixings. Designs that prevent straightforward replacement are not permitted. | [Please refer to a service manual or other document which shows sensor replacement. (See E006/DL01-VP.3 Test 1. Incorrectly referenced to Specification Clause 4.12.2.)]
4.2.17 | Calibration
| For vaccine storage chamber temperature measurement, system components are to be covered by a Certificate of Traceability and Calibration. | [Please provide certificate of calibration including a calibration report or a copy of the calibrated temperatures. The whole logger i.e. from the sensor to the data read-out must be calibrated, not just the sensor. An uncertainty estimation should be included. (See Specification Clause 4.2.15 Sensors.)]
| The traceability declaration is to confirm that the measurement standards and instruments used during calibration of the components are traceable to an ISO/IEC 17025 accredited testing laboratory, to NIST, or to another internationally recognized standards agency. The certificate shall be accompanied by a copy of the reference instrument calibration certificate. | [Please refer to a copy of the reference instrument calibration certificate in the calibration report.]
| Field calibration of sensors must not be required at any time, whether due to battery replacement in designs where the operational lifetime of the battery is shorter than the required 10-year life of the logger, or due to sensor replacement. As field calibration of sensors is not a reasonable expectation, the intent of this requirement is to maintain temperature sensor accuracy through expected field maintenance events. | [Please show an example temperature measurements before and after of a sensor replacement.]
4.2.18 | Energy storage lifetime
| Logger energy storage (e.g. electrical batteries or capacitors) is required to enable continuous data recording during times when M2M and appliance supply power is unavailable e.g. power outage for ILRs, night-time for SDDs. Various approaches to energy storage are acceptable. | [Please describe the details of your energy storage design. (See E006/DL01-VP.3 Test 1.)]
| The functional lifetime of any non-replaceable energy storage components of the logger (e.g. batteries) shall be no less than the appliance within which it is integrated, i.e. 10 years for refrigerators. Replaceable energy storage components must have a functional lifetime of no less than five years. | [What is the design lifetime of the energy storage and if there are replaceable components, please refer to the relevant section of the service manual to show how easy it is to replace any component.]
| In the event of a battery discharge or failure, the logger must perform logging functions when supply power is available to the appliance. | [Please reference the appropriate section in the test report. (See E006/DL01-VP.3 Test 8.)]
4.2.19 | Energy storage recharging
| Logger energy storage components may be rechargeable. Intelligent charging schemes should be utilized to extend the functional lifetime and safety of the energy storage components. | [Please describe how the design complies with this requirement.]
| For rechargeable energy components: If a logger’s energy storage is fully depleted, it shall have the ability to recharge sufficiently in at most eight hours - assuming power supply - to then operate without power supply for at least 48 hours. | [Please show data to confirm time to recharge and then operate for at least 48 hours. (See E006/DL01-VP.3 Test 2.)]
| Loggers may utilize long-lifetime energy storage components that do not utilize supply power for recharging. | [What type of long-lifetime non-rechargeable energy storage components are utilized by the logger?]
| Replacement of energy storage components: if the batteries are intended to be replaceable, such a task shall be possible by a trained technician on-site; shall not require recalibration of sensors. | [Please refer to the relevant section of the user manual or service manual to show how easy it is to replace batteries.]
4.2.20 | Electromagnetic compatibility
| Operation of the system and of the individual system components shall be unaffected in the normal electromagnetic compatibility environment in which the system is intended to work, considering disturbances generated by adjacent apparatus which are compliant with the requirements of IEC 61000-6-1 for immunity and IEC 61000-6-3 or IEC 61000-6-8 for emissions. Information to ensure uninterrupted use of the device shall be stated in the user instructions. | [Please provide an EMC test report to show compliance with Clause 4.2.20. (See E006/DL01-VP.3 Test 1.)]
4.3 | Environmental requirements
All components shall be adequately protected from the environment in which they will be deployed, including temperature, ambient humidity, input voltage and dust and water ingress.
4.3.1 | Ambient temperature
| All components shall tolerate high and low temperatures during transportation and storage, specifically in the range of -30°C to +70°C with the system components inactivated except for relative time keeping, which must begin at the point of manufacture as specified in Clause 4.2.1. | [Please demonstrate or provide evidence that all components can tolerate temperatures in the range of -30°C to +70°C during transportation and storage. (Such evidence could be combined with the sub-clause below.) (See E006/DL01-VP.3 Test 1.)]
| All components shall tolerate high and low temperatures during use, specifically in the range of -10°C to +55°C with the system components activated. | [Please demonstrate or provide evidence that all components can function in the range of -10°C to +55°C during use. (Such evidence could be combined with the above sub-clause.)]
4.3.2 | Ambient humidity
| During transport, storage and use, logger and M2M components may be subject to a wide range of uncontrolled ambient relative humidity (RH) from 0 up to 95% RH. The components shall be designed to remain undamaged and fully functional if exposed to these conditions in hot zone temperatures (+43°C). | [Please demonstrate or provide evidence that the logger can tolerate humidities within the range of 0 to 95%RH during transport, storage and use. (See E006/DL01-VP.3 Test 1.)]
4.3.3 | Input voltage fluctuations
| All components shall be designed to operate in unreliable and erratic power supply environments that may present risk of electrical damage to sensitive components. For AC supply (mains) devices, the appliance’s voltage protection device or system shall sufficiently protect all logger and M2M components without additional installation setup or user intervention. | [Please explain, demonstrate or provide evidence on how the logger is protected by an appropriate voltage stabilizer or similar protection hardware.]
4.3.4 | Dust and water ingress
| Logger components: There is no minimum IP rating requirement for the logger and its components given it shall be physically integrated within the appliance. However, the legal manufacturer(s) of the appliance and logger system shall take care in the design and integration of the logger within the appliance such that the design eliminates or minimizes the effects of dusty, humid environments with potential water splashing. This is especially relevant for sensors that may be subject to condensation. | [Please explain how the design and integration of the logger within the appliance such that the design eliminates or minimizes the effects of dusty, humid environments and potential water splashing. (See E006/DL01-VP.3 Test 1.)]
| Conformal coating of Printed Circuit Board Assembly (PCBA): any PCBA not protected by an IP-65 rated enclosure shall be protected from dust, water, and humidity through the use of a conformal coating. Manufacturers should consider coating guidelines in IPC document IPC-HDBK-830. | [Please show whether the PCBA is protected by an IP-65 rated enclosure or whether it is otherwise protected from dust, water, and humidity through the use of a conformal coating. (See E006/DL01-VP.3 Test 1.)]
| M2M power and data connections: the USB Type-C data and barrel power connections shall have a minimum rating of IP64 when left unconnected and when cables are connected to an E-EMD. | [Please provide the IP certificate or report for the USB Type-C data and barrel power connections when left unconnected and when cables are connected to an E-EMD.]
4.3.5 | Corrosion
| The legal manufacturer shall utilize non-corrodible and robust material (e.g. plastics or treated metal) for all components. | [Please state the component materials and their protection against corrosion. (See E006/DL01-VP.3 Test 1.)]
4.4 | Physical characteristics
4.4.1 | Overall dimensions
| Not critical. | [No requirements.]
4.4.2 | Weight
| Not critical. | [No requirements.]
4.5 | Interface requirements
Logger interfaces consist of internal interfaces to the appliance, which are mostly unspecified, a Machine-to-Machine (M2M) interface, and a file and format specification for data available on that interface. The purpose of the Machine-to-Machine (M2M) interface is to provide: • Local access to logger data for E-EMDs and other external devices utilized by local users (e.g. laptops, mobile phones). • Supply power from the appliance to external devices such as E-EMDs. • Supply power from external devices to the logger for the purpose of data access when logger power is depleted.
4.5.1 | Power supply
| The logger shall have reliable power supplies to enable continuous monitoring with and without supply power from the appliance. The following modes of power supply shall be possible: • Dedicated energy storage: the logger shall have a means of continuously recording when there is no power supply available via the appliance as specified in Clauses 4.2.6 and 4.2.7. For specific energy storage requirements, see Clauses 4.2.18 and 4.2.19. • M2M power supply: per Clause 4.5.7 Data Access During Power Outage, in the event the appliance lacks supply power and/or the logger power supply is depleted, it shall be possible to power the logger using power supplied by external devices via the M2M USB-C connection. This functionality is intended to ensure that loggers with depleted energy storage (e.g. batteries) can still be accessed by technicians or EMDs to extract data related to the last known performance conditions of the appliance. | [Please state mode(s) of power supply and how it may continuously log when there is no power available.]
| Optionally, the logger may source power directly from the appliance: • Appliance power supply: power may be supplied by the appliance to allow for operation and charging of any logger energy storage components. • Failsafe: the power supply to the logger shall not interfere with normal operation or starting of the appliance. Additionally, design of the logger shall mitigate failure modes of the logger that could cause it to draw greater loads from the appliance. | [Please state whether these optional mode(s) of power supply are available to the logger.]
4.5.2
| The M2M consists of two connections: • Data and logger power connector: USB Type-C female receptacle shall be used for data download by external devices and power supply from external devices to the logger in the event the logger’s energy storage is depleted. • Power connector: Barrel-type male plug with captive cable connected to the appliance shall be used and have the following specifications: o Type: Barrel o Sleeve diameter: 5.5 mm o Sleeve length: 9.5 mm o Pin diameter: 2.1 mm o Polarity: pin positive, sleeve negative o Cable type: captive to appliance o Cable length: 20 cm o Serviceability: captive cable shall be easily replaceable by a trained technician and available a spare part from the legal manufacturer or reseller. | [Please show compliance to the exact specification by photo and technical drawing for the - USB Type-C female receptacle - Barrel-type male plug with captive cable connected to the appliance shall be used and have the following specifications: Serviceability: Please refer to Service Manual Section to show how to replace the captive cable. (See E006/DL01-VP.3 Test 1.)]
| Under this requirement, connecting the appliance M2M data receptacle with external devices shall be achieved using male plugs on both ends of a USB-C data cable. Connecting the appliance M2M power cable with external devices shall be achieved by directly connecting the captive barrel cable and plug to the external EMD power receptacle or via a compatible extension cable that would be provided separately with the external EMD. | [Please show compliance to this clause by reference to the User Manual Section.]
4.5.3 | M2M Power transmission modalities and limits
| Both of the following power transmission modalities shall be possible: Appliance powering of E-EMDs: The M2M power connection shall provide power to external EMDs via the barrel connector when the appliance is supplied with power. The voltage shall be 5 V ± 5% (4.75 to 5.25 V). To ensure M2M power supply does not negatively affect the performance of the appliance within which it is integrated, the supply shall have a current limit of 1.1 A ± 50 mA (1.05 to 1.15 A) for E003 appliances. For DC supply appliances, the efficiency of any voltage conversion must be greater than 80% at maximum load. E001 and E002 appliances do not have a maximum output current limit. For transportable E003 appliances, the M2M power connection is not required, but is strongly recommended for devices that are recharged by connecting to AC or DC power. | [Please refer to a precise document section showing: - The measured efficiency of any voltage conversion (for DC supply appliances). - The test report showing the recorded voltage and current (including a maximum current limit of 1.1 A) from Test 6. A "pass" statement alone is not acceptable. (See E006/DL01-VP.3 Test 6.)]
Note that for the appliance the output is nominally 5.5 W and the allowable draw from an E-EMD is 5 W. This is done to avoid conflicting logic at the current limit threshold.
| External device powering appliance logger: The M2M connection shall accept power from an external device via the USB Type-C connector so that the logger data may be accessed even if the logger’s energy storage is depleted. The M2M need only support 5 V 500 mA default USB 2.0 levels required by the USB Type-C standard; the higher currents and voltages enabled by the USB Power Delivery standard are optional. | [Please show that the M2M connection has been tested to accept power from an external device via the USB Type-C connector. (An in-house test is acceptable.) (See E006/DL01.3 Clause 4.5.7.)]
4.5.4 | M2M data access
| The logger shall make all recorded data accessible to on-site users and devices via USB Type-C data connection (see Clause 4.5.2). The following communications modalities are required or optional: • Logger-EMD access (required): the M2M shall allow an external E-EMD to access and download data from the logger. • External device access (required): the M2M shall allow external devices such as laptops or mobile phones to access and download data from the logger. • Wireless logger-device communication (optional): the logger/M2M may, as an option, provide wireless access to recorded data in the logger via wireless communication systems such as Wi-Fi® or Bluetooth®. | [Please state which communications modalities are or can be available.]
All data generated and/or saved by the logger is considered to be owned by the employer, and the legal manufacturer or reseller must not restrict access to data by the employer or the employer’s representatives. No data (unprocessed or processed) collected from loggers will be shared with any third parties by the legal manufacturer or reseller without explicit and informed consent by the employer. The following are additional requirements of the logger to enable on-site data access.
4.5.5 | USB Mass Storage Device Class & uninterrupted logging
| The logger shall function as a USB Mass Storage device, complying with USB Mass Storage Device Class protocol requirements as specified by the USB Implementers Forum and formatted as FAT16 or FAT32. The choice of USB Mass Storage rather than other data communication protocols is based on wide adoption and compatibility of this transfer modality across devices and operating systems. With these benefits may come some limitations in that if the logger is mounted by an external EMD, it may not be able to record new appliance data. | [Please state the type of formatting. (See E006/DL01-VP.3 Test 1 and Test 3.)]
| Therefore, the logger shall be designed such that there is never a recording gap greater than 15 minutes due to it being mounted by an external device. For example, it could be designed with temporary memory to store recorded data while the USB is mounted, or to be able to disconnect from USB to write to its filesystem. In these cases, the logger should delay connecting over USB until file writes are complete. | [Please explain whether the logger has a temporary memory to store recorded data while the USB is mounted, or to be able to disconnect from USB to write to its filesystem.]
| The logger may force a USB disconnection if an E-EMD or other USB host has mounted for longer than 180 seconds. | [Please state whether the logger will do this.]
4.5.6 | Compatibility
| Appliances complying with this logger and M2M specification shall be compatible with all brands and models of external EMDs complying with WHO/PQS/E006/EM01 and common external devices like laptops and mobile phones. | [State the brands and models of external EMDs (complying with WHO/PQS/E006/EM01) and common external devices like laptops and mobile phones that have been check-tested at the M2M connection. (See E006/DL01-VP.3 Test 1.)]
4.5.7 | Data access during power outage
| During power outages, the logger shall still make recorded data available via the M2M to the EMD or other connected devices e.g. laptop. EMD or other connected devices shall provide power to facilitate data transmission. | [Show that the logger can make recorded data available via the M2M to an EMD or other connected devices e.g. laptop during a power outage. (An in-house test report is acceptable.) (See E006/DL01-VP.3 Test 4.)]
| In the event the logger has no appliance supply power, and its energy storage is depleted, it shall be possible for external devices to power and access the logger using external device power via the M2M USB-C connection (see Clause 4.5.3). | FOR INFORMATION. (See Specification Clause 4.5.3.)
4.5.8 | Access drivers and software
| Downloading data from the logger via the M2M to an external device (e.g. external EMD or laptop) shall not require drivers or software from another source such as the Internet, flash drive, or compact disc.] | [Please confirm that no drivers or software from any source are required to download data. (See E006/DL01-VP.3 Test 3.)]
4.5.9 | Read and write access to logger
| Read access to the logger’s recorded data is essential, and limited write access is acceptable for the purposes of firmware updates or configuration updates. The requirements below shall be satisfied: • Read access: logger data files shall be presented to the EMD or local-access device (laptop, mobile phone, etc.) as a drive that does not accept write operations (USB Mass Storage Class). • Write access: if firmware or configuration updates are supported, these shall be accomplished by methods that prevent inadvertent alteration of firmware or data stored on the logger. Some possible methods include using a separate USB communication channel as part of a USB composite device for the update, or requiring a button press by a trained technician concurrently with connecting a USB host. (See also spare parts programming requirements in Clause 4.9.1.) | [Please confirm: - logger data files do not accept write operations - if firmware or configuration updates are supported, there can be no inadvertent alteration of firmware or data stored on the logger.]
4.5.10 | Remote data transmission frequency
| The logger and M2M may utilize wireless communication modalities (e.g. Bluetooth®, Wi-Fi®, etc.) that provide users wireless access to “read-only” data. Mobile software (e.g. mobile phone apps) may be provided to enable users to access this data wirelessly on-site. (Any hardware utilized to enable wireless data shall be built within the appliance without any components, such as antennas, exposed on the surface of the appliance that may become damaged over time.) | [Please refer to the user manual Section that shows how to access data via wireless communication. (See E006/DL01-VP.3 Test 1.)]
4.5.11 | Logger analysis software
| If a manufacturer of the logger or third-party developer provides software or an app for analysing logger data, this software shall be available for download from the Internet via the manufacturer or provider’s website or app store. | [Please state where this software or app can be downloaded. (See E006/DL01-VP.3 Test 1.)]
4.5.12 | Folder and file structure
| When mounted by a USB host, the logger shall present as a USB Mass Storage Device with each of the following available in the root directory: | COMPLETE AS BELOW:
| • User-readable 60-day aggregated summary report (detailed in Clause 4.5.14), in Portable Document Format (PDF). | [Please refer to a PDF example of this summary.]
| • Raw data file containing all required and optional raw data parameters, in JavaScript Object Notation format (JSON) for the most recent 0- to 60-day period. Raw data files must follow the published JSON schema referenced in the PQS data standard (WHO/PQS/E006/DS01) that is published on the PQS E006 “category documentation” web page. | [Please refer to an example of downloaded raw data.]
| • Synchronization data file containing the relative timestamp at time of mount, RTC Wakeup value (RTCW), present door opening data objects (IDRV and/or IDRF as appropriate), and present compartment temperature data objects (TVC and/or TFRZ as appropriate), as well as most recent full data record from the raw data file in JSON format. Synchronization file must follow the published JSON schema referenced in the PQS data standard (WHO/PQS/E006/DS01) that is published on the PQS E006 “category documentation” web page. | [Please refer to an example of a synchronization data file containing the relative timestamp at time of mount, RTC Wakeup value (RTCW), present door opening data objects (IDRV and/or IDRF as appropriate), and present compartment temperature data objects (TVC and/or TFRZ as appropriate), as well as most recent full data record from the raw data file in JSON format.]
| • Folder named “DATA_HISTORY”: This folder is where historic raw data files are located, for access to the minimum 1-year record of required and optional data parameters. | [Please refer to the user manual Section that shows how to access historic raw data files.]
The root directory may contain other files like manuals or diagnostics information but should contain no more than 10 visible files/folders, to prevent user confusion.
4.5.13 | File naming conventions
| The files required in Clause 4.5.12 shall be named based on the following conventions: | COMPLETE AS BELOW: (See E006/DL01-VP.3 Test 7.)
| 60-day aggregated summary report filename convention: The filename for the user-readable 60-day aggregated summary report shall be structured as: 0123456789_60DTR_SUMMARY_PnDTnHnMnS.pdf Where 0123456789 is to be replaced by the alphanumeric logger unique identifier, and PnDTnHnMnS is to be replaced by the relative timestamp as of the time of USB mount in ISO 8601 duration format. (See Clause 4.5.14 for Summary Report format and content requirements.) | [Please refer to the data download example that confirms this formatting.]
| Active logging raw data file filename convention: There shall be only one CURRENT_DATA file in the root directory. The filename for the most current raw data file shall be structured as: 0123456789_CURRENT_DATA_PnDTnHnMnS.json Where 0123456789 is to be replaced by the alphanumeric logger unique identifier, and PnDTnHnMnS is to be replaced by the relative timestamp as of the time of USB mount in ISO 8601 duration format. | [Please refer to the data download example that confirms this formatting.]
| Historic raw data file(s) filename convention: The filename for historic raw data files shall be structured as: 0123456789_DATA_PnDTnHnMnS.json Where 0123456789 is to be replaced by the alphanumeric logger unique identifier, and PnDTnHnMnS is to be replaced by the relative timestamp associated with the most recent data object stored in the file, in ISO 8601 duration format. | [Please refer to the data download example that confirms this formatting.]
| Synchronization file filename convention: There shall be only one SYNC file in the root directory. The filename for the synchronization filename shall be structured as: 0123456789_SYNC_PnDTnHnMnS.json Where 0123456789 is to be replaced by the alphanumeric logger unique identifier, and PnDTnHnMnS is to be replaced by the relative timestamp as of the time of USB mount in ISO 8601 duration format. | [Please refer to the data download example that confirms this formatting.]
4.5.14 | User-readable 60-Day aggregated summary reports
| A summary report in PDF format shall be automatically generated and available for download whenever an EMD or another external device mounts the logger via the M2M. The report shall include data from the last 60 days. The PDF report shall have all columns visible on the same page horizontally. If there are multiple pages, each page should at a minimum contain the appliance serial number and report creation time. | COMPLETE AS BELOW: (See E006/DL01-VP.3 Test 7.)
| The report shall provide the following administration information: • Appliance PQS Code • Appliance Serial Number • Report creation time (relative time) • Alarm threshold information | [Please refer to the data download example that confirms this information is clear on the document.]
| The report shall provide data related to environment, performance and use in a tabular format organized with rows representing 24-hour periods and columns listing the following relevant data, with the most recent days at the top. It is acceptable for the first listed day to either be a partial “in progress” day or for it to represent the most recent 24-hour period. Dates and times listed in the report may be represented in absolute time, but the file naming convention must be followed as specified in Clause 4.5.13. The report must include at a minimum the following (summarized and recorded on a 24-hour basis), with additional information optionally added by suppliers: | [Please refer to the data download example that confirms this formatting.]
| • Average vaccine chamber storage temperature (°C) • Status • Low temperature o Minimum temperature (°C) o Total time below limit (hh:mm) e.g. 2°C for refrigerators. o Total duration the vaccine chamber is in a low-temperature alarm state (hh:mm) • High temperature o Maximum temperature (°C) o Total time above limit (hh:mm) e.g. 8°C for refrigerators. o Total duration the vaccine chamber is in a high-temperature alarm state (hh:mm) • Number of door openings • Cumulative total door open time in minutes per day • AC power availability (time ON within 24-hour period, in hh:mm). For AC supply appliances, this is based on the ‘AC supply voltage availability’ data object. For DC supply appliances, this column is not required. • Cumulative compressor run time (time ON within 24-hour period, in hh:mm) • Average ambient temperature (°C) • Error codes (A sample ‘Summary Report’ file for ILRs is included in Annex 1.) | [Please refer to the data download example that confirms this information is clear on the document. (See E006/DL01-VP.3 Clause 5.3.8.)]
4.6 | Human factors
4.6.1 | M2M location on appliance
The M2M connections shall be permanently fixed to the appliance in a location that can be easily accessed by users and technicians for the purpose of connecting external EMDs and downloading data, but while also avoiding connected wires from obstructing the use or service of the appliance. Guidance relevant to different appliance categories provided below:
| • E001 (Walk-in Cold/Freezer Rooms): The interface shall be located where the main controls of the appliance are located. All connections shall be no less than 0.5 meters from the ground. | [Please refer to a photo or diagram showing the precise position of the M2M interface.]
| • E002 (Refrigerated Vehicles): The interface shall be positioned where the main controls of the cooling system are located. | [Please refer to a photo or diagram showing the precise position of the M2M interface.]
| • E003 (Stationary and Transportable Refrigerators & Freezers): The interface shall be located on the appliance such that the data and power connections are in the same location, are easily accessible by users without tools and any connected wires to external devices do not interfere with the general use and access of the appliance and its storage compartment(s). For stationary appliances, all connections shall be no less than 0.3 meters from the ground. | [Please refer to a photo or diagram showing the precise position of the M2M interface.]
4.6.2 | M2M port markings
| The M2M USB-C data connection receptacle shall be clearly labelled with “Logger Data Access” such that a user or technician understands that the port is for accessing data on the appliance’s logger and not for the charging of mobile phones or other devices. | [Please provide a photo of the label next to the USB-C port. (Can be one photo with the clause below.) (See E006/DL01-VP.3 Test 1.)]
| The M2M barrel power connection plug shall be clearly labelled with “5 V Monitoring Power” such that a user or technician understands that the cable and plug is for powering monitoring devices and not for the charging of mobile phones or other devices. | [Please provide a photo of the label next to the M2M barrel power connection plug. (Can be one photo with the clause above.) (See E006/DL01-VP.3 Test 1.)]
| All labels shall be sufficiently durable and robust such that they remain readable for the lifetime of the appliance. | NOTE: This is a requirement.
4.7 | Materials
4.7.1 | Ozone depleting chemicals
| During manufacture and assembly of the printed circuit boards and final assembly of the product, do not use any substance included in Annex A, B or C of the Montreal Protocol. | [Please state the printed circuit board materials and whether they comply with the Montreal Protocol. (See E006/DL01-VP.3 Test 1.)]
4.7.2 | Other restricted materials
| The product and its constituent components, including batteries, shall not contain lead, mercury, cadmium, hexavalent chromium, polybrominated biphenyls (PBB) or polybrominated biphenyl ethers (PBDE). | [Please provide a RoHS report or equivalent. (See E006/DL01-VP.3 Test 1.)]
4.8 | Warranty
| The logger and M2M components shall be covered by a replacement warranty in the event of any component failure arising from defective design, materials, or workmanship. The warranty of all logger and M2M components shall meet or exceed the warranty offered with the appliance. | [Please state or refer to a document showing the warranty with terms and definitions. (See E006/DL01-VP.3 Test 1.)]
4.9 | Maintainability
| The logger and M2M connections shall be designed for minimal routine maintenance over the course of its designed lifetime, which shall match the designed lifetime of the appliance within which it is integrated. | [Please refer to the User Manual Section to show routine maintenance. (See E006/DL01-VP.3 Test 1.)]
| The appliance-logger-M2M system shall also be designed to allow for ease of replacement and repair by a skilled technician whereby key components are easily accessible and removable with simple and relatively available tools. | [Please refer to the Service Manual Section or other document to show ease of replacement.]
4.9.1 | Spare parts
| The legal manufacturer or reseller shall make available all spare parts necessary to repair and service the logger and M2M system. | [Please refer to the User Manual Section or Technician Manual Section or a document to show a list of spare parts. (See E006/DL01-VP.3 Test 1.)]
| If the legal manufacturer or reseller preloads appliance administrative data about the appliance or logger, replacement logger spare parts shall also be configurable to store the same information as was stored in the original integrated logger. This reconfiguration of replacement logger components shall be done on-site at the time of the repair or by other means such that appliance administrative information is available on the new logger after the replacement of original logger components. | [Please refer to the User Manual Section or Technician Manual Section or a document to show reconfiguration of replacement logger.]
4.10 | Disposal and recycling
| The legal manufacturer shall provide information to the buyer on the hazardous materials contained within the components and suggestions for resource recovery/recycling and/or environmentally safe disposal as part of the recovery and disposal information provided for the associated cold chain equipment. For the legal manufacturers based in the European Union, WEEE compliance in accordance with European Union Directive 2002/96/EC is mandatory. | [Please refer to the User Manual Section or to a document explaining what hazardous materials are contained within the system and show suggestions for resource recovery/recycling and/or environmentally safe disposal. (For the European Union WEEE compliance in accordance with European Union Directive is mandatory.) (See E006/DL01-VP.3 Test 1.)]
4.11 | Instructions
| Appliances complying with this specification must include guidance on the logger and M2M interface in the appliance’s user manual and technician manual in Arabic, English, French, Mandarin Chinese, Russian and Spanish. An English version of all instructions and manuals are required to be supplied at time of laboratory testing and prequalification. Instructions should include easy to understand visuals whenever possible to avoid reliance on text. | [(See E006/DL01-VP.3 Test 1.)]
| The user manual shall include the following information: | COMPLETE AS BELOW:
| • Health and safety guidance | [Pleaser refer to the User Manual Section that shows health and safety guidance.]
| • Basic operations description | [Pleaser refer to the User Manual Section that shows a description of basic operations.]
| • Warranty and contact information for legal manufacturer or reseller. | [Please refer to the user manual Section that shows warranty and contact information for legal manufacturer or reseller.]
The technician manual shall include the following information:
| Health and safety guidance | [Please refer to the Technician Manual Section that shows health and safety guidance.]
| Detailed operations description | [Please refer to the Technician Manual Section that shows detailed operations description.]
| Correct handling to avoid logger, M2M or appliance damage and for the safety of handling persons | [Please refer to the Technician Manual Section that shows correct handling.]
| Installation procedures | [Please refer to the Technician Manual Section that shows how an external device including an E-EMD may be connected both for data access and for power.]
| Preventive maintenance tasks (e.g. daily, weekly, and monthly) | [Please refer to the Technician Manual Section that shows preventive maintenance tasks (e.g. daily, weekly, and monthly).]
| Corrective maintenance procedures | [Please refer to the Technician Manual Section that shows corrective maintenance procedures.]
| Diagnostic and repair procedures | [Please refer to the Technician Manual Section that shows diagnostic and repair procedures.]
| Itemised list of spare parts including part numbers | [Please refer to the Technician Manual Section that shows itemised list of spare parts including part numbers.]
| End-of-life resource recovery and recycling procedures | [Please refer to the Technician Manual Section that shows end-of-life resource recovery and recycling procedures.]
| User training guidance. | [Please refer to the Technician Manual Section that shows user training guidance.]
| Printed user-operations and maintenance instructions specifically directed at the health centre staff shall be pictorial. All key information is to be summarized on a single pictogram sheet that may be fixed onto the appliance. The sheet is to be sufficiently durable to last the life of the appliance. | [Please provide a copy of the summarized pictogram sheet with the information as required by this clause.]
| All instruction manuals shall be supplied in printed format and easily available and downloadable in digital format on the Internet. Digital manual materials may also be provided with the appliance-logger-M2M system on a separate USB storage drive or within the file system of the logger. | [Please provide a copy (in English) of the user and any other manuals. State where instruction manuals can be easily available / downloadable and whether they provided on the a separate USB storage drive.]
4.12 | Training
Printed user-operations and maintenance instructions specifically directed at the health
5. | centre staff shall be pictorial. All key information is to be summarized on a single
pictogram sheet that may be fixed onto the appliance. The sheet is to be sufficiently
| If requested, the legal manufacturer or reseller shall at a minimum provide remote training in the countries where their product is deployed via suitable teleconferencing tools. Remote training shall cover the installation, on-site use, maintenance of the hardware, and download of data via the M2M data interface. | [Please state how remote training could be available. (See E006/DL01-VP.3 Test 1.)]
| Training programs shall cover the minimum requirements described in the following sections: | COMPLETE AS BELOW:
4.12.1 | Installation and user training
| Practical, hands-on installation training shall be available for the installation of the appliance-logger-M2M system. | [Please state how the customer can access installation training for the appliance-logger-M2M system.]
Training outcomes for participants:
| • Correct installation and setup. • Commissioning and handover of installed equipment in the form of job completion checklists and/or reports. | [Please confirm your training will cover these items: (See E006/DL01-VP.3 Test 1.)]
Minimum training requirements:
| • Unpacking of the equipment and verification that the correct equipment components and accessories are available for successful installation. • Use and interpretation of installation manuals. • Hands-on assembly and installation of different components, such as M2M cables. • Use of the logger and M2M including connecting external devices, accessing data, downloading data, disconnecting external devices. • Information on types of files and folders available on logger. • Interpretation of recorded data and summary report. • Daily, weekly, and monthly tasks and best practices for temperature monitoring, data accessing, downloading, analysis, and corrective action. • System functionality and troubleshooting. • Standard commissioning procedures. | [Please confirm your training will cover all these items: (See E006/DL01-VP.3 Test 1.)]
4.12.2 | Technical training
| Training in correct installation of the appliance-logger-M2M system, system administration, setup, operation, troubleshooting, maintenance, interpreting of data, and training of further technicians, and system administrators shall be available to employers. | [Please confirm that this training shall be available to employers.]
Training outcomes for participants:
| • Correct installation, setup, and operation of the appliance-logger-M2M system. • Commissioning and handover of installed equipment in the form of job completion checklists and/or reports. • Correct interpretation of recorded data and summary reports. • Advanced system maintenance, troubleshooting, and repair. • Training of further technicians on the installation, setup, and operation of the appliance-logger-M2M system. | [Please confirm your training will cover all these items: (See E006/DL01-VP.3 Test 1.)]
Minimum training requirements:
| • Demonstrate the appliance-logger-M2M system, including main components, controls, connections, and functionalities. • Demonstrate the process of installing and setting up the appliance-logger-M2M system. • Use of the appliance-logger-M2M system, including connecting external devices, accessing data, downloading and exporting of data, disconnecting external devices. • Interpretation of recorded data and summary reports. | [Please confirm your training will cover all these items:]
4.12.3 | System administrator training
| Training in system administration, setup, operation, troubleshooting, maintenance, interpreting of data, and training of further system administrators shall be available to employers. | [Please confirm that this training shall be available to employers.]
Training outcomes for participants:
| • Correct set up and operation of the appliance-logger-M2M system. • Correct interpretation of recorded data and summary reports, including any software tools that are available for aggregating data or transmitting data to logistics information management systems. • Training of further system administrators in the setup and operation of the appliance-logger-M2M system. | [Please confirm your training will cover all these items:]
Minimum training requirements:
| • Demonstrate the appliance-logger-M2M system, including main components, controls, connections, and functionalities. • Use of the appliance-logger-M2M system, including connecting external devices, accessing data. downloading and exportation of data, disconnecting external devices. • Interpretation of recorded data and summary reports. • Daily, weekly, and monthly tasks and best practices for temperature monitoring, data accessing, downloading, analysis, and corrective action. • Standard corrective and preventative maintenance tasks. • Standard commissioning procedures. • Frequently asked questions. • Processes and best practices for training of further technicians in the installation, setup, and operation of the appliance-logger-M2M system. • Processes and best practices for training of further system administrators in the setup and operation of the appliance-logger-M2M system. | [Please confirm your training will cover all these items: (See E006/DL01-VP.3 Test 1.)]
4.12.4 | Training materials
To support training outlined above, the following materials shall be available:
| • Appliance-logger-M2M system installation & commissioning: o Training materials on appliance-logger-M2M system installation and setup. o Training materials on standard commissioning procedures, including the use of commissioning and/or functionality checklists. | [Please provide training materials covering these topics. (See E006/DL01-VP.3 Test 1.)]
| • Appliance-logger-M2M system use and troubleshooting: o Training materials on appliance-logger-M2M system operation, including connecting external devices, accessing data, downloading data, disconnecting external devices. o Training materials on interpretation of recorded data and summary reports. o Training materials on basic and advanced system troubleshooting and maintenance procedures. | [Please provide training materials covering these topics.]
| • Materials for training trainers: o Training materials on how to conduct training of further technicians. o Training materials on how to conduct training of further system administrators. | [Please provide training materials covering these topics.]
| All training materials shall be supplied in digital format and easily available and downloadable on the Internet, in a format easy to print on standard A4 letter paper. Digital manual materials may also be provided with the appliance-logger-M2M system on a separate USB storage drive. | [Please provide a link to the training materials.]
4.13 | Verification
| In accordance with PQS verification protocol WHO/PQS/E006/DL01-VP.3. | [Please refer to the test report detailing results in accordance with WHO/PQS/E006/DL01-VP.3 and see "VP Sheet" in this submission template.]
5.0 | Packaging
| Loggers are intended to be integrated into appliances by the appliance’s legal manufacturer. Appliance packaging is covered by the relevant appliance specification. Packaging of loggers for shipment to appliance’s legal manufacturer is outside the scope of this specification. | NO REQUIREMENT.

## VP
E006 Submission to PQS for approval: Equipment Monitoring Devices for Equipment Monitoring Systems |  |  | Brand & Model:
| PQS verification protocol: | E006/DL01-VP.3
| To verify compliance with PQS specification: | E006/DL01.3
Instructions in [square brackets] are for your guidance. Please delete these instructions and replace with evidence or a precise reference (including page and section numbers) showing compliance with that particular PQS specification requirement. In general a "pass" or "complies" statement alone is not acceptable.
One document may provide evidence for several clause requirements. Please carefully reference the sub-sections of any such document so that it is clear that the clause requirement has been met. This will help speed the pre-qualification process.
Greyed-out cells do not require a response.
NOTE: If it is not clear or doubtful whether a requirement has been met, then this is a non-compliance and pre-qualification might be delayed.
Please ensure the Spec sheet is completed.
Section 7. | VP Clause | Applicant's submission | Assessor's remarks [Name & Date] | Applicants response [Name & Date] | Assessor's remarks [Name & Date] | Applicants response [Name & Date] | Assessor's remarks [Name & Date]
5.1 | Evidence of conformity assessment
| Appliances containing loggers shall carry the CE mark, UL mark and/or equivalent internationally accepted evidence of conformity assessment. Appliances that are not ordinarily marked (e.g. cold rooms, refrigerated vehicles) are exempt, but relevant components should have evidence of conformity assessment. Certifications from accredited tests labs are preferable. If conformity is self-declared, then a technical construction file shall be provided as part of the dossier to validate conformity. | [Please provide evidence of any conformity mark if the appliance has one.]
5.2 | Number of samples
| The legal manufacturer or reseller must supply the testing laboratory with a full duplicate set of the Product Dossier already supplied to WHO in accordance with the requirements of specification Clause 7. Two samples are to be provided as described below. | [Please provide the testing laboratory a draft dossier for checking and amendment. This will help to ensure that the dossier is complete.]
5.3.2 | Test 1: Type examination
| The test report must clearly show that the requirements of each section and sub-section in the type examination are individually and clearly met. | [Please refer to the test report and see the specification sheet in this template where precise references to each specification requirements are required. (Test 1 requirements refer to many specification clause requirements. See E006/DL01.3.)]
5.3.3 | Test 2: Logger power wakeup and recharging | The information requested below can be shown in one downloaded data file or in several downloaded data files as applicable so that the evidence is convincing to the PQS reviewer. The submission to PQS can helped with annotated notes. (See E006/DL01.3 Clause 4.2.6.)
| Fully depleted logger without supply power was “awoken” by the connection of external computer and made its contents accessible. | [Please refer to the precise section in the test report showing awakened logger and downloaded data files. A "pass" statement alone is not acceptable. (See E006/DL01.3 Clause 4.5.7.)]
| Logger successfully recorded 48 hours of continuous data following eight hours of charging. | [Please refer to a precise section in the test report showing the recorded data (e.g. in the form of a downloaded data file) for 48 hours. A "pass" statement alone is not acceptable. (See E006/DL01.3 Clause 4.2.19.)]
5.3.4 | Test 3: Logger data access
| Logger mounts as a FAT16 or FAT32 USB Drive | [Please refer to the precise section in the test report which shows the USB file properties (e.g. in the form of a screenshot) of the data download. A "pass" statement alone is not acceptable. (See E006/DL01.3 Clause 4.5.5.)]
| User able to access and download all files on logger without need for drivers or software. | [Please refer to the precise section in the test report showing a downloaded data file with independent confirmation by the tester that no driver or software was necessary. (See E006/DL01.3 Clause 4.5.8.)]
| User unable to perform write-access functions, including deleting, renaming, adding files and folders. | [Please refer to the precise section in the test report showing user unable to perform write-access functions, including deleting, renaming, adding files and folders; with independent confirmation by the tester. (See E006/DL01.3 Clause 4.5.9.)]
5.3.5 | Test 4: Endurance of timekeeping and recording functions, including timestamp drift
The evidence requested below can be shown as one file if this is applicable.
| Ability to access logger and download data without appliance supply power. | [Please refer to the precise section in the test report showing an example data file downloaded during this test without appliance supply power. (See E006/DL01.3 Clause 4.5.7.)]
| Verification that logger has continued to save required data objects through the period of power disconnection. | [Please refer to the above suggestion as per the previous subclause to demonstrate this pass criteria requirement. A "pass" statement alone is not acceptable. (See E006/DL01.3 Clause 4.5.7.)]
| Absolute value of timestamp drift calculated in step 8 is less than 114 parts per million (ppm), translating to less than 60 minutes of timestamp drift per year. | [Please refer to a precise section in the test report showing the measured timestamp drift. A "pass" statement alone is not acceptable. (See E006/DL01.3 Clause 4.2.4.)]
5.3.6 | Test 5: Door opening | The evidence requested below can be shown as one file if this is applicable.
| Instantaneous door opening durations IDRV (and/or IDRF for a freezer compartment) match the length the door was open at the time of USB mount in Step 1. | [Please refer to a precise section in the test report showing the recorded data (e.g. in the form of a downloaded data file) for the door open event. A "pass" statement alone is not acceptable. (See E006/DL01.3 Clause 4.2.5.)]
| Logged door opening durations DORV (and/or DORF for a freezer compartment) match the total length the door was open. | [Please refer to the above suggestion as per the previous subclause to highlight this verification. A "pass" statement alone is not acceptable. (See E006/DL01.3 Clause 4.2.5.)]
5.3.7 | Test 6: M2M load safety
| M2M output voltage is in the range 4.75 V to 5.25 V during the entire test. | [Please refer to a precise section in the test report showing the recorded voltage and current for Test 6. A "pass" statement alone is not acceptable. (See E006/DL01.3 Clause 4.5.3.)]
| M2M output current limit is in the range 1.05 A to 1.15 A. | [As above.]
5.3.8 | Test 7: Logger data review | The evidence requested below can be shown as one file if this is applicable. (See E006/DL01.3 Clause 4.5.13.)
| Data file naming conventions comply with specifications. | [Please refer to a precise section in the test report or other document showing compliance with this requirement. A "pass" statement alone is not acceptable.]
| Recorded data complies with specifications in terms of completeness of required data objects, formatting and recording period. | [Please refer to a precise section in the test report or other document showing compliance with this requirement. A "pass" statement alone is not acceptable.]
| Summary report complies with specifications. | [Please refer to a precise section in the test report or other document showing summary report. A "pass" statement alone is not acceptable.]
| Logger has sufficient storage to record at least one year of administrative, environmental, performance, and use data. | [From the amount of data recorded over the test period, estimate the storage space required for one year. Show that the storage space available is at least as large as this calculation. (See E006/DL01.3 Clause 4.5.14.)]
5.3.9 | Test 8: Disruption to timekeeping | The evidence requested below can be shown as one file if this is applicable.
| Fully depleted logger without supply power was “awoken” by the connection of external computer and made its contents accessible. | [Please refer to a precise section in the test report or other document showing compliance with this requirement (e.g. a downloaded data file). A "pass" statement alone is not acceptable. (See E006/DL01.3 Clause 4.5.7.)]
| Logger records data when supplied with power. | [Please refer to a precise section in the test report or other document showing compliance with this requirement (e.g. a downloaded data file). A "pass" statement alone is not acceptable. (See E006/DL01.3 Clause 4.2.3.)]
| RELT timestamps in data file are strictly increasing. | [Please refer to a precise section in the test report or other document showing compliance with this requirement (e.g. a downloaded data file). A "pass" statement alone is not acceptable. (See E006/DL01.3 Clause 4.2.3.)]
| RTCW values are constant during periods of continuous timekeeping and change when timekeeping is disrupted. The value recorded after a disruption to timekeeping corresponds to the RELT timestamp when operation resumed (the value is after the last RELT timestamp before operation resumed). | [Please refer to a precise section in the test report or other document showing compliance with this requirement (e.g. a downloaded data file). A "pass" statement alone is not acceptable. (See E006/DL01.3 Clause 4.2.3.)]
5.4 | Test criteria for qualification
| A final report shall be issued after all testing is complete. The report of the tests shall contain the following data and analyses: | (See E006/DL01.3 Clause 4.13.)
| Summary: Conclusions and recommendations. | [Please refer to the precise section in the report.]
| Test 1: Comments on samples received, tabulated data, compliance with required specifications (plus supporting documentation) and photographs of samples. | [Please refer to the precise section in the report.]
| Test 2: Results of logger power wakeup and recharging test. | [Please refer to the precise section in the report.]
| Test 3: Results of logger data access test. | [Please refer to the precise section in the report.]
| Test 4: Results of endurance of timekeeping and recording functions test. | [Please refer to the precise section in the report.]
| Test 5: Results of door opening test. | [Please refer to the precise section in the report.]
| Test 6: Results of M2M load safety test. | [Please refer to the precise section in the report.]
| Test 7: Results of logger data review. | [Please refer to the precise section in the report.]
| Test 8: Results of disruption to timekeeping test. | [Please refer to the precise section in the report.]
| Annexes: A preapproved test protocol verifying that the procedures set out in this document have been followed. Description of the test apparatus. Test chamber temperature and humidity records. Copies of reference thermometer and humidity measurement calibration certificate(s). Additional supporting documentation requested and received from the legal manufacturer or reseller during the course of the type-testing. | [Please refer to the precise section in the report.]