## Spec
E006 Submission to PQS for approval: Equipment Monitoring Devices for Equipment Monitoring Systems |  |  | Brand & Model:
| Compliance with PQS specification: | E006/EM01.2
| As shown by testing in accordance with PQS verification protocol: | E006/EM01-VP.2
ASSESSOR’S SUMMARY:
Instructions in [square brackets] are for your guidance. Please delete these instructions and replace with evidence or a precise reference (including page and section numbers) showing compliance with that particular PQS specification requirement. In general a "pass" or "complies" statement alone is not acceptable.
One document may provide evidence for several clause requirements. Please carefully reference the sub-sections of any such document so that it is clear that the clause requirement has been met. This will help speed the pre-qualification process.
Greyed-out cells do not require a response.
NOTE: If it is not clear or doubtful whether a requirement has been met, then this is a non-compliance and pre-qualification might be delayed.
| Please also complete the VP sheet. |  | REVIEWER’S SUMMARY: |  | REVIEWER’S SUMMARY: |  | REVIEWER’S SUMMARY:
Section 7. | Product dossier | Applicant's submission | Reviewer's remarks [Name & Date] | Applicants response [Name & Date] | Reviewer's remarks [Name & Date] | Applicants response [Name & Date] | Reviewer's remarks [Name & Date]
The legal manufacturer or reseller is to provide WHO with a pre-qualification dossier containing the following:
| Dossier examination fee in US dollars. | [Please arrange payment with WHO PQS.]
| General information about the legal manufacturer, including name and address. | [Please state.]
| Unique identification reference for the EMD. | [Please state.]
| Brand name of the EMD. | [Please state.]
| Full specifications of the EMD being offered, covering all the requirements set out in this document, including details of EMD marking and traceability. | [Please list the document(s provided in the submission for pre-qualification).]
| Certified photocopies of all type-approvals obtained for E-EMDs or appliances containing I-EMDs, such as CE marking etc. | [Please refer to document(s).]
| Certified copy of the legal manufacturer’s current ISO 9001 quality system certification. | [Please refer to a document.]
| Where relevant, certified photocopies of the legal manufacturer’s ISO 14001 certification, EMAS registration or registration with an equivalent environmental audit scheme. Conformity with an environmental audit scheme is not mandatory; however preference will be given to manufacturers who are able to demonstrate compliance with good environmental practice. | [Please refer to a document.]
| Evidence of conformity with ISO/IEC 27001:2013 for remote data systems, including evidence of independent security audit completed within last two years. | [Please refer to a document.]
| Laboratory test report(s) proving conformity with the product specifications. | [Please refer to a test report where the EMD has been tested in accordance with PQS verification protocol WHO/PQS/E006/EM01-VP.2.]
| Indicative cost of the product per unit, per 10 units and per 100 units, including coolant-packs, EXW (Incoterms 2010). | [Please state or refer to a document.]
| Indicative costs for cellular data transmission (range is acceptable), data hosting, and portal subscription for the default three-year period. | [Please state or refer to a document.]
Section 4. | Requirements
4.1 | General
| Many of the following requirements are specified based on whether the EMD is integrated or external to the appliance. They apply to both local-only and remote communication EMDs. | [State whether the EMD is integrated or external to the appliance and whether it is a Level 2 or Level 3 EMD.]
4.1.1 | Compatibility of external EMDs
| External EMDs shall be compatible with all brands of appliances complying with the logger and M2M specification. Accordingly, appliances complying with the logger and M2M specification shall be compatible with all brands of external EMDs and common external devices like laptops and mobile phones. Reference logger and M2M specification for appliance-device connections (WHO/PQS/E006/DL01) and the CCE Data Standard (WHO/PQS/E006/DS01). | [Please state whether the external EMD is compatible to all brands of appliances and laptops and mobile phones. State any limitations in this respect.]
4.1.2 | Field-fitted external EMDs
| Legal manufacturers or resellers may prequalify an external EMD against both this specification and the RTMD specification for the purpose of providing flexible solutions that may utilize the M2M data and power connections and also field-fitted sensors for monitoring appliances without a logger and M2M. The requirements of both specifications shall be satisfied and verified through both verification protocols. Legal manufacturers or resellers interested in pursuing such an approach are encouraged to consult with PQS early on in development. | [Please state whether you are seeking pre-qualification for your EMD solely against this specification (E006/EM01.2) or also against the RTMD specification (E006/TR03.2).]
4.2 | Performance
4.2.1 | Absolute time keeping
| Time format shall comply with the latest version of the Data Standard specification: WHO/PQS/E006/DS01. • Local-only EMDs: If the EMD does not have the means or ability to sync its time via remote communication networks, absolute time keeping is not required. • Remote EMDs: The EMD shall maintain absolute time synced via remote communication networks and server clocks. Syncing shall be automatic and frequent enough so that drift time kept on the EMD differs from global time by no more than 1 minute. | [Please state exact compliance - local only or remote.]
4.2.2 | Absolute time drift
| Drift in the absolute time on the EMD relative to the server time shall be less than 60 minutes per year at ambient temperatures ranging from +10°C to +43°C. | [Please state or refer to the measured time drift.]
4.2.3 | Local time
| The EMD shall have the ability to be programmed based on the local time zone within which it is deployed. Any on-site display of timestamps shall show local time if such local programming has occurred. | [Please refer to the user manual to show how it is possible to change the time zone.]
4.2.4 | Absolute timestamping at time of data download from logger
| The E-EMD shall record an absolute timestamp at the time of data receipt from logger such that the logger’s relative time data may be converted to UTC time within remote data systems. | [Please refer or show how the logger’s relative time data may be converted to UTC time (within remote data systems) e.g. via a dashboard.]
| At the time of initiating the USB mount of the logger, the E-EMD shall record an absolute timestamp, and shall append this timestamp to the end of the filename of all files downloaded from the logger, separated by an underscore. This permits downstream server systems and data consumers to associate the E-EMD absolute timestamp with the correct logger relative timestamp. The downloaded files with the updated filename are then transmitted unmodified to server infrastructure. For example, a “sync” file after being downloaded from a logger will have an updated filename with the following format: 0123456789_SYNC_PnDTnHnMnS_YYYYMMDDThhmmssZ.json | [Please provide a description of the timestamping and data transmission process.]
| It is recommended that the E-EMD also record and transmit a JSON record associated with each logger download event which contains the logger ID, logger relative timestamp, and the EMD absolute timestamp associated with the USB mount time of that download event. This could be transmitted in a standalone EMD-generated file, or directly as a JSON payload. | [Please state whether a JSON record is transmitted in a standalone EMD-generated file, or directly as a JSON payload in accordance with this recommendation.]
4.2.5 | Data management.
| All required and optional data objects shall be recorded and formatted in compliance with the latest published version of the Data Standard specification: WHO/PQS/E006/DS01 | [Please state whether all recorded data objects are recorded in accordance with WHO/PQS/E006/DS01.]
4.2.6 | Data recording
| For all data objects sampled on a regular, time-basis (e.g. temperature), data recording shall occur at least every 15 minutes and include a timestamp for the recording. | [Please state or refer to a document showing the capture interval for all data objects or state the range of capture intervals for this object if user adjustment is possible.]
| For all data objects that are sampled on an event-basis (e.g. door opening), the EMD shall save a record of those events in the associated time-based data object. | [Please show an example time-based data record and describe how event-based data are associated with the data in the record.]
| The EMD shall be able to access data available on any connected (via M2M interface or integrated) logger such that data and alarms may be triggered and shared locally and remotely if utilizing remote communication modalities. | [Please refer to the user manual and show an example (screen shot?) of how data and alarms may be shared locally and remotely (if utilizing remote communication modalities).]
4.2.7 | Data transfer mode
| External EMDs shall act as a “host” device and periodically mount the logger USB device and access its data. The EMD shall not mount the logger for longer than 180 seconds to avoid prolonged periods where the logger cannot record data objects. | [Please show the results of a test of downloading data from a logger including the duration of the USB connection / mount period.]
4.2.8 | Generating data objects
For EMDs, every data object is listed as either required or optional for generation. For all required data objects for generation, the EMD shall have the ability to generate data related to those data objects. For all required data objects for recording, the EMD shall have the ability to record data related to those data objects.
4.2.9 | Required data objects for EMD generation
The EMD shall have the ability to generate the following data objects:
| Administration: • Absolute time (for remote EMDs only): Coordinated Universal Time (UTC) time derived from an independent verified source (e.g. cellular tower, GPS, Internet time server), standardized according to ISO 8601 Internet Date Time profile, using days, hours, minutes, and seconds (YYYYMMDDThhmmssZ). • EMD administrative data (for remote EMDs only): data related to the EMD shall include manufacturer, model, software version, manufacturer serial number, date of production, WHO PQS code. | [Please confirm the EMD will generate data for Absolute time; Coordinated Universal Time (UTC) according to ISO 8601 Internet Date Time profile, using days, hours, minutes, and seconds (YYYYMMDDThhmmssZ) and EMD administrative data including manufacturer, model, software version, manufacturer serial number, date of production and WHO PQS code.]
| Performance: • EMD Battery Remaining: estimated number of days of battery remaining to operate the EMD normally, measured in days (DDDD.D). | [Please show an example of the estimated battery remaining time on the EMD display. (Cross-reference VP sheet Section 5.3.6.)]
| Error Codes: • Local-only EMDs: The EMD shall generate and record error codes for any condition that may impair its normal operation (e.g. broken or disconnected M2M data or power connections, self-test failure). The EMD shall record all error codes that are available on the logger. Such error codes shall be accessible to local, on-site users via the EMD for at least 60 days. • Additional requirement for remote EMDs: Error codes related to the EMD shall be communicated to Remote Data Management Systems. | [Please show examples of all error codes on the EMD display.]
4.2.10 | Optional data objects for EMD generation
The EMD may have the ability to generate itself the following data objects:
| Administration: • Location / GNSS Coordinates: data related to the physical location of the EMD. This is optional because in some regions health facility location information can be a security concern. Any legal manufacturer or reseller integrating geolocation features shall make this known to WHO PQS at the time of prequalification and have the ability to remove this functionality if requested by an employer. • Data Integrity: File data integrity data such as cyclic redundancy check (CRC), hash, or cryptographic file signature that can be used for data validation. | [Please state whether the EMD can generate the following optional data objects: - Location of the EMD. - Protects integrity of data.]
| Environment: • Ambient Relative Humidity: relative humidity of the environment in which the appliance is deployed, measured in RH% at least once every 15 minutes. • Ambient temperature: temperature of the immediate surroundings within which the EMD is operating, measured in degrees Celsius (°C) at least once every 15 minutes. | [Please state whether the EMD can generate the following optional data objects: ambient temperature and ambient humidity.]
| Performance: • Holdover, Autonomy, or Independence time remaining: estimated number of days of holdover/autonomy/independence time remaining (DDD.D). | [Please state whether the EMD can generate the following optional data objects: Holdover, Autonomy and Independence-time remaining.]
4.2.11 | EMD data recording
The EMD shall have the ability to record all required data necessary for local visualization and remote data transmission. These data include EMD-generated data objects defined in Clauses 4.2.9 and 4.2.10, as well as downloaded logger data files/objects.
4.2.12 | Continuous operation without power supply
| The EMD shall continue to perform required functions in this specification, such as alarming, during periods when there is no mains or solar power supply. This minimum period of continuous operation shall be no shorter than: • Integrated EMDs: o Stationary solar-powered appliances including cold/freezer rooms: 1.5x rated autonomy time at +43 °C or 96 hours at +43 °C, whichever is greater. o Stationary mains-powered appliances including cold/freezer rooms: 2.0x rated holdover time at +43 °C or 96 hours at +43 °C, whichever is greater. o Transportable powered appliances: 1.5x rated independence time or 18 hours at +43 °C, whichever is greater. o Refrigerated vehicles: 1.5x rated non-idle run time of the refrigeration unit. • External EMDs: 240 hours, such that E-EMDs can be paired with “long” holdover mains appliances that have a minimum of 120 hours of holdover. | [Please provide evidence for the EMD duration when there is no power supply. (See E006/EM01-VP.2 Clause 5.3.3)]
| During such periods of no power supply, EMDs may utilize methods to conserve energy, but shall maintain the following minimum functions: • The EMD shall continue to trigger alarms within 15 minutes of the alarm event occurring. • The visual display may be dimmed or shut off but shall be reactivated by a user in an intuitive way (e.g. touching screen or pressing a button). • Frequency of remote communication may be limited so long as there is not a change in active alarm condition or error code. When an alarm or error code is triggered, a remote EMD shall immediately communicate this change in alarm or error status via remote communication networks. If an alarm or error condition no longer exists, this change in status shall be immediately communicated. By transmitting only changes to alarm or error status, it may not be necessary to repeatedly transmit an unchanged alarm status, thereby conserving resources. | [Please confirm that during such periods of no power supply, the EMD will continue to trigger alarms and, • that the visual display can still be reactivated by a user and, • there is not a change in active alarm condition or error code. and, • if an alarm or error condition no longer exists, this change in status shall be immediately communicated. (See E006/EM01-VP.2 Clause 5.3.3)]
4.2.13 | Energy storage lifetime
| Energy storage (e.g. electrical batteries or capacitors) is required to enable continued monitoring (specified in Clause 4.2.12) during times when supply power is lost. Various approaches to energy storage are acceptable but shall comply with the minimum requirements in this and the following sections: | [Please state for how long continuous monitoring is available when the supply power is lost.]
| Components for storing energy may be replaceable or non-replaceable and rechargeable or non-rechargeable. Minimum functional lifetimes are specified based on whether the EMD is integrated or external: • Integrated EMD: The functional lifetime of any non-replaceable energy storage components shall be no less than the appliance within which it is integrated, i.e. 10 years for refrigerators. Replaceable energy storage components must have a functional lifetime of no less than 5 years. • External EMD: The functional lifetime of any non-replaceable energy storage components shall be no less than 10 years. Replaceable energy storage components must have a functional lifetime of no less than 5 years. | [Please provide evidence for the battery (or other) lifetime when there is no power supply.]
| Replacement of energy storage components: If the batteries are intended to be replaceable, such a task shall be possible by a trained technician on-site; replacement shall not require recalibration of sensors. | [Please refer to the user manual Section which show how to replace the battery and confirm that sensors do not need to be re-calibrated.]
4.2.14 | Energy storage recharging
| EMD energy storage components may be rechargeable. Intelligent charging schemes should be utilized to extend the functional lifetime of the energy storage components. For rechargeable power components: If an EMD’s energy storage is fully depleted, it shall have the ability to recharge sufficiently in at most eight hours - assuming power supply - to then operate without power supply for at least 48 hours. | [Please provide evidence to show that if an EMD’s energy storage is fully depleted, it shall have the ability to recharge sufficiently in at most 8 hours (assuming power supply is available) and to then operate without power supply for at least 48 hours.]
4.2.15 | Calibration
| For I-EMDs, system components for vaccine storage chamber temperature measurement are to be covered by a Certificate of Traceability and Calibration. The traceability declaration is to confirm that the measurement standards and instruments used during calibration of the components are traceable to an ISO/IEC 17025 accredited testing laboratory, to NIST, or to another internationally recognized standards agency. The certificate shall be accompanied by a copy of the reference instrument calibration certificate. | [Please provide certificate and test data show calibration of the EMD in accordance with this clause.]
4.2.16 | Electromagnetic compatibility
| Operation of the system and of the individual system components shall be unaffected in the normal electromagnetic compatibility environment in which the system is intended to work, considering disturbances generated by adjacent appliances which are compliant with the requirements of IEC 61000-6-1 for immunity and IEC 61000-6-3 or IEC 61000-6-8 for emissions. Information to ensure uninterrupted use of the device shall be stated in the user instructions. | [Please provide evidence for Electromagnetic compatibility.]
4.2.17 | Remote data systems general requirements
| Remote data clauses are not applicable to local-only EMDs. Remote Data Systems are the server-based, remote software systems that manage monitored data. The legal manufacturer or reseller is required to comply with the following minimum specifications and abide by guidelines when negotiating agreements with employers. | [Please state whether EMD is a Level 2 or Level 3. For a Level 2 EMD, please state "Not Applicable for Clauses 4.2.18 through 4.2.22. For a Level 3 EMD, please see Clause 4.2.18 through to Clause 4.2.22.]
4.2.18 | Geographic location of remote data hosting
| The acceptable location (country) for data hosting and associated server storage shall be negotiated between the legal manufacturer or reseller and the employer at the time of procurement. If the employer prefers to host data within their country, the legal manufacturer or reseller should make efforts to support this approach. However, the ultimate responsibility for establishing secure, functional data hosting within a programme country is the responsibility of the employer unless otherwise negotiated with the legal manufacturer or reseller. | [Applicant to consult with "employer" at the time of procurement.]
4.2.19 | Scope and period of remote hosted data access
The legal manufacturer or reseller shall make hosted data accessible for a period of time to employers and those they designate with access rights. Requirements and general guidelines of the scope of data and time period for access include:
| • The legal manufacturer or reseller shall offer in their dossier submission a standard service package that includes three years of data hosting and is compliant with all Remote Data System requirements. Actual service periods shall be negotiated with the employer at the time of procurement. | [Please give an example of a standard service package including three years of data hosting (compliant with all Remote Data System requirements).]
| • The effective time period by which the legal manufacturer or reseller of the Remote Data System shall satisfy the functional requirements within this specification shall be negotiated with the employer at the time of procurement. | [Please negotiate with the "employer" at the time of procurement.]
| • All data includes raw and summary data (e.g. alarms) related to the EMD and the appliance and logger being monitored by that EMD. | [Please provide an example of a raw and summary data download and refer to the user manual to explain how to obtain a similar download.]
| • If so desired by the legal manufacturer or reseller, older raw data may be archived after a period agreed to by the employer, but this raw data shall still be accessible upon request during the data hosting service period. | [Please explain or refer to the user manual how older raw data may be archived while still being able to access this raw data.]
| • At the end of an employer’s hosting service contract period, the legal manufacturer or reseller shall provide advance notice to the employer and allow for a data transition period whereby the employer may export any or all data before it is no longer hosted or accessible via the legal manufacturer’s or reseller’s Remote Data System. | [Please negotiate with the "employer" at the time of procurement.]
4.2.20 | Data ownership
| Ownership of all data collected and/or transmitted by the EMD is vested with the employer, and the employer has control over the use, storage, data protection requirements, transmission, internal processing and terms of access and use of the data by third parties throughout the full lifespan of the data. No data (unprocessed, processed, generated by or entered into the Remote Data System), whether anonymized or not, will be shared with any third parties by the legal manufacturer or reseller without explicit and informed consent by the customer. | [Please demonstrate or refer to a Section in the user manual how the data is protected from unauthorized access.]
4.2.21 | Data delivery to external systems
| When requested by the employer, the legal manufacturer or reseller providing remote communication services shall transmit cold chain data to the employer per the requirements in Clause 5 of WHO/PQS/E006/DS01. | [Please give an example of transmitted cold chain data in accordance with the requirements in Clause 5 of WHO/PQS/E006/DS01.]
4.2.22 | Data security
| The legal manufacturer or reseller shall demonstrate evidence of conformity with ISO/IEC 27001:2013. | [The legal manufacturer or reseller shall demonstrate evidence of conformity with ISO/IEC 27001:2013.]
4.3 | Environmental requirements
All EMD components shall be adequately protected from the environment in which they will be deployed, including ambient temperature, ambient humidity, input voltage and dust and water ingress. Some requirements are determined by whether the device is integrated within the appliance or standalone from it.
4.3.1 | Ambient temperature
| All EMD components shall tolerate high and low temperatures during transportation and storage, specifically in the range of -30°C to +70°C with the system components inactivated. | [Please demonstrate or provide evidence that all EMD components can tolerate temperatures in the range of -30°C to +70°C during transportation and storage.]
4.3.2 | Ambient humidity
| During transport, storage, and use, the EMD may be subject to a wide range of uncontrolled ambient relative humidity (RH) from 0 to 95% RH. The EMD shall be designed to remain undamaged and fully functional if exposed to these conditions in hot zone temperatures (+43°C). | [Please demonstrate or provide evidence that the EMD can tolerate humidities in the range of 0 to 95%RH during transportation, storage and use.]
4.3.3 | Input voltage fluctuations
| The functionality of the system and of the individual system components shall not be affected by voltage fluctuations due to issues including (but not limited to) problematic power grid and intense electrical storm activity. • Integrated EMD: at a minimum, the EMD shall be protected by the appliance’s voltage protection hardware such as an integrated or external voltage stabilizer. • External EMD: the legal manufacturers or reseller of external EMD shall provide adequate voltage protection hardware as part of the procurement if utilizing M2M power supply is not feasible, and an external mains power supply is provided. | [Please explain, demonstrate or provide evidence on how the internal or external EMD is protected by an appropriate voltage stabilizer or similar protection hardware.]
4.3.4 | Dust and water ingress
| Integrated EMD requirements: • Any component directly exposed to the external environment (e.g. visual display on the external surface of the appliance) shall have no less than IEC 60529: IP64 protection. • There is no minimum IP rating requirement for EMD components that are internal to the appliance and are not directly exposed to the external environment. However, the legal manufacturer(s) of the appliance and integrated EMD shall take care in the design and integration of the EMD within the appliance such that the design eliminates or minimizes the effects of dusty, humid environments with potential water splashing. • Conformal coating of Printed Circuit Board Assembly (PCBA): any PCBA not protected by an IP-65 rated enclosure shall be protected from dust, water, and humidity through the use of a conformal coating. Manufacturers should consider coating guidelines in IPC document IPC-HDBK-830. External EMD requirements: • No less than IEC 60529: IP64 protection. | [For an integrated EMD, if it has component(s) exposed to the external environment (e.g. a visual display), please provide a test report showing compliance to IEC 60529. If all EMD components are purely internal to the appliance, explain how the design and integration of the EMD eliminates or minimizes the effects of dusty, humid environments with potential water splashes. For an external EMD, please provide a test report showing compliance to IEC 60529: IP64.]
4.3.5 | Corrosion
| The legal manufacturer shall utilize non-corrodible and robust material (e.g. plastics or treated metal) for all components. | [Please state the component materials and their protection against corrosion.]
4.4 | Physical characteristics
4.4.1 | Overall dimensions
| No requirements. | [Please state the dimensions for the EMD and any ancillary components e.g. power connectors.]
4.4.2 | Weight
| No requirements. | [Please state the weight the EMD and any ancillary components e.g. power connectors.]
4.5 | Interface requirements
4.5.1 | External EMD mount
| External EMDs shall provide a means of being fixed to the appliance. | [Please provide a photo or refer to a Section in the user manual.]
4.5.2 | Power supply
| The following modes of power supply shall be possible: • Integrated EMDs: I-EMDs may utilize power supplied by the appliance to allow for operation and charging of any EMD energy storage components. The power supply to the EMD shall not interfere with normal operation of the appliance. • External EMD: E-EMDs utilize power supplied by the appliance via the M2M power connection, but shall draw no greater than 5 W (1.0 A ± 5% at 5 V). | [For both internal and external EMDs, please state the type and specifications (including measured power consumption) for the type of power supply provided. Please demonstrate or provide evidence that the power supply to the EMD does not interfere with normal operation of the appliance. (See E006/EM01-VP.2 Clause 5.3)]
4.5.3 | M2M power connection
| External EMDs shall have a female jack connection that accommodates a barrel plug connector with the following specifications: • Type: Barrel • Sleeve diameter: 5.5 mm • Sleeve length: 9.5 mm • Pin diameter: 2.1 mm • Polarity: pin positive, sleeve negative The legal manufacturer or reseller of external EMDs shall provide a 1.5-meter-long extension cable to connect the shorter captive cable fixed to the appliance with the external EMD. This cable shall be compatible with the male plug output from the M2M, and the input jack required on the external EMD. | [If E-EMD, please show that the female jack connection meets these specifications.]
4.5.4 | Failsafe load design
| Design of both integrated and external EMDs shall mitigate failure modes that could cause it to draw greater than 5 W from the appliance or M2M power connection. | [Please demonstrate or provide evidence on how the design limits the power consumption to 5 W or less.]
4.5.5 | External mains supply power & cable
| External EMDs may also have their own dedicated power supply. If a dedicated power cable is provided to power the EMD (not from the appliance), the supply power cable shall be fitted with a plug that matches the mains electrical socket standard in the country of intended use (to be specified by the relevant procurement agent at the time of ordering). Minimum connection lead length shall be 1.5 meters. | [Please provide a photo or show an example of the power supply for an external EMD. Please state the connection lead length.]
4.5.6 | External solar supply power & cable
| External EMDs may also have their own dedicated solar power supply system. Solar power systems shall comply with requirements in the latest published version of the Solar Power system for low electrical requirements specification (WHO/PQS/E006/PVdc01). | [If the external EMDs has its own dedicated solar power supply system, please provide evidence for or show compliance to WHO/PQS/E006/PVdc01.]
4.5.7 | M2M data connection for external EMDs
| The M2M data connection that an E-EMD uses to collect data from an appliance shall be compatible with the latest published version of the logger and M2M specification (WHO/PQS/E006/DL01). This includes support for both FAT16 and FAT32 USB Mass Storage Device filesystem formats required in that specification. Also note that some legacy 30-day temperature loggers with USB interfaces use FAT12 format: supporting FAT12 is encouraged but not required. External EMDs shall have a USB Type-C female receptacle for the purpose of connecting a USB Type-C plug connection cable (male-male ends) with the appliance’s M2M data port. The legal manufacturer or reseller of external EMDs shall provide a compatible USB-C cable with male plugs on both sides. The cable must be at least 1.5 meters long. | [Please state the type of M2M connection and demonstrate that this is compatible with WHO/PQS/E006/DL01. Please provide photo(s) of all USB-C connections and state the length of the cable.]
4.5.8 | Data access port
| A data port may be incorporated as part of an External EMD for the purpose of locally accessing data that is recorded and stored within the EMD. If employed, this data port shall be a USB-C female receptacle interface such that the EMD can be mounted as a USB mass storage device. The connection shall provide access to all recorded data and if any custom software is needed to access data, that software must be freely and easily available. | [Please provide photo(s) and state the type of all data access ports.]
4.5.9 | Remote communication modalities
| This section is not applicable to local-only EMDs. The ability of EMDs to communicate with Remote Data Systems via an Internet connection provides additional levels of functionality and value to immunization programmes. However, remote communication capabilities are not required of EMDs and are specified as an additional feature with flexibility on communication modality (e.g. GSM / SIM-based, Wi-Fi®, etc.). If a legal manufacturer or reseller chooses to provide remote communication capabilities, the requirements in this clause and the next two shall be satisfied. All data objects, alarms, error codes, etc. recorded by the logger and EMD shall be communicated to Remote Data Systems. | [For Level 3 EMDs, please demonstrate or provide evidence that all data objects, alarms, error codes, etc. recorded are communicated to Remote Data Systems.]
| EMD connection to Remote Data Systems may be achieved through various modalities of transmitting information onto the Internet based on the following guidance and requirements: • Cellular-based communication: EMDs may communicate via mobile telecommunication networks that utilize both local and global SIM cards. The employer should specify their preference at the time of procurement. The legal manufacturer or reseller shall be able to offer functionality with local, global, and M2M/IoT SIM cards as requested by the employer. After commissioning, it shall be possible to change from local to global SIM cards and vice-versa. When switching from local to global SIM cards after commissioning, it is acceptable that a configuration of Access Point Name (APN) may be required. A 3-year communications contract length shall be a standard offering with the flexibility to negotiate other terms at the time of procurement. • Other communication capabilities: transmitting data to the Internet may also be achieved via Satellite, Wi-Fi®, and other networking solutions. The legal manufacturer or reseller should take care to ensure the remote transmission service provided is appropriate and will deliver on the functionality required under this specification for the communication environment of the employer’s country. | [Please state how the EMD may be connected to and give details for Remote Data Systems. Please state the options for communication contracts.]
4.5.10 | Remote data transmission frequency
| • Transmission frequency: Transmission frequency shall be frequent enough such that there is no greater than a 15-minute delay from when an alarm event occurs to when it is communicated and known remotely. • Less-frequent data transmission: if no alarm conditions exist, data transmission frequency may be reduced to conserve resources so long as alarms are still triggered remotely within 15 minutes of the alarm condition occurring within the appliance. Under this condition, new data not previously communicated remotely should be uploaded at least once every 24 hours, unless otherwise agreed at time of procurement. | [Please state transmission frequency or if variable, the range of frequencies and conditions under which they vary.]
4.5.11 | Connection and reconnection to remote data systems
| When powered on, a Cellular-based EMD that has been procured with a SIM card and cellular service enabled shall automatically connect to cellular networks and transmit data to the server without any additional configuration beyond physical setup. Connection to non-cellular, Internet-connected networks (e.g. Wi-Fi®) should be simple for users or technicians to complete with sufficient knowledge of the communication modality used and local network access privileges. In the event of any communication interruptions (e.g. loss of power supply), the EMD shall automatically reconnect to whichever Internet-enabled network is used once the interruption is resolved (e.g. power supply resumes). | [Please confirm that the EMD automatically connects to cellular networks and transmit data to the server without any additional configuration beyond physical setup (plug and play) or demonstrate that connection to non-cellular, internet-connected networks is straightforward. Also confirm that the EMD automatically reconnects to the same internet-enabled network once the interruption is resolved.]
4.6 | Human factors
4.6.1 | Visual display | (See E006/EM01-VP.2 Clause 5.3.6))
| Visual display requirements for both local-only and remote-enabled EMDs are specified below. • An on-site display shall be part of the EMD. An on-site display is not required to be permanently affixed to an appliance. • Monitored data may be displayed via secondary devices (e.g. smartphone applications), but this shall not replace the ability to visually display information via the EMD or appliance. • One local display may display information for more than one on-site appliance. • The display should be designed in such a way that essential performance and safety information is available, but not overwhelming to users who may have limited experience with technology. • The display shall allow users to access required information in as few “clicks” or interactions as possible while maintaining simplicity and readability. “Home screens” should be basic and only communicate essential information such as date, time, and vaccine compartment temperature. • Display contents: the visual display shall have means of displaying the following information: o Local date and time o Display languages: displays shall utilize English as a default, with translations readily available in all UN languages (Arabic, English, French, Mandarin, Russian, and Spanish). Display languages specific to certain countries are to be provided if requested by the buyer. | [Please provide photo(s) of the on-site display and demonstrate (references to Sections in the user manual may be appropriate) that the display shows: - essential performance and safety information, - the display allows users to access required information in a few “clicks”, - photo(s) of the home screen, - local date and time, - displays English as a default, - displays the most recently recorded temperature of coldest point in the vaccine compartment, - alarm status, - alarm history (at least the last 30 days of heat and freeze alarm data, including alarm duration and maximum/minimum temperature), - alarm disabled status per Clause 4.6.7, - error status with activation indicating a component error with what type of error, - historical storage chamber temperature summaries including the last 30 days of daily maximum, minimum, and average storage chamber temperature data, - whether there is supply power available to the appliance, - whether the appliance logging system is operating normally, or an error has occurred, - logger-generated error codes (for at least 60 days),
| o Vaccine compartment temperature: most recently recorded temperature of coldest point in the vaccine compartment. o Alarm status: activation of obvious visual cue that an alarm condition is occurring and specifically what type of alarm, e.g. “freezing temperature alarm”; the visual cue should serve to notify users and trigger action based on programmatic guidance for responding to specific alarms. All alarms listed in the latest published version of the Data Standard specification (WHO/PQS/E006/DS01) shall be visually communicated on the display when triggered. The required contents of the home screen listed here shall still be visible during an active visual alarm. o Alarm history: At least the last 30 days of heat and freeze alarm data, including alarm duration and maximum/minimum temperature. o Disabled alarms: Display alarm disabled status per Clause 4.6.7. o Error status: activation of obvious visual cue that an EMS system or component error is actively occurring and specifically what type of error. o Historical storage chamber temperature summaries: At least the last 30 days of daily maximum, minimum, and average storage chamber temperature data. | - logger battery charge OR battery replacement warning indicator, - EMD status, - EMD battery charge level indicator OR battery replacement warning indicator.]
o Appliance supply power: Indication whether there is supply power available to the appliance. This visual indicator may be driven by the existence of supply power to the EMD from the appliance, or the EMD may optionally implement more sophisticated approaches to represent appliance power availability more accurately. o Logger connection status: Whether the appliance logging system is operating normally, or an error has occurred. o Logger-generated error codes: The EMD shall display all error codes that are available on the logger. Such error codes shall be accessible to local, on-site users via the EMD for at least 60 days. o Logger battery charge (if applicable): Level indicator OR battery replacement warning indicator; activated when the logger battery charge is below 20% of its rated capacity for non-rechargeable batteries, or for rechargeable batteries when the capacity is less than 80% of rated capacity. o EMD status: Whether the EMD is operating normally, or an error has occurred. o EMD battery charge (if applicable): Level indicator OR battery replacement warning indicator; activated when the battery charge is below 20% of its rated capacity for non-rechargeable batteries, or for rechargeable batteries when the capacity is less than 80% of rated capacity.
4.6.2 | Additional visual display requirements for remote EMDs
| For EMD with remote communication capabilities, the following are additional display requirements: • Cellular signal strength indicating strength of connection with remote communication network. • Date and time of last successful data transmission. • KPI visualization. Metrics to assess overall appliance performance shall be accessible to users and indicate the period for which the KPIs were calculated. | [Please provide photo(s) of the on-site display and demonstrate (references to Sections in the user manual may be appropriate) that the display shows: - Cellular signal strength, - Date and time of last successful data transmission. - KPI visualization.]
4.6.3 | Audio-visual alarms
| A combined audio-visual alarm shall be employed to notify users or employers whenever a temperature or performance event occurs outside user-programmed settings. These requirements apply to both local-only and remote-enabled EMDs. Required alarms and their default threshold settings shall comply with the latest published version of the Data Standard specification (WHO/PQS/E006/DS01). It is acceptable for E-EMDs to have up to 15-minute latency in alarm triggering due to the expected typical M2M data download interval. Local adjustment of alarm configuration is optional. If implemented in local-only EMDs, users shall have the ability to adjust these default settings using the EMD interface without the need for an Internet connection or other devices. Functional requirements for alarms are: | (See E006/EM01-VP.2 Clause 5.3.6)
| • Visual alarm: The visual indication of an alarm shall be displayed prominently such that users cannot mistake the alarm for any other displayed information. The display of the alarm shall be generally described as “obvious” and “apparent” to any user that is within the immediate vicinity of the display. | [Please provide photo(s) of the on-site display and demonstrate (references to Sections in the user manual may be appropriate) how the display shows alarm visualization.]
| • Audio alarm: 70dB(A) at one-metre distance from the sounder. The pattern of the signal is to be an intermittent pulse. The timing and/or pattern of the pulse shall be set to ensure that it cannot be confused with other emergency alarms as specified in ISO 8201. EMDs with an adjustable sound profile will be acceptable provided the means for adjustment is not easily accessible. | [Please describe the type of alarm pattern and show measurements that it is 70dB(A) at 1 m from the sounder. (See E006/EM01-VP.2 Clause 5.3.4)]
| • Alarm activation: An audio-visual alarm shall be triggered locally automatically and immediately whenever a defined alarm condition is detected by the EMD. | (See E006/EM01-VP.2 Clause 5.3.4)]
| • Alarm deactivation: An audio-visual alarm is to be automatically deactivated whenever an alarm event is no longer occurring (e.g. internal compartment has returned to a safe temperature). | [Please show an example (screen shot?) of a deactivated alarm event-based data object. (See E006/EM01-VP.2 Clause 5.3.4)]
| • Alarm muting: The user shall have a means of muting the audible alarm on-site. The mute function shall be easily accessible to users. Muting shall not deactivate the alarm, but rather stop all audio alarming sounds for the duration of the current alarm. | [Please make reference to the appropriate Section in the user manual that shows how to mute the alarm. (See E006/EM01-VP.2 Clause 5.3.4)]
| • Disable/enable alarm monitoring: Local users shall have the ability to indefinitely disable audio-visual alarms in the event the appliance is inactive or not storing vaccines. When alarms are disabled, the disabled status must be prominently visible on the display so that users will not forget to re-enable when the appliance is placed back in use. | [Please refer to the user manual section which shows how to locally disable the audio-visual alarm. (See E006/EM01-VP.2 Clause 5.3.4)]
| • Alarm recording: Each of the alarm actions listed here shall be recorded by the EMD. | [Please give examples of alarm events being recorded by the EMD, e.g. a data local data download or a view of remote data systems containing the data.]
| • Alarm data summary: Data summarizing the alarm event shall comply with Data Standard (WHO/PQS/E006/DS01). | [Please show compliance with WHO/PQS/E006/DS01, Annex 1.]
4.6.4 | Remote audio-visual alarms
| For EMD with remote communication capabilities, the following are additional alarm requirements. • Alarm communication latency: within 15 minutes of an alarm state being recognized by an EMD, the EMD shall communicate the alarm to remote data systems. • The EMD may optionally transmit alarm muting or other acknowledgement actions to remote data systems as described in the Data Standard (WHO/PQS/E006/DS01). | [Please state the delay between the alarm event and the EMD sending the alarm to remote data systems.]
4.6.5 | Remote alarm notifications
| It shall be possible to send remote notifications such as SMS text and email, to users via communication networks, but the details of such a service shall be negotiated at the time of procurements with the employer. If desired by the employer, SMS text and email are the minimum requirements for remote notifications. Such communications shall be directly actionable. | [Please state the options for sending SMS text and email to users via communication networks.]
4.6.6 | Remote alarm muting
| Alarms shall have the ability to be muted remotely by each of the following: 1) sending a remote response (e.g. through SMS, email), and 2) utilizing a cancellation function in an online dashboard. The latency of this remote action and the alarm being muted on-site shall be on a best effort basis. | [Please refer to the user manual Section to show how the alarm may be muted remotely.]
4.6.7 | Remote disable/enable alarm monitoring
| The audio-visual alarm system shall have the ability to be indefinitely suspended by remote intervention in the event the appliance is inactive or not storing vaccines. If activated, this suspension of alarms shall be visually communicated via online dashboards and on-site. Restarting alarm monitoring shall be simple to execute via the online dashboard. | [Please refer to the user manual Section to show how the audio-visual alarm system can be indefinitely suspended by remote intervention in the event the appliance is inactive or not storing vaccines; - and how the alarm monitoring can be restarted.]
4.6.8 | Online dashboards
| Online dashboards are Internet-based software applications that allow users and administrators to access and analyse data remotely. Numerous advanced dashboard features and functionalities may be possible, but the legal manufacturer or reseller shall offer a solution with the following minimum requirements: • Adjust thresholds for local and remote alarms & KPIs calculated on EMDs. • Mute active alarm(s) per Clause 4.6.6. • Disable alarm monitoring per Clause 4.6.7. • Grant access privileges to staff of the employer. • Access and download all recorded data in at least CSV format. • View an activity log of all system changes and actions taken by any party through the dashboard. | [Please refer to the user manual section to show how - • Adjust thresholds for local and remote alarms & KPIs calculated on EMDs. • Mute active alarm(s) per Clause 4.6.6. • Disable alarm monitoring per Clause 4.6.7. • Grant access privileges to staff of the employer. • Access and download all recorded data in at least CSV format. • View an activity log of all system changes and actions taken by any party through the dashboard.]
4.6.9 | Additional guidelines for advanced dashboards
| Feature-rich online dashboards are not required, but if offered, should allow users to easily: • See any immediate/urgent issues/alarms. • Generate summary views and graphs, allowing end-users to quickly see basic analysis on device performance, alarm status and some baseline data, such as high and low temperatures recorded during the day. • View all deployed appliances/EMDs as a list and spatially (on a map). • Select a specific EMD view information regarding that EMD and appliance at the individual sensor level. • Create and download reports. • Export data in CSV format and any images in JPEG and PDF formats, with export of data in JSON format optional. • Establish an alarm escalation tree: The system should allow a minimum of three layers of alarm notification escalation. The system administrator should be able to remotely specify (through the dashboard interface) recipients at each level of this escalation tree, their preferred form of notification (e.g. SMS, voice call, email), and the allowed time period for response before the next escalation. Manufacturers are also encouraged to use common user-interface design language for dashboards, and to create documentation to simplify learning and retention of dashboard features. A possible technique is presented in EM01.2 Annex 1. | [If offered, please refer to the user manual Section to show how to - • See any immediate/urgent issues/alarms. • Generate summary views and graphs, showing basic analysis on device performance, alarm status and some baseline data, such as high and low temperatures. • View all deployed appliances/EMDs on a map. • Select a specific EMD view information including at the individual sensor level. • Create and download reports. • Export data in CSV format and any images in JPEG and PDF formats, with export of data in JSON format optional. • Establish an alarm escalation tree.]
4.6.10 | Low-bandwidth platforms
| Low-bandwidth interfaces are optional but strongly recommended to allow users access if/when utilizing slow or weak Internet connections. | [Please state the size or range of low-bandwidth interface.]
4.6.11 | Mobile platform
| A mobile platform/app for users to access Remote Data System information is optional. | [Please state whether and how a mobile platform/app for users to access Remote Data System information could be available.]
4.6.12 | Additional threshold notifications
| Beyond the default alarm and KPI condition thresholds outlined in the Data Standard (WHO/PQS/E006/DS01), it may be possible to create additional notifications that are triggered by data object thresholds in the event “warning” notifications are desired by the employer prior to alarm thresholds being reached. The same features of activation/deactivation, muting, enable/disable and recording of notifications shall apply to these additional notifications. There is currently no WHO definition for these thresholds and thus this functionality is optional to provide an extra level of customization based on industry and programme needs over time. | [Pleaser refer to the user manual Section to show how it is possible to create additional notifications that are triggered by data object thresholds.]
4.7 | Materials
4.7.1 | Ozone depleting chemicals
| During manufacture and assembly of the printed circuit boards and final assembly of the product, do not use any substance included in Annex A, B or C of the Montreal Protocol. | [Please state the printed circuit board materials and whether they comply with the Montreal Protocol.]
4.7.2 | Other restricted materials
| The product and its constituent components, including batteries, shall not contain lead, mercury, cadmium, hexavalent chromium, polybrominated biphenyls (PBB) or polybrominated biphenyl ethers (PBDE). | [Please provide a RoHS report or equivalent. See E006/EM01-VP.2 Clause 5.3.1)]
4.8 | Warranty
| The EMD shall be covered by a replacement warranty in the event of any component failure arising from defective design, materials, or workmanship. Specific requirements depending on whether the EMD is integrated or external: • Integrated EMD requirements: the warranty of the system and all components shall meet or exceed the warranty offered with the appliance(s). Warranty provisions shall not be separate for the monitoring system and shall be included as part of the warranty provided with the appliance(s). • External EMD requirements: all system components are to be covered by a three-year warranty in the event of any component failure. The employer will take full ownership of the EMD, warranty and aftersales conditions at the time of delivery to the port of entry. The operation and maintenance of the system will be the responsibility of the employer for the duration of the EMD's lifetime unless otherwise agreed upon. | [Please state or refer to a document showing the warranty with terms and definitions.]
4.8.1 | Minimum warranty for energy storage components
| The minimum warranty for all energy storage components shall be no less than three years. | [Please state the minimum warranty duration for all energy storage components.]
4.9 | Maintainability
| On-site maintenance of the system and the system components will be required over the lifetime of the EMD. The legal manufacturer or reseller shall declare how their EMD design mitigates long-term maintenance challenges associated with hot, humid, dusty climates with poor power quality and availability. • Integrated EMD requirements: the integrated design of the EMD shall consider ease of maintenance. Components that may require replacement over the lifetime of the EMD or appliance shall be easily accessible and serviceable. • External EMD requirement: the EMD shall be serviceable for common maintenance and repair needs. Key components shall be replaceable. | [Please state how the EMD design mitigates against long-term maintenance including challenges associated with hot, humid, dusty climates with poor power quality and availability. Also - • Integrated EMD requirements: refer to the user manual Section to show how to replace components that may require replacement over the lifetime of the EMD or appliance. • External EMD requirement: the EMD shall be serviceable for common maintenance and repair needs. Key components shall be replaceable.]
4.9.1 | Equipment servicing provision
| Integrated EMD requirements: • The EMD shall be designed to achieve a low-maintenance life of no less than the operational lifetime of the appliance within which it is integrated (e.g. 10 years for ILR and SDD refrigerators). • Batteries and sensors that do not require replacement throughout the lifetime of the appliance are preferable; however, they shall be serviceable or replaceable if the need arises. External EMD requirements: • The EMD shall be designed to achieve a low-maintenance life apart from routine cleaning, battery or sensor replacement and re-calibration. The minimum operational lifetime of the EMD shall be at least 10 years. | [For integrated EMD requirements, explain how the EMD is designed to achieve a low-maintenance life of no less than the operational lifetime of the appliance and - • how batteries and sensors are serviceable or replaceable. For External EMD requirements: • The EMD shall be designed to achieve a low-maintenance life apart from routine cleaning, battery or sensor replacement and re-calibration. The minimum operational lifetime of the EMD shall be at least 10 years.]
4.9.2 | Spare parts
| The legal manufacturer or reseller shall provide a complete list of the spare parts, including part numbers for all key components of the EMD. | [Please refer to the user manual Section or a document to show a list of spare parts.]
4.10 | Disposal and recycling
| The manufacturer shall provide information to the buyer on the hazardous materials contained within the system and suggestions for resource recovery/recycling and/or environmentally safe disposal as part of the recovery and disposal information provided for the associated cold chain equipment. For the European Union WEEE compliance in accordance with European Union Directive is mandatory. | [Please refer to a user manual Section or to a document explaining what hazardous materials are contained within the system and show suggestions for resource recovery/recycling and/or environmentally safe disposal. (For the European Union WEEE compliance in accordance with European Union Directive is mandatory.)]
4.11 | Instructions
Each integrated or standalone EMD shall include a user manual and technician manual in Arabic, English, French, Mandarin Chinese, Russian and Spanish. An English version of all instructions and manuals are required to be supplied at time of laboratory testing and prequalification. Instructions to include easy to understand visuals whenever possible to avoid reliance on text.
The user manual shall include the following information:
| • Health and safety guidance | [Pleaser refer to the user manual Section that shows health and safety guidance.]
| • Basic operations description | [Pleaser refer to the user manual Section that shows a description of basic operations.]
| • Advanced instructions for system administrators to configure EMD and dashboard (if applicable) | [Pleaser refer to the user manual Section that shows advanced instructions for system administrators to configure EMD and dashboard (if applicable).]
| • Preventive maintenance tasks (e.g. daily, weekly, and monthly). | [Please refer to the user manual Section that shows preventive maintenance tasks (e.g. daily, weekly, and monthly).]
| • Warranty and contact information for legal manufacturer or reseller. | [Please refer to the user manual Section that shows warranty and contact information for legal manufacturer or reseller.]
The technician manual shall include the following information:
| • Health and safety guidance | [Please refer to the user manual Section that shows health and safety guidance.]
| • Detailed operations description | [Please refer to the user manual Section that shows detailed operations description.]
| • Correct handling to avoid EMD or appliance damage and for the safety of handling persons | [Please refer to the user manual Section that shows correct handling to avoid EMD or appliance damage and for the safety of handling persons.]
| • Installation procedures | [Please refer to the user manual Section that shows installation procedures.]
| • Compatible solar power system voltage (external EMD only, and only if applicable) | [Please refer to the user manual Section that shows compatible solar power system voltage (external EMD only).]
| • Preventive maintenance tasks (e.g. daily, weekly, and monthly) | [Please refer to the user manual Section that shows preventive maintenance tasks (e.g. daily, weekly, and monthly).]
| • Corrective maintenance procedures | [Please refer to the user manual Section that shows corrective maintenance procedures.]
| • Diagnostic and repair procedures | [Please refer to the user manual Section that shows diagnostic and repair procedures.]
| • Itemised list of spare parts including part numbers | [Please refer to the user manual Section that shows itemised list of spare parts including part numbers.]
| • End-of-life resource recovery and recycling procedures | [Please refer to the user manual Section that shows end-of-life resource recovery and recycling procedures.]
| • User training guidance. | [Please refer to the user manual Section that shows user training guidance.]
| Printed user-operations and maintenance instructions specifically directed at the health centre staff shall be pictorial. All key information is to be summarized on a single pictogram sheet that may be fixed onto the appliance. The sheet is to be sufficiently durable to last the life of the appliance. | [Please provide a copy of the summarized pictogram sheet with this information.]
| All instruction manuals shall be supplied in printed format and easily available and downloadable in digital format on the Internet. Digital manual materials may also be provided with the EMD on a separate USB storage drive. | [Please provide a copy (in English) of the user and any other manuals. State where instruction manuals can be easily available / downloadable and whether they provided on the a separate USB storage drive.]
4.12 | Training
Printed user-operations and maintenance instructions specifically directed at the health
5. | centre staff shall be pictorial. All key information is to be summarized on a single
pictogram sheet that may be fixed onto the appliance. The sheet is to be sufficiently
| If requested, the legal manufacturer or reseller shall at a minimum provide remote training in the countries where their product is deployed via suitable teleconferencing tools. Remote training shall cover the installation, on-site use, maintenance of the hardware, configuration and use of software and dashboard interface and system troubleshooting. | [Please state how remote training could be available.]
| Training programmes shall cover the minimum requirements described in the following sections. Specific requirements related to integrated EMD and external EMD include: • I-EMD requirements o Training shall be included as part of the provided training for the appliance(s). • E-EMD requirements o Training shall be made available either through the installer of the appliance(s), or through other agents of the legal manufacturer or reseller of the EMD. Agreement on this arrangement shall be made among the legal manufacturer or reseller of the EMD and appliance(s), the procurement agent, and the employer. o The training programmes shall include physical installation and maintenance of the standalone hardware. | [Please state the topics covered by the remote training.]
4.12.1 | Installation and user training
| Practical, hands-on installation training shall be available for the installation and use of EMDs. Some provisions may not be applicable if the EMD is integrated versus external. | [Please state how the customer can apply for/access practical, hands-on training shall be available for the installation and use of EMDs.]
Training outcomes for participants:
| • Correct installation and setup of the complete EMD • Basic system troubleshooting • Commissioning and handover of installed equipment in the form of job completion checklists and/or reports | [Please confirm your training will cover all these items:]
Minimum training requirements:
| • Unpacking of the equipment and verification that the correct equipment components and accessories are available for successful installation of the EMD. • Use and interpretation of electrical circuit diagrams and installation manuals. • Hands-on assembly and installation of different components, such as M2M cables, sensors, sensor cables, etc. • Hands-on assembly and installation of components such as M2M data and power cables, including proper EMD positioning and cable management. • Use of the EMD including setting alarm conditions, muting alarms, configuring monitored asset parameters, and accessing monitored information through the Internet-based web portal. • Interpretation of all displays and recorded data • EMD system functionality test and troubleshooting • Standard commissioning procedures, including necessary metadata collection/reporting to associate EMD with proper facility name and administrative areas. | [Please confirm your training will cover all these items:]
4.12.2 | Technical training
| Training in correct installation of the EMD, system administration, setup, operation, troubleshooting, maintenance, interpreting of data, and training of further technicians, and system administrators shall be available to employers. | [Please confirm that this training shall be available to employers.]
Training outcomes for participants:
| • Correct installation, setup, and operation of the EMD • Correct interpretation of displays and recorded data • Advanced system maintenance and troubleshooting • Commissioning and handover of installed equipment in the form of job completion checklists and/or reports • Training of further technicians on the installation, setup, and operation of the EMD. | [Please confirm your training will cover all these items:]
Minimum training requirements:
| • Demonstrate the EMD, including main components, controls, connections, and functionalities • Use of the EMD, including setting alarm conditions, configuring appliance parameters, configuring users, SMS recipients and accessing information through the Internet • Interpretation of all displays and recorded data • Daily, weekly, and monthly tasks and best practices for remote temperature monitoring, data accessing, analysis, downloading and corrective action • Standard troubleshooting procedures • Standard corrective and preventative maintenance tasks • Frequently asked questions • Processes and best practices for training of technicians in the installation, setup, and operation of the EMD system | [Please confirm your training will cover all these items:]
4.2.13 | System administrator training
| Training in system administration, setup, operation, troubleshooting, maintenance, interpreting of data, and training of further system administrators shall be available to employers. | [Please confirm that this training shall be available to employers.]
Training outcomes for participants:
| • Correct set up and operation of the complete EMD • Correct interpretation of displays and recorded data • Standard system maintenance and troubleshooting • Training of further system administrators in the setup and operation of the EMD | [Please confirm your training will cover all these items:]
Minimum training requirements:
| • Demonstrate the EMD, including main components, controls, connections, and functionalities • Use of the EMD, including setting alarm conditions, configuring appliance parameters, configuring users, SMS recipients and accessing information through the Internet • Interpretation of all displays and recorded data • Daily, weekly, and monthly tasks and best practices for remote temperature monitoring, data accessing, analysis, downloading and corrective action • Standard troubleshooting procedures • Standard corrective and preventative maintenance tasks • Frequently asked questions • Processes and best practices for training of technicians in the installation, setup, and operation of the EMD system • Processes and best practices for training of further system administrators in the set up and operation of the EMD system. | [Please confirm your training will cover all these items:]
4.12.4 | Training materials
To support training outlined above, the following materials shall be available:
| • EMD installation & commissioning: o Training materials on EMD installation and setup, including mounting of the base unit, positioning, and mounting of sensors, connection to external power supply (where applicable), routing and protection of sensor and power cabling. o Training materials on standard commissioning procedures, including the use of commissioning and/or functionality checklists. | [Please provide training materials covering these topics.]
| • EMD use and troubleshooting: o Training materials on device and dashboard (if applicable) operation, including setting alarm conditions, configuring monitoring parameters, configuring users, SMS recipients and accessing information through the Internet. o Training materials on interpretation of all displays and best practices for remote temperature monitoring, data accessing, analysis, downloading and corrective action. o Training materials on basic and advanced system troubleshooting and maintenance procedures. o Standard Operating Procedure posters for vaccine storage facilities, covering topics such as interpretation of all displays and best practices for remote temperature monitoring, data accessing, analysis, downloading and corrective action. | [Please provide training materials covering these topics.]
| • Materials for training trainers: o Training materials on how to conduct training of further technicians. o Training materials on how to conduct training of further system administrators. | [Please provide training materials covering these topics.]
| All training materials shall be supplied in printed format and easily available and downloadable in digital format on the Internet. Digital manual materials may also be provided with the EMD on a separate USB storage drive. | [Please provide a link to the training materials.]
4.13 | Verification
| In accordance with PQS verification protocol WHO/PQS/E006/EM01-VP.2. | [Please refer to the WHO/PQS/E006/EM01-VP.2 test report and see "VP Sheet" in this submission template.]
5.0 | Packaging
| Integrated and external EMDs shall comply with the following packing requirements. If the EMDs is integrated within an appliance, the packaging requirements are those found in the PQS specifications for those appliances. Materials used for packaging the finished appliance are to be free of ozone-depleting compounds as defined in the Montreal Protocol. The packaging is to be a sturdy export-quality and of a commercial standard that will provide adequate protection of the goods for carriage by air, sea and/or road to final destinations worldwide, including remote locations under adverse climatic and storage conditions and high humidity. The packaging is to be not less than 17kN edge crush resistance with minimum 60% remaining with 90% humidity at a temperature of +70°C (tropical conditions). To avoid destructive unpacking prior to installation, legal manufacturer or reseller is encouraged to add a resealable observation opening in their packaging to aid inspectors in finding labelling and/or placing additional markings prior to installation. Instructions on the packaging alerting inspectors to use of the opening and what information will be revealed are also advised. The general specification of shipping containers will be subject to agreement with the individual procurement agent. | [Please confirm all of the following: - Packaging materials are free of ozone-depleting compounds - The type of Packaging material that will provide adequate protection of the goods for carriage including adverse climatic and storage conditions and high humidity in accordance with the PQS clause requirement. - that it has a resealable observation with instructions.

## VP
E006 Submission to PQS for approval: Equipment Monitoring Devices for Equipment Monitoring Systems |  |  | Brand & Model:
| Tested in accordance with PQS verification protocol: | E006/EM01-VP.2
| To verify compliance with PQS specification: | E006/EM01.2
Instructions in [square brackets] are for your guidance. Please delete these instructions and replace with evidence or a precise reference (including page and section numbers) showing compliance with that particular PQS specification requirement. In general a "pass" or "complies" statement alone is not acceptable.
One document may provide evidence for several clause requirements. Please carefully reference the sub-sections of any such document so that it is clear that the clause requirement has been met. This will help speed the pre-qualification process.
Greyed-out cells do not require a response.
NOTE: If it is not clear or doubtful whether a requirement has been met, then this is a non-compliance and pre-qualification might be delayed.
Please ensure the Spec sheet is completed.
Section 7. | VP Clause | Applicant's submission | Assessor's remarks [Name & Date] | Applicants response [Name & Date] | Assessor's remarks [Name & Date] | Applicants response [Name & Date] | Assessor's remarks [Name & Date]
5.1 | Evidence of conformity assessment
| E-EMDs and appliances containing I-EMDs shall carry the CE mark, UL mark and/or equivalent internationally accepted evidence of conformity assessment. Appliances that are not ordinarily marked (e.g. cold rooms, refrigerated vehicles) are exempt, but relevant components should have evidence of conformity assessment. Certifications from accredited tests labs are preferable. If conformity is self-declared, then a technical construction file shall be provided as part of the dossier to validate conformity. | [Please provide evidence of any conformity mark if the appliance has one.]
5.2 | Samples and equipment to be provided by manufacturer
| The legal manufacturer or reseller must supply the testing laboratory with a single sample and a full duplicate set of the Product Dossier already supplied to WHO in accordance with the requirements of specification Clause 7. For cold rooms and refrigerated vehicles, the legal manufacturer may provide an I-EMD with representative sensors and instructions for simulating appliance power supply to the sample. | [Please provide the testing laboratory a draft dossier for checking and amendment. This will help to ensure that the dossier is complete.]
5.3.1 | Test 1: Type examination
| The test report must clearly show that the requirements of each section and sub-section in the type examination are individually and clearly met. | [Please refer to the test report and see the specification sheet in this template where precise references to each clause is required.]
5.3.2 | Test 2: Power supply and energy storage recharging | The information requested below can be shown in one downloaded data file or in several downloaded data files as applicable so that the evidence is convincing to the PQS reviewer. The submission to PQS can enhanced with annotated notes.
| E-EMD did not consume more than 5 W + 5% = 5.25 W (1 A current + 5% = 1.05 A) during the powered portion of the test. | [Please refer to the precise section in the test report showing the current drawn throughout the 8 hours (this can be graphical). A "pass" statement alone is not acceptable.]
| EMD successfully recorded 48 hours of continuous data following eight hours of charging. | [Please refer to the precise section in the test report showing the recorded data (e.g. in the form of a spreadsheet) for 48 hours. A "pass" statement alone is not acceptable.]
5.3.3 | Test 3: Endurance of operation without power
| Verification that EMD has continued to save required data objects through the period of power disconnection, and that those data objects can be visualized using the local display. | [Please refer to a precise section in the test report showing the recorded data (e.g. in the form of a spreadsheet) for the appropriate time in accordance with the type and application of the EMD. Also provide photo of an example of data object in the display. A "pass" statement alone is not acceptable. (Cross-reference Spec sheet Section 4.2.12.)]
5.3.4 | Test 4: Alarm and alarm muting functions (refrigerator)
| Freeze alarm results in an audible/visual alarm that meets the 70 dB(A) sound level requirement. | [Please refer to a precise section in the test report showing the recorded data for the freeze alarm event and an actual measurement of the alarm sound level. A "pass" statement alone is not acceptable.]
| Locally muting the alarm successfully results in muting the alarm for the duration of the current alarm condition. | [The tester should check and confirm the mute function is successful.]
| Disabling alarm monitoring successfully. | [The tester should check and confirm that freeze alarm monitoring can be disabled. It is useful if the steps necessary to disable alarm monitoring could be shown.]
5.3.5 | Test 5: Alarm and alarm muting functions (freezer)
| Heat alarm results in an audible/visual alarm that meets the 70 dB(A) sound level requirement. | [Please refer to a precise section in the test report showing the recorded data for the heat alarm event and an actual measurement of the alarm sound level. A "pass" statement alone is not acceptable.]
| Locally muting the alarm successfully results in muting the alarm for the duration of the current alarm condition. | [The lab tester should check and confirm the mute function is successful. The report should show how to mute the alarm.]
| Disabling alarm monitoring successfully. | [The lab tester should check and confirm that heat alarm monitoring can be disabled. The report should show the necessary steps to disable alarm monitoring.]
5.3.6 | Test 6: Display and user interface functions
| The required parameters and status information can be visualized on the local display. The required parameters are below: | (Cross-reference Spec sheet Section 4.2.9 & Section 4.6.1.)
| Vaccine compartment temperature | [The report should show an example photo of the EMD displaying the vaccine compartment temperature. The laboratory must check that the displayed temperature is the coldest temperature in the vaccine storage compartment.]
| Appliance supply power is OK | [The report should show an example photo of the EMD display indicating that appliance power is ok.]
| EMD battery charge state | [The report should show an example photo of the EMD display indicating battery charge state.]
| (Remote EMDs only) Cellular signal strength | [The report should show an example photo of the EMD display indicating signal strength.]
| (Remote EMDs only) Date and time of last successful data transmission | [The report should show an example photo of the EMD display indicating the date and time of last successful data transmission.]
| The appliance supply power status on the EMD updates to reflect the lack of supply power. | [The report should show an example photo of the EMD display indicating the updated lack of appliance supply power.]
| Disconnect the E-EMD from the M2M data interface and verify that within 30 minutes, the display indicates a logger connection fault. | [The report should show an example photo of the EMD display indicating a logger connection fault.]
| Verify that historic storage chamber temperature maximum, minimum, and average values for each of the past 30 days are available using the local display | [The report should show example photos of the EMD display indicating historic storage vaccine compartment temperatures of maximum, minimum, and average values for each of the past 30 days.]
| Verify that details regarding alarms (e.g. type of alarm and triggered duration) for each of the past 30 days are available using the local display, including the alarm triggered in Test 4 or Test 5 | [The report should show example photos of the EMD display indicating details regarding alarms (e.g. type of alarm and triggered duration) for each of the past 30 days are available using the local display, including the alarm triggered in Test 4 or Test 5.]
5.3.7 | Test 7: Data access port
| Logger mounts as a USB Drive. | [The report should confirm that the E-EMD has an appropriate access USB.]
| User able to access and download all files on EMD and, if drivers or software are needed, they are freely and easily available. | [The report should confirm by examples of data spreadsheets that the user is able to access and download all files on and confirm that any drivers or software (if needed) are freely and easily available.]
5.4 | Test criteria for qualification
A final report shall be issued after all testing is complete. The report of the tests shall contain the following data and analyses:
| Summary: Conclusions and recommendations. | [Please refer to the precise section in the report.]
| Test 1: Comments on samples received, tabulated data, compliance with required specifications (plus supporting documentation) and photographs of samples. | [Please refer to the precise section in the report.]
| Test 2: Results of Power Supply and Energy Storage Recharging test. | [Please refer to the precise section in the report.]
| Test 3: Results of Endurance of operation without power test. | [Please refer to the precise section in the report.]
| Test 4 or Test 5, as appropriate: Results of Alarm and alarm muting functions test. | [Please refer to the precise section in the report.]
| Test 6: Results of Display and user interface functions test. | [Please refer to the precise section in the report.]
| Test 7: Results of Data access port test. | [Please refer to the precise section in the report.]
| Annexes: A preapproved test protocol verifying that the procedures set out in this document have been followed. Description of the test apparatus. Test chamber temperature and humidity records. Copies of reference thermometer and humidity measurement calibration certificate(s). Additional supporting documentation requested and received from the legal manufacturer or reseller during the course of the type-testing. | [Please refer to the precise section in the report.]