## Sheet1
E006 Irreversible freeze indicator submission template (V2.0 13 MAR 2025)
| [Company name] | [Model name] | [Date]
| Dossier submitted by: | [Please enter your name, position within the company, and email address]
| Tested in accordance with PQS verification protocol | E006/IN03-VP.1
| To verify compliance with PQS specification below | E006/IN03.1
ASSESSOR’S SUMMARY:
Instructions in [square brackets] are for your guidance. Please delete these instructions and replace with evidence or a precise reference (including page and section numbers) showing compliance with that particular PQS specification requirement. In general a "pass" or "complies" statement alone is not acceptable.
One document may provide evidence for several clause requirements. Please carefully reference the sub-sections of any such document so that it is clear that the clause requirement has been met. This will help speed the pre-qualification process.
Greyed-out cells do not require a response.
NOTE: If it is not clear or doubtful whether a requirement has been met, then this is a non-compliance and pre-qualification might be delayed.
|  |  | REVIEWER’S SUMMARY: |  | REVIEWER’S SUMMARY: |  | REVIEWER’S SUMMARY:
Section 7. | Product dossier | Applicant's submission | Reviewer's remarks [Name & Date] | Applicants response [Name & Date] | Reviewer's remarks [Name & Date] | Applicants response [Name & Date] | Reviewer's remarks [Name & Date]
The legal manufacturer or reseller is to provide WHO with a pre-qualification dossier containing the following:
| Dossier examination fee in US dollars. | [Please contact PQS to arrange payment.]
| General information about the legal manufacturer, including name and address. | [Please enter this information, include for all parties, e.g. manufacturer, reseller, plus other.]
| Unique identification reference for the product type. | [Please enter this information.]
| Full specifications of the product being offered, covering all the requirements set out in this document, including details of product marking and traceability and documentary evidence of claimed battery life. | [The dossier must have clear traceable references to all information required below.]
| Certified photocopy of Certificate of Traceability and Calibration traceable to an ISO/IEC 17025 accredited testing laboratory, to NIST, or to another internationally recognized standards agency. | [Please enter the reference for the supporting document.]
| Certified photocopies of all type-approvals obtained for the appliance and/or its components, including CE marking and the like. | [Please enter the reference for supporting documents. A simple statement of compliance is not acceptable.]
| Certified photocopies of the legal manufacturer’s ISO 9001 quality system certification. | [Please enter the reference for the supporting document.]
| Where relevant, certified photocopies of the legal manufacturer’s ISO 14001 certification, EMAS registration or registration with an equivalent environmental audit scheme. Conformity with an environmental audit scheme is not mandatory; however, preference will be given to manufacturers who are able to demonstrate compliance with good environmental practice. | [Please enter the reference for the supporting document.]
| Where available, laboratory test report(s) proving conformity with the PQS equipment performance specifications. | [Please enter the reference for supporting document(s) .]
| One sample of the product and of the instruction leaflet in English language. | [Please send 1 sample to each of the addresses to be provided.]
| Indicative cost of the product per 100 units, per 1,000 units and per 10,000 units EXW (Incoterms 2010). | [Please state or refer to a document.]
Section 4. | Requirement
4.1 | General
The product will be packed with freeze-sensitive vaccines during transport, or during storage in fixed locations, and will be used to warn of exposure to temperatures below -0.5°C.
4.2 | Performance
4.2.1 | Operating temperature range
| Upper limit: +43°C. Lower limit (without triggering alarm): -0.5°C. Lower limit (without device failure): -20°C. | [Please enter the reference for supporting evidence in the dossier.]
4.2.2 | Accuracy
| ±0.5°C or better at 0°C | [Please enter the reference for the calibration report in the dossier. See also Clause 4.2.6]
4.2.3 | Power source
| None, or non-replaceable battery. | [Please state power source.]
4.2.4 | Sensor
| Electronic sensor or irreversible phase change indicator. | [Please state whether whether electronic or phase change.]
4.2.5 | Mode of operation
| The product is to be triggered by exposure to a temperature of -0.5°C, ±0.5°C, for 60 minutes ± 5 minutes maximum. It must not be possible for the end user to re-set the device after a freezing event. | [Please enter the reference for test evidence in the dossier.]
4.2.6 | Calibration
| Electronic products are to be covered by a Certificate of Traceability and Calibration. The traceability declaration is to confirm that the measurement standards and instruments used during calibration of the product are traceable to an ISO/IEC 17025 accredited testing laboratory, to NIST, or to another internationally recognized standards agency. | [Please enter the reference for the calibration report in the dossier. See also Clause 4.2.2.]
4.2.7 | Casing
| • Electronic devices: The sensor and any other working parts are to be housed in a non-corrodible water-resistant casing. • Passive devices: The device must be unaffected by overall wetting. | [Please describe casing design.]
4.2.8 | IP rating
| • Electronic devices: Protection of the product not less than IEC 60529: IP64. • Passive devices: Not applicable. | [If applicable, please enter the reference for test evidence in the dossier.]
4.2.9 | Battery life
| Minimum acceptable battery life for electronic products, with the product switched on, measured at any point in the operating temperature range, is to be 3 years. | [Please state the provided battery life, and where in the dossier to find evidence that the product meets this specification]
4.2.10 | Shelf life
| Minimum 3 years from date of manufacture, inclusive of operational life. | [Please state product shelf life and reference where that information is provided in the dossier.]
4.2.11 | Electromagnetic compatibility
| Operation of the device must be unaffected in the normal electromagnetic compatibility environment in which it is intended to work, taking into account disturbance generated by adjacent apparatus which is compliant with relevant ISO, EN, or other internationally recognized standards. Information required to ensure uninterrupted use of the device must be contained in the user instructions. | [Please state where in the dossier to find evidence that the product meets this specification.]
4.3 | Environmental requirements
4.3.1 | Ambient temperature range during transport and storage
| 5°C to +55°C with device inactivated. | [Please state where in the dossier to find evidence that the product meets this specification.]
4.3.2 | Ambient humidity range during transport, storage and use
| 0 to 95% RH. | [Please state where in the dossier to find evidence that the product meets this specification.]
4.3.3 | Electrical storm activity
| Operation of an electronic device must be unaffected by intense electrical storm activity. | [Please state where in the dossier to find evidence that the product meets this specification.]
4.3.4 | Impact resistance
| Product to withstand 5 drops from 1 metre onto a concrete floor, when cooled to a temperature of +3°C, without physical damage or loss of calibration. | [Please state where in the dossier to find evidence that the product meets this specification.]
4.3.5 | Vibration
| Product to withstand 30 minutes on a programmable vibrating table without physical damage or loss of calibration. | [Please state where in the dossier to find evidence that the product meets this specification.]
4.4 | Physical characteristics
4.4.1 | Overall dimensions
| Not exceeding 100 x 50 x 25mm. | [Please state dimensions of product.]
4.4.2 | Weight
| Not critical. | [Please state weight of product.]
4.5 | Interface requirements
None.
4.6 | Human factors
4.6.1 | Activation
| The product may be supplied already activated. Alternatively the user may be required to activate the device. Manual activation, however achieved, must be irreversible. | [Please state exact Section / Page number in the user manual which explains how the product is activated.]
4.6.2 | User interface for phase change products
| When the product is triggered this must be indicated by an irreversible colour change from a light to a dark colour. The colour change must take the form of a rapid and complete transformation from light to dark – products that retain an intermediate state will not be acceptable. The colour change must be readily distinguishable by users with all forms of colour blindness. | [If applicable: please state exact Section / Page number in the user manual which explains how the product is interpreted.]
4.6.3 | User interface for products with LED indicator(s) only
| The product is to have aminimum of one indicator light. The indicator or indicators must provide the means of unambiguous combinations of steady or flashing lights: • that the product is activated; • that the battery is functioning; • whether the temperature of the load has remained above 0°C OR; • whether the temperature of the load has fallen below 0°C. Clear instructions on interpreting the display must be printed on the product in a graphical form that is not language dependent. | [If applicable: please state exact Section / Page number in the user manual which explains how the product is interpreted.]
4.6.4 | User interface for products with LCD displays
| The product is to have an LCD display, with or without LEDs, capable of showing the following information: • that the product is activated; • that the battery is functioning; • whether the temperature of the load has remained above 0°C OR; • whether the temperature of the load has fallen below 0°C. The status of the load must be clearly and permanently indicated on the LCD. Indicator symbols must not be language-dependent and must be easily understood by untrained users. Acceptable indicators include, but are not confined to, the following: 'Tick’ or ‘OK’ symbol for temperatures above 0°C ‘Cross’ or ‘Crossed OK’ symbol for exposure to 0°C | [If applicable: please state exact Section / Page number in the user manual which explains how the product is interpreted.]
4.6.5 | Legibility
| It must be possible for a person with normal visual acuity (with or without glasses) to read the indicator both in bright sunlight and in tungsten/fluorescent lighting at 100 lux on the working plane, both before and after exposure to the trigger temperature. | [Please state where in the dossier to find evidence that the product meets this specification.]
4.6.6 | Mounting device
| The product should have a means for attaching it to the vaccine load – for example a self-adhesive strip or an eyelet. | [Please state where in the dossier to find evidence that the product meets this specification.]
4.7 | Materials
4.7.1 | Ozone depleting chemicals
| During manufacture and assembly of the printed circuit boards and final assembly of the product do not use any substance included in Annex A, B or C of the Montreal Protocol. | [Please submit evidence for compliance.]
4.7.2 | Other restricted materials
| The appliance and its constituent components, must not contain lead, mercury, cadmium, hexavalent chromium, polybrominated biphenyls (PBB) or polybrominated biphenyl ethers (PBDE). | [Please submit evidence for compliance.]
4.8 | Warranty
| The product is to be covered by a one year warranty in the event of any component failure not caused by mechanical damage. | [Please refer to your legally binding warranty document with terms.]
4.9 | Servicing provision
| The product is to be maintenance-free. | [Please confirm compliance with this PQS specification requirement.]
4.10 | Disposal and recycling
| The manufacturer is to provide information to the buyer on the hazardous materials contained within the system and suggestions for resource recovery/recycling and/or environmentally safe disposal. For the European Union WEEE compliance in accordance with European Union Directive 2002/96/EC is mandatory. | [Please show the information provided to the buyer on hazardous materials and for recycling and safe disposal.]
4.11 | Instructions
| One illustrated instruction sheet, in Arabic, English, French, Mandarin Chinese, Russian and Spanish, is to be supplied in the quantities requested by the end user at the time of ordering. The instructions are to be in large format on single sided ISO A4 or 8.5 inch x 11 inch poster format so that they can be used for staff training. | [Please enter the reference for the user manual and confirm compliance to the language requirement and state all options for receiving copies of the user instructions (including on a USB stick).]
4.12 | Training
| No requirement. | [There is no PQS requirement but please state whether training could be available e.g. via video tutorial etc.]
4.13 | Verification
| Verification: In accordance with PQS Verification Protocol E06/IN03.VP.1. | [Please enter the reference(s) for the laboratory test report(s).]
5. | Packaging
| Materials used for packaging the finished product are to be free of ozone depleting compounds as defined in the Montreal Protocol. | [Please describe the packaging and refer to the user manual section/page number on how this packaging may be disposed / recycled.]