## Spec
E006 Submission to PQS for apporval: Threshold indicators (V1.1 13MAR25) |  |  | [Brand & Model here]
| Tested in accordance with PQS verification protocol | E006/IN04-VP.1 (09 JUL 2018)
| To verify compliance with PQS specification below | E006/IN04.1 (28 JUN 2018)
ASSESSOR’S SUMMARY:
Instructions in [square brackets] are for your guidance. Please delete these instructions and replace with evidence or a precise reference (including page and section numbers) showing compliance with that particular PQS specification requirement. In general a "pass" or "complies" statement alone is not acceptable.
One document may provide evidence for several clause requirements. Please carefully reference the sub-sections of any such document so that it is clear that the clause requirement has been met. This will help speed the pre-qualification process.
Greyed-out cells do not require a response.
NOTE: If it is not clear or doubtful whether a requirement has been met, then this is a non-compliance and pre-qualification might be delayed.
| Please ensure the VP sheet is completed. |  | REVIEWER’S SUMMARY: |  | REVIEWER’S SUMMARY: |  | REVIEWER’S SUMMARY:
Section 7. | Product dossier | Applicant's submission | Reviewer's remarks [Name & Date] | Applicants response [Name & Date] | Reviewer's remarks [Name & Date] | Applicants response [Name & Date] | Reviewer's remarks [Name & Date]
The legal manufacturer or reseller must provide WHO with a prequalification dossier containing the following:
| Dossier examination fee in US dollars. | [Please contact PQS.]
| General information about the legal manufacturer, including name and address. | [Please state or refer to a document.]
| Unique identification reference for the product type. | [Please state.]
| Full specifications of the product being offered, covering all the requirements set out in this document, including details of product marking and traceability. | [Please list the document(s provided in the submission for pre-qualification).]
| Certified photocopies of the legal manufacturer’s ISO 9001 quality system certification. | [Please refer to a document.]
| Where relevant, certified photocopies of the legal manufacturer’s ISO 14001 certification, EMAS registration or registration with an equivalent environmental audit scheme. Conformity with an environmental audit scheme is not mandatory; however, preference will be given to manufacturers who are able to demonstrate compliance with good environmental practice. | [Please refer to a document.]
| Laboratory test report(s) proving conformity with the product specifications. | [Please refer to a document.]
| One sample of the product shipped in accordance with the manufacturer's instructions together with product instruction insert – both in English language. | [Please send to PQS and direct to the product reviewers – details to be confirmed.]
| Indicative cost of the product per 100 units, per 1,000 units and per 10,000 units EXW (Incoterms 2015). | [Please state or refer to a document.]
Section 4. | Requirement
4.1 | General
| Threshold indicators take the form of a visual indicator which may be attached to a backing card. | [Please provide photo(s) and describe the form and function of your threshold indicator.]
4.2 | Performance
4.2.1 | Activation
| The indicator may be activated by means of a physical removal of an “activation strip” or other positive user intervention. Activatable indicator must visibly demonstrate when the indicator is in an activated or inactivated state. No visual change must occur unless the indicator is activated. | [State whether continuously active or otherwise. If the indicator needs to be activated, refer to the user instruction that explains this process and also explains how the user can be clear that the indicator is inactive or active.]
4.2.2 | Mode of operation:
| The indicator must exhibit an irreversible rapid visual change when the temperature has exceeded the threshold temperature within 15 minutes as specified in Clause 4.2.3 and Clause 4.2.4. The visual change can take the form of any of the following: – change in shade – change in colour – reveal or obscure text or symbols | [Please provide photos and describe the visual change.]
4.2.3 | Threshold temperatures
| +XY°C indicator: Full visual change must take place when the temperature reaches XY°C, where XY = [threshold temperature]. | [Please refer to test evidence that demonstrate this. Cross reference: E006/IN04.VP.1 Clause 5.2.3 Test 3.]
4.2.4 | Tolerance
| ±1.0°C. The tolerance is set such that no more than 1% of indicators shall show visual change at the specified temperature minus 1.0°C, and no less than 99% shall show visual change at the specified temperature plus 1.0°C (e.g. for a 40°C indicator, ≤1% shall reach endpoint at 39°C and ≥99% shall show visual change at 41°C). | [Please refer to test evidence with a statistical analysis that demonstrate this. Cross reference: E006/IN04.VP.1 Clause 5.2.3 Test 3.]
4.2.5 | Operating (activated) temperature range:
| Upper limit (without indicator reversion): 70°C. Lower limit (without product failure): – 30°C. | [Please refer to test evidence that demonstrate this. Cross reference: E006/IN04.VP.1 Clause 5.2.4 & Clause 5.2.5, Test 4 & Test 5.]
4.2.6 | Quality control
| Indicator production must follow Good Manufacturing Practices with appropriate testing controls. Manufacturers and supply chain users can verify indicator performance by sampling and testing of multiple production lots in accordance with AQL standards. | [Please state the AQL standards used and give details of how this is applied to production line samples. Test results must be provided.]
4.2.7 | Power source
| None. | NOTE: This must be a passive device.
4.2.8 | Shelf life
| Minimum 3 years from date of manufacture, including operational life, when stored under recommended conditions. | [Please provide evidence for the shelf-life of the indicator. This might include, but is not limited to, showing samples that have been stored for 3 years and are still fully functional.]
4.3 | Environmental requirements
4.3.1 | Ambient temperature range during transport and storage (not activated)
| -30°C to +70°C. | [Please refer to test evidence that demonstrate this. Cross reference: E006/IN04.VP.1 Clause 5.2.2 Test 2.]
4.3.2 | Ambient humidity range during transport and storage (not activated)
| 5%RH to 95% RH. | [Please provide evidence for stability at these humidities.]
4.3.3 | Durability
| The product is to be constructed of materials that are adequately robust and durable for the intended use. The indicator should be unaffected by anticipated impact, pressure or bending e.g. caused by coolant packs or the weight of vaccine vials. This includes indicators affixed to a vaccine vial. | [Please provide evidence for the robustness and durability. Cross reference: E006/IN04.VP.1 Clause 5.2.6 Test 6.]
4.3.4 | Moisture
| The indicator and backing card (if supplied) to be water resistant. | [Please refer to test(s) which demonstrate water resistance. Cross reference: E006/IN04.VP.1 Clause 5.2.6 Test 6.]
4.4 | Physical characteristics
4.4.1 | Overall dimensions
| To be defined with and without backing card if supplied. | [Please state the dimensions or refer to a specification document.]
4.4.2 | Weight
| Not critical. | [Please state the weight.]
4.5 | Human factors
4.5.1 | Appearance
| When the threshold temperature is exceeded, there must be an irreversible visual change which must take the form of a complete, rapid and permanent transformation. Products that retain an intermediate state will not be acceptable. The visual change must be of high contrast so that it is readily distinguishable by users with colour blindness. Examples of visual changes are specified in Clause 4.2.2. Care should be exercised if pictogram symbols are used. The visual change must avoid possibilities for ambiguities and confusion in the culture and region of use. For example, if a ‘Tick’ or ‘OK’ symbol is obscured when Clause 4.2.3 conditions are met, the absence alone of a symbol may not be -clearly understood as “do not use”. Alternatively, an indicator could reveal a “do not use” pictogram when Clause 4.2.3 conditions are met. | [Please refer to test evidence that demonstrates that there is an irreversible visual change in the form of a complete, rapid and permanent transformation and that the visual change is of high contrast. Cross reference: E006/IN04.VP.1 Clause 5.2.4 & Clause 5.2.5, Test 4 & Test 5.]
4.5.2 | Backing card
| If supplied, the backing card material should be water resistant and contain appropriate instructions as stated in Clause 4.9. | [Please refer to test evidence that demonstrate that the backing card material is water resistant . Cross reference: E006/IN04.VP.1 Clause 5.2.6 Test 6.]
4.5.3 | Legibility
| It must be possible for a person with normal visual acuity (20/20 *with or without corrective lenses) to interpret the visual change of the indicator both in bright sunlight and in tungsten/fluorescent lighting at 100 lux on the working plane, before and after activating and after reaching the end point. | [Please refer to test evidence that demonstrates it is possible to correctly interpret the visual change of the indicator both in bright sunlight and in artificial lighting. Cross reference: E006/IN04.VP.1 Clause 5.2.7 Test 7.]
4.5.4 | Mounting
| The product should have a means of attachment, for example, a self-adhesive strip. The indicator response should not be affected by orientation. | [Please state what means of attachment are available.]
4.6 | Materials and disposal
| Materials used must be non-toxic and non-irritant to the end user and harmless to the environment. The product must meet all requirements concerning toxicity of materials and packaging in force in the country of manufacture. | [Please state any potentially toxic materials both in use and disposal. Also refer to user manual where it shows correct procedure for disposal.]
4.7 | Warranty
| The product is to be maintenance-free and all batches of the product must be warranted to conform to the requirements of this specification for the specified shelf life of at least three years for the product when stored in accordance with the manufacturer's instructions. | [Please refer to the warranty with full terms and conditions.]
4.8 | Instructions
| Illustrated instruction sheets in English, to be supplied in every carton. Instructions printed in other languages to be supplied on client's request. | [Please refer to the user instructions.]
| Storage instructions for active and non-active indicators. | [Please refer to a particular section in the user instructions.]
| How to activate the indicator or a clear instruction that the indicator is always active. | [Please refer to a particular section in the user instructions.]
| Visual state of the indicator when the temperature has not exceeded manufacturer's specified threshold temperature. | [Please refer to a particular section in the user instructions.]
| Visual state of the indicator when the temperature has exceeded manufacturer's specified threshold temperature. | [Please refer to a particular section in the user instructions.]
| If supplied, backing cards to have the above instructions which can be diagrammatic. | [Please refer to a particular section in the user instructions.]
4.9 | Training
| The product manufacturer to provide training for the vaccine manufacturer so that the manufacturer can correctly handle, apply and test the indicator. | [Please state the options for how training can be made available.]
4.10 | Verification
| In accordance with PQS Verification Protocol E006/IN04-VP.1. | [Please refer to the test report.]
5. | Packaging
| Materials used for packaging the finished product are to be free of ozone-depleting compounds as defined in the Montreal Protocol. | [Please state the materials used for packaging the finished product and whether they are free of CFC compounds as defined in the Montreal Protocol. Also refer to user manual where it shows correct procedure for disposal of packaging.]

## VP
E006 Submission to PQS for apporval: Threshold indicators (V1.1 13MAR25) |  |  | [Brand & Model here]
| Tested in accordance with PQS verification protocol | E006/IN04-VP.1 (09 JUL 2018)
| To verify compliance with PQS specification below | E006/IN04.1 (28 JUN 2018)
ASSESSOR’S SUMMARY:
Instructions in [square brackets] are for your guidance. Please delete these instructions and replace with evidence or a precise reference (including page and section numbers) showing compliance with that particular PQS specification requirement. In general a "pass" or "complies" statement alone is not acceptable.
One document may provide evidence for several clause requirements. Please carefully reference the sub-sections of any such document so that it is clear that the clause requirement has been met. This will help speed the pre-qualification process.
Greyed-out cells do not require a response.
NOTE: If it is not clear or doubtful whether a requirement has been met, then this is a non-compliance and pre-qualification might be delayed.
Please ensure the Spec sheet is completed.
| VP Clause | Applicant's submission | Assessor's remarks [Name & Date] | Applicants response [Name & Date] | Assessor's remarks [Name & Date] | Applicants response [Name & Date] | Assessor's remarks [Name & Date]
5.2.1 | Test 1: Type examination
| Acceptance criteria: Type examination indicates full conformity with all specification requirements. | [Applicant to check that the test report contains all clause component of the type examination with evidence for compliance. Please provide an accurate test report reference.]
5.2.2 | Test 2: Stability during storage
| Acceptance criterion: No visual change in any test sample at the end of Step 4 or Step 7 with the exception of the always-active samples exposed to +70°. All samples to show visual change at the end of Step 8. | [Applicant to check that the test report results clearly show that the threshold indicator meets the storage criteria and then functions correctly. Please provide an accurate test report reference.]
5.2.3 | Test 3: Temperature accuracy test
| Acceptance criteria: Visual change after completion of Step 3 to be complete within ± 1.0 °C of the threshold temperature within the time limit required by specification Clause 4.2.5. Visual change to be maintained throughout Step 4. | [Applicant to check that the test report results clearly show that the threshold indicator accurately transitions at the pre-determined temperature and remains in the transitioned state. A statistical analysis of the test results is helpful. Please provide an accurate test report reference. Cross-reference E006/IN04.1 Clause 4.2.3 & Clause 4.2.4.]
5.2.4 | Test 4: Low ambient temperature test
| Acceptance criterion: All test samples to continue to display the visual change. | [Applicant to check that the test report results clearly show that the threshold transition is permanent when exposed to a cold ambient. Please provide an accurate test report reference.]
5.2.5 | Test 5: High ambient temperature test
| Acceptance criterion: All test samples to maintain the visual change following removal from test chamber. | [Applicant to check that the test report results clearly show whether the threshold transition is permanent when exposed to a hot ambient. Please provide an accurate test report reference.]
5.2.6 | Test 6: Soak test
| Acceptance criterion: No water damage observed. All samples to meet the criteria of Test 3: Step 3. Visual appearance to be maintained throughout Test 3: Step 4. | [Applicant to check that the test report results clearly show whether the threshold indicator transition is accurate after soaking. Please provide an accurate test report reference.]
5.2.7 | Test 7: Observer perception test
| Acceptance criterion: All five observers are able to sort the two groups of test samples with 100% accuracy under both lighting regimes. The mean time taken to complete the task should not exceed two minutes. | [Applicant to check that results from the test report clearly show whether users can correctly interpret the indicator under different lighting conditions. Please provide an accurate test report reference.]
5.3 | Test criteria for qualification
The report of the tests must contain the following data and analyses:
| Summary: Conclusions and recommendations. | [Applicant to check that this information is in the report.]
| Test 1: Results of type examination, tabulated data and photographs of samples. | [Applicant to check that this information is in the report.]
| Test 2: Results of storage stability test. | [Applicant to check that this information is in the report.]
| Test 3: Results of temperature accuracy test. | [Applicant to check that this information is in the report.]
| Test 4: Results of low temperature test. | [Applicant to check that this information is in the report.]
| Test 5: Results of high temperature test. | [Applicant to check that this information is in the report.]
| Test 6: Results of water-resistance test. | [Applicant to check that this information is in the report.]
| Test 7: Results of observer perception test. | [Applicant to check that this information is in the report.]
| Annexes: Information from manufacturer as examined in Test 1. Additional supporting documentation requested and received from the legal manufacturer or reseller during the course of the type-testing. Test chamber temperature records and laboratory instrumentation calibration records to be kept on file for later scrutiny. | [Applicant to check that this information is either in the report as requested and that other information that is kept on file can be available for later scrutiny.]