## Spec
E006 Vaccine vial monitor submission template (V1.1 14 MAR 2025)
| [Company name] | [Model name] | [Date]
| Dossier submitted by: | [Please enter your name, position within the company, and email address]
| Tested in accordance with PQS verification protocol | E006/IN05/VP.4 (19 OCT 2020)
| To verify compliance with PQS specification below | E006/IN05.4 (19 OCT 2020)
ASSESSOR’S SUMMARY:
Instructions in [square brackets] are for your guidance. Please delete these instructions and replace with evidence or a precise reference (including page and section numbers) showing compliance with that particular PQS specification requirement. In general a "pass" or "complies" statement alone is not acceptable.
One document may provide evidence for several clause requirements. Please carefully reference the sub-sections of any such document so that it is clear that the clause requirement has been met. This will help speed the pre-qualification process.
Greyed-out cells do not require a response.
NOTE: If it is not clear or doubtful whether a requirement has been met, then this is a non-compliance and pre-qualification might be delayed.
|  |  | REVIEWER’S SUMMARY: |  | REVIEWER’S SUMMARY: |  | REVIEWER’S SUMMARY:
Section 7. | Product dossier | Applicant's submission | Reviewer's remarks [Name & Date] | Applicants response [Name & Date] | Reviewer's remarks [Name & Date] | Applicants response [Name & Date] | Reviewer's remarks [Name & Date]
The legal manufacturer or reseller is to provide WHO with a pre-qualification dossier containing the following:
| Dossier examination fee in US dollars. | [Please contact PQS.]
| General information about the legal manufacturer, including name and address. | [Please state or refer to a document.]
| Unique identification reference for the product type. | [Please state.]
| Full specifications of the product being offered, covering all the requirements set out in this document, including details of product marking and traceability. | [Please refer to specification details.]
| Details of the legal manufacturer’s internal AQL sampling procedures in respect of ISO 3951 and ISO 2859. | [Please refer to a document giving complete details.]
| Certified photocopies of the legal manufacturer’s ISO 9001 quality system certification. | [Please refer to a document.]
| Where relevant, certified photocopies of the legal manufacturer’s ISO 14001 certification, EMAS registration or registration with an equivalent environmental audit scheme. Conformity with an environmental audit scheme is not mandatory; however, preference will be given to manufacturers who are able to demonstrate compliance with good environmental practice. | [Please refer to a document.]
| Laboratory test report(s) proving conformity with the product specifications. | [Please refer to a document.]
| A minimum of five samples of each of the types of VVM shipped in accordance with manufacturer’s instructions, together with product instruction insert in English language. | [Please send to PQS and to the assessors – details to be confirmed.]
| Indicative cost of the product per 10,000 units, per 100,000 units and per 1,000,000 units EXW (Incoterms 2015). | [Please state or refer to a document.]
Section 4. | Requirement
4.1.1 | General
The VVM comprises an outer reference surface and an inner active surface. The change of the active surface is limited to a change of shade, from light to dark. Any colour is permitted for the VVM design, but changes in hue when activated are not permitted.
4.1.2 | Format and dimensions
| Please see Annex 1 the VVM design specification. | [Please state the dimensions or refer to a specification document.]
4.1.3 | Activation
| The VVM may be continuously active or made into an activated state e.g. by means of a physical removal of a “trigger strip” or other positive user intervention. Manufacturer must provide instructions and controls to ensure that it is visibly obvious whether the VVM is in an activated or inactivated state. Only activatable VVMs require visibly obvious cues. | [State whether continuously active or otherwise refer to instructions on how it is obvious which state the VVM is in.]
4.2 | Performance
4.2.1 | Optical density change
The optical density change of the indicator is illustrated in the Figure 1 of the PQS Specification. At the start point the shade of the active surface is lighter than the reference surface. The end point is indicated when the shade of the active surface matches the shade of the reference surface. The end point is exceeded when the shade of the active surface is darker than the reference surface. Clauses 4.2.2 to 4.2.5 describe allowed parameters for the shade change.
4.2.2 | Optical density at start point and end point
| • At the start point, the optical density of the active surface must be lower than the optical density of the reference surface by a difference of at least 0.23. | [Please refer to tests which demonstrate this.]
| • The end point is reached when the difference in the average optical density obtained from readings at two different points on the reference surface and the optical density of two different points of the active surface is 0.00. The end point is exceeded when the OD of the active surface is higher than OD of the reference surface. | [Please refer to test(s) which demonstrate this.]
| • The specifications for the Start OD of both the reference and active surfaces and the OD difference between the two surfaces (R-I) with their permitted tolerances must be specified by the manufacturer of the VVMs. Tolerances should include any effect from the liner. | [Please refer to specifications for the Start OD of both the reference and active surfaces and the OD difference between the two surfaces (R-I) with their permitted tolerances.]
4.2.3 | Homogeneity of the reference surface
| The optical density of one 2 mm diameter portion of the reference surface must be within 0.03 of the optical density at any other two 2 mm diameter portions of the reference surface, when measured with a densitometer (with a 2 mm diameter aperture). | [Please refer to test(s) which demonstrate this.]
4.2.4 | Variation of the reference surface within the lot
| The average optical density of at least two measurements in different positions on the reference surface of one sample must be within 0.03 of a similar average optical density of the reference surface of any other sample within the same lot. | [Please refer to test(s) which demonstrate this.]
4.2.5 | VVM reaction rates
| Reaction rates specific to different example types of VVM, relating to groups of vaccines according to their stability at two specific temperature points in Table 1a and Table 1b. Note that VVM0.5 and VVM½ may be used interchangeably. | [Please refer to the mathematical model of the reaction rate for your VVM.]
Table 1a specifications:
| • (Applies to all except VVM0.5) At +37°C, RH 33% ± 5% and RH 75% ± 5%, the primary limits are set such that at the specified time no more than 5% of VVMs shall reach end point at the lower temperature limit of 35°C and no less than 95% shall reach end point at the upper temperature limit of 37°C. Further, secondary limits are applied to restrict how far beyond the primary specification the end points are allowed to be. At the specified time no more than 0.1% of VVMs shall reach end point at the lower secondary temperature limit of 34.5°C and no less than 99.9% shall reach end point at the upper secondary temperature limit of 37.5°C. | [Please refer to test(s) which demonstrate this.]
| • (Applies to VVM0.5 only) At +25°C, RH 33% ± 5% and RH 75% ± 5%, the primary limits are set such that at the specified time no more than 5% of VVMs shall reach end point at the lower temperature limit of 23°C and no less than 95% shall reach end point at the upper temperature limit of 25°C. Further, secondary limits are applied to restrict how far beyond the primary specification the end points are allowed to be. At the specified time no more than 0.1% of VVMs shall reach end point at the lower secondary temperature limit of 22.5°C and no less than 99.9% shall reach end point at the upper secondary temperature limit of 25.5°C. | [Please refer to test(s) which demonstrate this.]
| • At +5°C (applies to VVM2, VVM1 and VVM0.5) and +25°C (applies to VVM7, VVM11, VVM14 and VVM30) specifications (ambient humidity in submerged foil/polythene pouch): Limits are set such that at the specified time no more than 5% of VVMs shall reach end point at the lower temperature limit (2°C or 22°C, respectively) and no less than 95% shall reach end point at the upper temperature limit (5°C or 25°C, respectively). | [Please refer to test(s) which demonstrate this.]
Table 1b specifications:
| • At +55°C, RH 33% ± 5% and RH 75% ± 5%, the primary limits are set such that at the specified time no more than 5% of VVMs shall reach end point at the lower temperature limit of 53°C and no less than 95% shall reach end point at the upper temperature limit of 55°C. Further, secondary limits are applied to restrict how far beyond the primary specification the end points are allowed to be. At the specified time no more than 0.1% of VVMs shall reach end point at the lower secondary temperature limit of 52.5°C and no less than 99.9% shall reach end point at the upper secondary temperature limit of 55.5°C. | [Please refer to test(s) which demonstrate this.]
| • At +45°C specifications (ambient humidity in submerged foil/polythene pouch): Limits are set such that at the specified time no more than 5% of VVMs shall reach end point at the lower temperature limit (42°C) and no less than 95% shall reach end point at the upper temperature limit (45°C). | [Please refer to test(s) which demonstrate this.]
Table 2 defines the allowable range of R-I at the specified time and specified temperature consistent with the VVM reaction rate specifications.
4.2.6 | Overall uncertainty
| The expected total uncertainty for measuring the difference between the optical density of the reference surface and active surface is ± 0.03. The measurement uncertainty for a single measurement is ± 0.02. Major sources of uncertainty are instrument error both for the reference surface and the active surface, repeatability, and variation in end point caused by an allowed temperature tolerance of ± 0.2°C in the temperature bath. (The VVMs should be tested in water baths or incubators whose mean kinetic temperature is controlled to within ± 0.2°C.) | [Please refer to a document giving complete details of the uncertainty estimations.]
4.2.7 | Reversion
| The indicator must not revert to a lighter shade at any point in its life when exposed to conditions likely to be found during normal use. After the endpoint is reached, the active surface must remain the same shade or darker than the reference surface. | [Please refer to test(s) which demonstrate this.]
4.2.8 | Integrity and location of VVMs
| Before a vial or ampule is opened, the VVM should not be removable; it should resist removal from the vaccine vial as much as a label meeting current vaccine labelling requirements. In addition, the performance of the VVM should not be changed by soaking in water for 8 hours. The optical density of water-exposed samples should conform to within ±0.04 or ±10% of initial (R – I) whichever is the greater. | [Please refer to test(s) which demonstrate this.]
The location of the VVM on the vial depends upon whether the vaccine must be discarded at the end of the immunization session in which it is opened, or whether any remaining contents in an opened vial can be retained for use in subsequent sessions. The following cases apply:
| • For multi-dose vials containing a vaccine that can be used in subsequent sessions: Regardless of the vaccine presentation (liquid, freeze-dried or two vial combinations of liquid and freeze-dried), the VVM must be permanently attached to the label of the vaccine vial and must remain readily observable before, during, and after use, until the entire contents of the vial have been used. | [Please refer to test(s) which demonstrate this.]
| • For vaccines that must be discarded at the end of the session or within 6 hours, whichever comes first: The VVM must be attached to the vaccine vial or ampule and must remain readily observable until the vial or ampule is opened, but not observable after opening. In order to achieve this requirement, the VVM must be located on the flip-off top of a vial or on the neck of an ampule. | [Please refer to test(s) which demonstrate this.]
| • On a product by product basis, WHO will advise both the vaccine and the VVM manufacturer where the VVM is to be located. Locating the VVM on the bottom of a vial or ampule is never acceptable – it must always be in a visible location. | [Please refer to test(s) which demonstrate this.]
4.2.9 | Application Surfaces
VVMs should be designed to be applied to the following substrates: Vaccine manufacturers must ensure permanent adhesion to the vaccine container. Note: Users should check there is adequate adhesion of the VVM to the vaccine container.
| • Glass (e.g. tops of ampules). • Vial labels. • Plastic containers of a composition for which permeation of adhesive components is not a risk. • Foil pouches. • For vial cap applications, VVMs should be designed to be applied to smooth, flat plastic surfaces with no embossed areas, recessed areas, or ridges. The use of excessive release agents in the manufacture of the vial caps should be avoided. | [Please refer to test(s) which demonstrate adequate adhesion of the VVM to these surfaces.]
4.3 | Traceability
| Each roll of VVMs must be labelled with its product identity (part number) together with its lot number. | [Please show examples of VVMs labelled with its product identity (part number) together with its lot number.]
4.4 | Interface requirements
None.
4.5 | Human factors
| The shade change must be monochromatic in its response to cumulative time-temperature exposure within the limits of the allowed shade variation. The observer must be able to distinguish between an unchanged indicator, an intermediate shade change and the end point of the indicator. There must be a gradual change in the shade of the active surface. | [Please refer to test(s) which demonstrate this.]
4.6 | Materials
| Materials used must be non-toxic and non-irritant to the end user and substantially harmless to the environment. The VVM must meet any requirements in force concerning toxicity of labels or packaging in the country of manufacture. | [Please state any potentially toxic materials both in use and disposal.]
4.7 | Servicing & warranty
| The product is to be maintenance-free and all batches of the product must be warranted to conform to the requirements of this specification for the specified shelf life of the product when stored in accordance with the manufacturer’s instructions. | [Please refer to the warranty statement.]
4.8 | Disposal and recycling
| The product will be disposed of in conjunction with the vial to which it is attached. | [Please refer to safe disposal instructions.]
4.9 | Instructions
| An instruction insert, providing vaccine manufacturers with all necessary storage, handling and application directions and Traceability directions (with reference to clause 4.3) is to be supplied with every carton containing rolls of VVMs. The insert is to be printed in English. If any vaccine manufacturer requires an instruction insert in an additional language, this will be a matter for independent negotiation between the VVM manufacturer and the vaccine manufacturer. | [Please refer to a copy of the instruction insert.]
4.1 | Training
| The VVM manufacturer must provide training for the vaccine manufacturer in order that the manufacturer can correctly handle, apply and test VVMs. | [Please state the options for how training can be made available.]
4.11 | Verification
| In accordance with PQS Verification Protocol E006/IN05.VP.4. | [Please refer to the test report.]
5. | Packaging
| Materials used for packaging the finished product are to be free of CFC compounds as defined in the Montreal Protocol. | [Please state the materials used for packaging the finished product and whether they are free of CFC compounds as defined in the Montreal Protocol.]