## Sheet1
E006 Submission to PQS for apporval: Integrated electronic thermometer, with or without alarm function, for vaccine refrigerators and freezers (V2.0 14MAR25) |  |  |  | [Brand & Model here]
| Tested in accordance with PQS verification protocol | E006/TH06-VP.2
| To verify compliance with PQS specification below | E006/TH06.2 (JULY 2010)
ASSESSOR’S SUMMARY:
SCOPE
This specification describes the performance requirements for an integrated electronic thermometer, with or without alarm function, for vaccine refrigerators and freezers. This product is intended for use as the primary means for monitoring storage conditions in vaccine refrigerators and freezers in intermediate stores and health facilities. It can be offered in one or more of the following configurations: • Type A: Permanently integrated device, supplied by the manufacturer of an electric refrigerator or freezer with a mains power supply. • Type B: Permanently integrated device, supplied by the manufacturer of a solar refrigerator or freezer powered by a battery pack charged from a photovoltaic array. • Type C: Permanently integrated device, supplied by the manufacturer of a solar direct drive refrigerator or freezer or a gas or kerosene-fueled absorption refrigerator or freezer.
Instructions in [square brackets] are for your guidance. Please delete these instructions and replace with evidence or a precise reference (including page and section numbers) showing compliance with that particular PQS specification requirement. In general a "pass" or "complies" statement alone is not acceptable.
One document may provide evidence for several clause requirements. Please carefully reference the sub-sections of any such document so that it is clear that the clause requirement has been met. This will help speed the pre-qualification process.
Greyed-out cells do not require a response.
NOTE: If it is not clear or doubtful whether a requirement has been met, then this is a non-compliance and pre-qualification might be delayed.
| Please ensure the VP sheet is completed. |  | REVIEWER’S SUMMARY: |  | REVIEWER’S SUMMARY: |  | REVIEWER’S SUMMARY:
Section 7. | Product dossier | Applicant's submission | Reviewer's remarks [Name & Date] | Applicants response [Name & Date] | Reviewer's remarks [Name & Date] | Applicants response [Name & Date] | Reviewer's remarks [Name & Date]
This product will be offered as an integrated component of a refrigerator or freezer complying with one of the E003/ series performance specifications. The product dossier requirements below are to be read as a supplement to the relevant E003/ performance specification.
The legal manufacturer or reseller must provide WHO with a prequalification dossier containing the following:
| Dossier examination fee in US dollars. | [Please contact PQS to arrange payment.]
| General information about the legal manufacturer, including name and address. | [Please enter this information, include for all parties, e.g. manufacturer, reseller, plus other.]
| Unique identification reference for the product type. | [Please enter this information.]
| Full specifications of the product being offered, covering all the requirements set out in this document, including details of product marking and traceability. | [The dossier must have clear traceable references to all information required below.]
| Certified photocopy of Certificate of Traceability and Calibration traceable to an ISO/IEC 17025 accredited testing laboratory, to NIST, or to another internationally recognized standards agency. | [Please enter the reference for the supporting document in the dossier.]
| Certified photocopies of all type-approvals obtained for the appliance and/or its components, including CE marking and the like. | [Please enter the reference for supporting documents in the dossier. A simple statement of compliance is not acceptable.]
| Certified photocopies of the legal manufacturer’s ISO 9001 quality system certification. | [Please enter the reference for the supporting document in the dossier.]
| Where relevant, certified photocopies of the legal manufacturer’s ISO 14001 certification, EMAS registration or registration with an equivalent environmental audit scheme. Conformity with an environmental audit scheme is not mandatory; however, preference will be given to manufacturers who are able to demonstrate compliance with good environmental practice. | [Please enter the reference for the supporting document in the dossier.]
| Where available, laboratory test report(s) proving conformity with the product specifications. | [Please enter the reference for the supporting document in the dossier.]
| Sample refrigerator or freezer incorporating the product as defined in the relevant E003/ series performance specification. | [Please send 1 sample to each of the addresses to be provided.]
Section 4. | Requirement
4.1 | General
| Electronic thermometer for monitoring storage conditions in vaccine refrigerators and freezers. These devices are supplied by the refrigerator or freezer manufacturer as a permanently integrated component of the product.The following configurations are covered by this specification: • Type A1: Electronic thermometer fitted to a mains powered refrigerator or freezer. • Type A2: As Type A1, but with a high and low alarm function. • Type B1: Electronic thermometer fitted to a photovoltaic refrigerator or freezer with a solar charged battery pack. This type is NOT suitable for direct drive photovoltaic refrigerators as described in PQS specification E003/RF05, nor is it suitable for gas or kerosene fueled absorption refrigerators. • Type B2: As Type B1, but with a high and low alarm function. • Type C: Electronic thermometer powered by a photovoltaic cell which forms part of the device. This type draws no power from the product in which it is installed and is therefore suitable for use with gas and kerosene fueled absorption refrigerators and freezers, and with direct drive photovoltaic refrigerators as described in PQS specification E003/RF05 and E003/RF06. | [Please state Type and how this complies.]
4.2 | Performance
4.2.1 | Operating temperature range
| Operating temperature range: Upper limit: +50°C with over-range protection as clause 4.3.1. Lower limit: -30°C with under-range protection as clause 4.3.1. | [Please enter the reference for supporting evidence in the dossier.]
4.2.2 | Accuracy
| Type A and Type B devices: ±0.5°C or better within the range -30°C to +20°C. Type C devices: ±1.0°C or better within the range -30°C to +20°C. | [Please enter the reference for the calibration report in the dossier. See also Clause 4.2.9.]
4.2.3 | Resolution
| ±0.1°C or better within the range -30°C to +20°C | [Please enter the reference for supporting evidence in the dossier.] Note that resolution is the smallest difference that can be measured, not necessarily the smallest increment on the display.
4.2.4 | Power source
| • Type A: Power is taken from the 110/240 volt 50/60 Hz mains supply. The device must be fitted with a rechargeable back-up battery in the event of power failure. The battery must be capable of powering the device for a minimum period of 72 hours in the absence of mains power. • Type B: Power is taken from the main solar-charged battery pack or from the ancillary battery of a direct drive solar refrigerator with ancillary battery (PQS specification E003/RF06). An additional rechargeable back-up battery is not required. • Type C: Power is supplied by a photovoltaic cell which forms part of the thermometer. The photovoltaic cell must be able to operate the device at an ambient lighting level of 100 lux, or when a torch is shone onto the device. A response delay of up to 30 seconds is acceptable when a torch is used to activate the device when it is dormant. | [Please state the power source for the thermometer.]
4.2.5 | Sensor
| One or more electronic sensors permanently located in the refrigerator or freezer cabinet to give a true worst-case reading of the load temperature. The worst-case reading in refrigerators is to be the lowest temperature measured during testing and the worst-case reading in freezers is to be the highest temperature measured during testing. The reading must refer to a point lying within the zone allocated for the storage of vaccines. | [Please state where the sensor is located. If there are more than one sensors, indicate which is connected to the external device readout.]
4.2.6 | Product response time
| T90 20 minutes maximum in accordance with EN12830:1999 | [Please enter the reference for test evidence in the dossier.]
4.2.7 | Calibration
| Each product is to be covered by a Certificate of Traceability and Calibration. The traceability declaration is to confirm that the measurement standards and instruments used during calibration of the product are traceable to an ISO/IEC 17025 accredited testing laboratory, to NIST, or to another internationally recognized standards agency. The certificate must be accompanied by a copy of the reference instrument calibration certificate. | [Please enter the reference for the calibration report in the dossier. See also Clause 4.2.3.]
4.2.8 | Unit of measurement
| Temperatures must be recorded in degrees Centigrade. | [Please state the unit of measurement.]
4.2.9 | Logging interval
| Devices which measure the storage temperature at discreet intervals are to sample at an interval not greater than 10 minutes. | [If applicable, please state the logging intervals for the logger.]
4.2.10 | Mode of operation
| The device is to display the most recently logged temperature inside the refrigerator or freezer cabinet. If the temperature in the cabinet exceeds the high alarm setting or drops below the low alarm setting the device is to signal the relevant alarm condition(s) until the alarm is reset by the user. | [Please state exact Section / Page number in the user manual that shows where the user accesses the temperature readout.]
4.2.11 | Alarm
| Type A2 and Type B2 devices must include a high breach and low breach visual alarm event indicator. An audible alarm is optional. | [If applicable, please state exact Section / Page number in the user manual that describes the alarm functions]
4.2.12 | Alarm settings
| Upper and lower alarm settings in Type A2 and Type B2 devices must be factory programmed into the device as follows: • Version for vaccine refrigerators: a) Low alarm setting: Exposure to a single temperature event of -0.5°C or below for 60 minutes. b) High alarm setting: Exposure to a single temperature event of +10°C or above for 10 hours. • Version for vaccine freezers: a) High alarm setting: Exposure to a single temperature event above -15°C for 60 minutes. A low alarm is not required. Once an alarm event has occurred, the alarm event indicator is to remain activated until cancelled by the user, even if the temperature inside the cabinet has subsequently returned to within acceptable limits. | [If applicable, please state exact Section / Page number in the user manual that describes the alarm functions]
4.2.13 | Casing
| Non-corrodible plastics or metal case. | [Please state material.]
4.2.14 | IP rating
| Protection of the product not less than IEC 60529: IP54 | [Please enter the reference for the supporting IP test report.]
4.2.15 | Battery
| For Type A devices only: As defined in clause 4.2.4. Rechargeable batteries must be easily replaceable without need for tools. | [If applicable, please enter the reference for supporting evidence for battery life. Please state the Section / Page number of the user manual that describes battery life indication and replacement instructions.]
4.2.16 | Electromagnetic compatibility
| Operation of the device must be unaffected in the normal electromagnetic compatibility environment in which it is intended to work, taking into account disturbance generated by adjacent apparatus which is compliant with relevant ISO, EN, ANSI, or other internationally recognized standards. Information required to ensure uninterrupted use of the device must be contained in the user instructions. | [Please enter the reference for the supporting EMC test report.]
4.3 | Environmental requirements
4.3.1 | Ambient temperature range during transport and storage
| -30°C to +55°C with device inactivated. | [Please enter the reference for supporting evidence.]
4.3.2 | Ambient humidity range during transport, storage and use
| 5 to 95% RH, non-condensing. | [Please enter the reference for supporting evidence.]
4.3.3 | Resistance to electrical storms
| The functionality of the device must not be affected by intense electrical storm activity. | [Please enter the reference for supporting evidence.]
4.4 | Physical characteristics
4.4.1 | Overall dimensions
| Not critical. | [Please state the dimensions.]
4.4.2 | Weight
| Not critical. | [Please state the weight.]
4.5 | Interface requirements
The device must be capable of being incorporated into the control panel of a refrigerator or freezer in a position to be decided by the refrigerator/freezer manufacturer.
4.6 | Human factors
4.6.1 | User interface
| General requirements: • The digital temperature read-out and optional visual alarm is to be integrated into the control panel on the refrigerator or freezer cabinet. • The display is to be legible by a person with normal vision (with or without glasses) standing in front of the cabinet, both in bright sunlight and in tungsten/fluorescent lighting at 100 lux on the working plane. Type A1 and Type A2 devices: • There must be a ‘replace battery’ warning display which is triggered when 10% or 90 days of battery life remains (whichever period is the longer). Type A2 and Type B2 devices: • The visual alarm must unambiguously show both high and low temperature breaches. • The optional audible alarm must achieve a minimum sound intensity of 100dB(A) at a distance of one metre from the sounder. The pattern of the signal is to be an intermittent pulse. • There must be a switch on the control panel to reset the alarm. It must not be possible to confuse this with the any other control on the refrigerator or freezer. | [Please explain or refer to a precise page and section in the user manual which shows: • the visual display with temperature and any visual or audible alarms. • how to reset the alarms • the 'replace battery' warning display if applicable]
4.6.2 | Activation
| • Type A1 and Type A2: The device is to be activated by connection of the refrigerator or freezer to a mains electricity supply. In the event of mains failure or loss of photovoltaic power it is to remain fully activated until the supply is restored or the charge on the backup battery expires. • Type B1 and Type B2: The device is to be activated by connection of the refrigerator or freezer to a photovoltaic electricity supply. Backup power in low sun conditions is supplied by the main battery pack. • Type C: The device is to be activated by the ambient lighting, whether natural or artificial, in the room in which the product is installed. In health facilities without electricity, nighttime activation of a dormant device will be achieved using a torch. | [Please refer to the Section / Page of the user manual that shows all options for activating the logger.]
4.6.3 | De-activation
| It should not be possible to de-activate the product inadvertently. | [Please refer to the Section / Page of the user manual that shows how the logger is de-activated and how it is protected from inadvertent de-activation.]
4.7 | Materials
4.7.1 | Ozone depleting chemicals
| During manufacture and assembly of the printed circuit boards and final assembly of the product do not use any substance included in Annex A, B or C of the Montreal Protocol. | [Please submit evidence for compliance.]
4.7.2 | Other restricted materials
| The product and its constituent components, including batteries, must not contain lead, mercury, cadmium, hexavalent chromium, polybrominated biphenyls (PBB) or polybrominated biphenyl ethers (PBDE). | [Please submit evidence for compliance.]
4.8 | Warranty
| The product is to be covered by the warranty offered by the manufacturer of the appliance into which the device is integrated. | [Please refer to the warranty document.]
4.9 | Servicing provision
| The product is to be maintenance-free apart from routine battery replacement. | [Please confirm compliance with this PQS specification requirement.]
4.10 | Disposal and recycling
| The manufacturer is to provide information to the buyer on the hazardous materials contained within the system and suggestions for resource recovery/recycling and/or environmentally safe disposal. For the European Union WEEE compliance in accordance with European Union Directive 2002/96/EC is mandatory. | [Please show the information provided to the buyer on hazardous materials and for recycling and safe disposal.]
4.11 | Instructions
| Instructions: User instructions are to be incorporated into the instruction manual issued by the manufacturer of the appliance into which the device is integrated. | [Please enter the reference for the user manual and website link for download.]
4.12 | Training
| Training: No requirement. | [There is no PQS requirement but please state whether training could be available e.g. via video tutorial etc.]
4.13 | Verification
| Verification: In accordance with PQS Verification Protocol E006/TH06-VP.2. | [Please enter the reference(s) for the laboratory test report(s).]
5. | Packaging
| Materials used for packaging the finished product are to be free of ozone-depleting compounds as defined in the Montreal Protocol. | [Please state the materials used for packaging and recommendations for environmentally safe disposal.]