## Sheet1
E006 Programmable remote temperature and event monitoring systems submission template (V1.2 24 FEB 2025)
| [Company name] | [Model name] | [Date]
| Dossier submitted by: | [Please enter your name, position within the company, and email address]
| Tested in accordance with PQS verification protocol | E006/TR03-VP.2
| To verify compliance with PQS specification below | E006/TR03.2
ASSESSOR’S SUMMARY:
Instructions in [square brackets] are for your guidance. Please delete these instructions and replace with evidence or a precise reference (including page and section numbers) showing compliance with that particular PQS specification requirement. In general a "pass" or "complies" statement alone is not acceptable.
One document may provide evidence for several clause requirements. Please carefully reference the sub-sections of any such document so that it is clear that the clause requirement has been met. This will help speed the pre-qualification process.
Greyed-out cells do not require a response.
NOTE: If it is not clear or doubtful whether a requirement has been met, then this is a non-compliance and pre-qualification might be delayed.
|  |  | REVIEWER’S SUMMARY: |  | REVIEWER’S SUMMARY: |  | REVIEWER’S SUMMARY:
Section 6. | Product dossier | Applicant's submission | Reviewer's remarks [Name & Date] | Applicants response [Name & Date] | Reviewer's remarks [Name & Date] | Applicants response [Name & Date] | Reviewer's remarks [Name & Date]
The legal manufacturer or reseller is to provide WHO with a pre-qualification dossier containing the following:
| Dossier examination fee in US dollars. | [Please contact PQS to arrange payment.]
| Type examination document (please refer to product verification protocol: E006/TR03-VP.2) that includes a detailed list of internationally recognized standards, certificates (not limited to the list in Section 2), and other evidence of conformity (e.g., screenshots of web portal to illustrate the required capabilities, data from tests of robustness and accuracy), to the specifications. This document must be verified and signed off by WHO PQS or an independent party authorized by WHO PQS. Section 5 in E006/TR03- VP.2 details the content of the type examination document. | [Please enter this information, include for all parties, e.g. manufacturer, reseller, plus other.]
| Copies of all certificates and other evidence of conformity listed in the document mentioned above. | [The dossier must have clear traceable references to all information required below.]
| Proof of adherence to an internationally recognized quality management system (e.g., relevant parts of the ISO 9001 series). | [Please enter the reference for the supporting document.]
| A minimum of three recent customer references for the Reseller or Legal Manufacturer involving order quantities greater than 100 units. The references need not be about the product being submitted for PQS approval. Rather, the references are to address and attest to the proficiency of customer service of the Reseller or Legal Manufacturer. | [Please enter the references or document reference in the dossier.]
| List of countries where the Legal Manufacturer has a service network capable of installing and maintaining the offered system. | [Please enter the reference for the supporting document.]
| Certified photocopy of Certificate of Traceability and Calibration traceable to an ISO/IEC 17025 accredited testing laboratory, to NIST, or to another internationally recognized standards agency for all system components intended for temperature measurement. | [Please enter the reference for the supporting document.]
| Where relevant, certified photocopies of the legal manufacturer’s ISO 14001 certification, EMAS registration or registration with an equivalent environmental audit scheme. Conformity with an environmental audit scheme is not mandatory; however preference will be given to manufacturers who are able to demonstrate compliance with good environmental practice. | [Please enter the reference for the supporting document.]
| Indicative cost of each system component EXW (Incoterms 2000). | [Please state or enter the reference for a supporting document.]
| In addition to the dossier, the Reseller or Legal Manufacturer is required to provide PQS with a sample unit of the system, along with detailed instructions on its setup and operation. The sample unit must comprise a minimum of: | [Please send 1 sample to each of the addresses to be provided.]
| • Two temperature sensors • One “door open” sensor (where offered) • One voltage sensor (where offered) • Two logging units (where part of system) • One base station • Web portal access information • Installation, commissioning, and user instructions for the system in English language. | [Please state components are provided.]
| Where appropriate, the Reseller or Legal Manufacturer may demonstrate the setup and operation of the sample unit in person to WHO PQS or to an independent party authorized by WHO PQS. | [Please state whether a demonstration of the logger would be possible.]
Section 4. | Requirements
4.1 | General
Hardware
4.1.1 | Sensors
| 1. Temperature sensor operating temperature range Upper limit: +50°C. Lower limit: -30°C. | [Please enter the reference for supporting evidence in the dossier.]
| 2. Temperature sensor accuracy. ±0.5°C or better, within the range -30°C to +20°C. | [Please enter the reference for supporting evidence.]
| 3. Temperature sensor resolution. ±0.2°C or better, within the range -30°C to +20°C. | [Please enter the reference for supporting evidence in the dossier.] Note that resolution is the smallest difference that can be measured, not necessarily the smallest increment on the display.
| 4. Temperature sensor response time. T90 within 20 minutes maximum in accordance with EN12830:1999. | [Please enter the reference for supporting evidence.]
| 5. Temperature sensor type. Electronic. | [Please state.]
| 6. “Door-open” event sensor operating parameters. Door open or door closed with a user-programmable delay function. A “door-open” event must be identified whenever the door panel is not fully seated in the closed position. | [Please state whether the system has this sensor.]
| 7. Voltage sensor operating parameters. To monitor all national standard combinations of single phase or three phase AC voltage and frequency, including the range of fluctuations encountered in the country where the equipment is installed. | [Please state whether the system has this sensor.]
| 8. Sensor lead length. Supplied as required to suit site conditions (does not apply for wireless sensors). | [Please state the lead length or whether the sensor is wireless.]
| 9. Sensor fixings. Remote sensors are to be supplied with fixings suitable for permanent attachment to the inside skin of a sectional cold room or freezer room or to the inside skin of a vaccine refrigerator or freezer. | [Please provide photo(s) of the fixing(s).]
| 10. IP rating for all sensors and connection leads. Protection of the product not less than IEC 60529: IP64. | [Please provide supporting test report evidence for IP rating.]
4.1.2 | Logging units
| 1. Logging unit operating parameters. Where logging units form part of the system they may incorporate integrated temperature sensors or they may be connected to external sensors. | [Please state whether integrated temperature sensors or external sensors.]
| 2. Logging unit power source - EITHER: 110/240 volt 50/60 Hz mains operated with replaceable, rechargeable battery backup with the following minimum charge capacities and minimum 3 years operational life: Segment 1 / 2 days Segment 2 / 5 days Segment 3 / 7 days | [Please state which option and reference to a test report for battery lifetime.]
| OR: non-replaceable, rechargeable battery with a minimum seven year operational life, with minimum three year warranty on battery, and same charge capacities as above. | [Please state which option and reference to a test report for battery lifetime.]
| OR: non-replaceable, non-rechargeable battery with a minimum seven year operational life, with minimum three year warranty on battery. | [Please state which option and reference to a test report for battery lifetime.]
| 3. Logging unit memory. EPROM or equivalent non-volatile solid-state memory device capable of storing a minimum of 30 days worth of data for each sensor, assuming a 10-minute logging interval. | [Please state type of memory and the total number of data capture events that can be stored. (A minimum of 4320 events per sensor are required.)]
| 4. Logging interval. The system administrator must be able to program the logging interval at discrete periods, starting with a minimum interval of ten minutes. | [Please state the range of logging intervals available and refer to the precise section / page in the user manual which explains how to set this interval.]
| 5. Logging unit casing. Non-corrodible and robust material (e.g., plastics or treated metal). | [Please state the material.]
| 6. Logging unit fixings. Logging units are to be supplied with fixings suitable for permanent attachment to the inside or outside skin of a sectional cold room or freezer room or to the inside or outside skin of a vaccine refrigerator or freezer. | [Please provide photo(s) of the fixing(s).]
| 7. Logging unit IP rating. Protection of the product not less than IEC 60529: IP64. | [Please provide supporting test evidence for IP rating.]
| 8. Connection to base station. Data transmission to the base station must be established automatically, without requiring any manual adjustment or fine-tuning other than plugging in sensors and cables, as necessary. If radio frequency links are used, they must be in accordance with ETSI EN 300- 220 or the ZigBee standard. In the event of interruption of power supply, the logging unit must be able to automatically reconnect with the base station once power supply is resumed. | [Please state how your logger is compliant with this requirement.]
4.1.3 | Base station
| 1. Base station channels. The minimum number of temperature sensors to be supported by each base station (including the sensors from the logging units) must be at least as follows: Segment 1 / 8 sensors Segment 2 / 4 sensors Segment 3 / 1 sensor For Segment 1, in addition to the temperature sensors, base stations for remote temperature monitoring systems must support at least 2 door sensors. | [Please state Segment and number of sensors.]
| A. Base station memory. Active base stations must incorporate EPROM or equivalent non-volatile solid-state memory capable of storing data from a minimum number of channels as specified above. The memory capacity must be sufficient to store 30 days worth of sensor data from all connected channels, assuming a 10-minute logging interval. The system is to be configured to overwrite data cyclically on a first-in-first-out basis when the memory is full. - In case of trouble with connectivity (and therefore with transmitting data to the server), the data must be stored in the base station. Upon resumption of connectivity, all the data stored during this period must be transmitted to the server. | [Please state type of memory and the total number of data capture events that can be stored. (A minimum of 4320 events per sensor are required.)]
| B. Base station power source. - EITHER: 110/240 volt 50/60 Hz mains operated with replaceable, rechargeable battery backup with the following minimum charge capacities and minimum 3 years operational life: Segment 1 / 2 days Segment 2 / 5 days Segment 3 / 7 days | [Please state which option and reference to a test report for battery lifetime.]
| OR: non-replaceable, rechargeable battery with a minimum seven year operational life, with minimum three year warranty on battery, and same charge capacities as above. | [Please state which option and reference to a test report for battery lifetime.]
| OR: non-replaceable, non-rechargeable battery with a minimum three year operational life. | [Please state which option and reference to a test report for battery lifetime.]
| C. Base station data connection capability. The product must include at least a USB connection capability for downloading data to a PC. Other connection capabilities are permitted, but USB is required. | [Please state which connection options are available.]
| D. Base station IP rating. Protection of the product not less than IEC 60529: IP50. | [Please provide supporting test evidence for IP rating.]
4.1.4 | Audio-visual alarm and response/mute button
| 1. Alarm and mute button location. The alarm and mute button must be at the same site as the remote temperature monitoring system, and must be capable of being positioned in an easily reachable location for the user. | [Please reference the precise section / page number in the user manual which shows how to mute the alarm.]
| 2. Sound intensity for audio alarms. 70dB(A) at a distance of one metre from the sounder. The pattern of the signal is to be an intermittent pulse. The timing and/or pattern of the pulse should be set to ensure that it cannot be confused with the fire alarm sounder standard applicable in the country of installation. Devices with an adjustable sound profile will be acceptable provided the means for adjustment is not accessible once the device is mounted in its final position. | [Please state type of alarm and sound intensity. If adjustable, refer to the precise section / page number in the user manual which explains how to adjust the type of alarm.]
| 3. Light intensity for visual alarms. Flashing red coloured light rated 100 mcd diffused. | [Please state whether the alarm is visual, colour and intensity.]
| 4. Alarm mode of operation. The audio-visual alarm is to be triggered whenever an alarm event occurs for a preset time period. It is to be cancelled by: i) physically pressing a “mute” button on the base station, ii) sending a remote response (e.g. through SMS) to the base station, or iii) through the web portal. Each of these responses is to be recorded by the system and used to calculate the “time of user response”, which is to be transmitted to the server and displayed in the web portal. | [Please reference the precise section / page number in the user manual that explains how to comply with all the options in this requirement.]
4.1.5 | Modem
| 1. Modem power source. The modem must utilize the same power source as the base station. | [Please state the modem power source.]
| 2. Modem functionality - Transmit data periodically from the base station to the server - Receive any remote updates, setting changes, or commands - Issue immediate alerts when alarms are triggered. e.g. by sending SMSes to pre-programmed mobile phone numbers or audio calling and leaving a recorded voice message. These alerts may be issued either directly by the modem or indirectly through the server. - Escalate the alert to higher-level decision-makers (e.g. through SMSes, audio calls, or emails) if the first alert is not responded to within a specified time limit. | [Please reference the precise section / page number in the user manual that explains how to comply with all parts in this requirement.]
| 3. Modem connectivity. Due to likely issues with connectivity (e.g. cellular connectivity issues), remote alerts and data transmission must automatically be possible through basic GSM standards. When available, higher-speed and more cost-efficient transmission (e.g. 3G) must be automatically utilized. | [Please state the type of connectivity.]
4.1.6 | Display
| A display should be part of the base station, and must provide the user with at least the following information: • That the device is activated • Battery charge level indicator OR battery replacement warning indicator activated when no less than 20% of battery life remains • The most recently read temperature for each sensor connected directly or indirectly to the base station • Whether or not the component is in an alarm condition • Date, time, and cellular signal strength are optional information that could be included. The display may show all this information together or the user may be required to access the information by means of a button mounted on the device. Logging units with more than one attached sensor must be able to show temperature and alarm data from all connected sensors. | [Please reference the precise section / page number in the user manual that explains how to comply with all parts in this requirement.]
4.1.7 | Calibration
| System components are to be covered by a Certificate of Traceability and Calibration. Producer-issued calibration certificates are acceptable along with a general NIST-traceable certificate for the series of logging units/base stations. However, in the case of producer-issued certification, the testing laboratory must be accredited with ISO/IEC 17025. The traceability declaration is to confirm that the measurement standards and instruments used during calibration of the components are traceable to an ISO/IEC 17025 accredited testing laboratory, to NIST, or to another internationally recognized standards agency. The certificate must be accompanied by a copy of the reference instrument calibration certificate. | [Please refer to supporting test evidence for calibration.]
4.1.8 | Power leads
| Each component connected to mains power is to be supplied with a factory-sealed plug connector to suit the mains electrical socket standard in the country of intended use (to be specified by the relevant UN-purchasing agency at the time of ordering). | [Please state compliance with this clause.]
4.1.9 | Electromagnetic compatibility
| Operation of the system and of the individual system components must be unaffected in the normal electromagnetic compatibility environment in which the system is intended to work, taking into account disturbance generated by adjacent apparatus which is compliant with relevant ISO, EN, or other internationally recognized standards. Information required to ensure uninterrupted use of the device must be contained in the user instructions. | [Please refer to supporting test evidence for electromagnetic compatibility.]
Software
4.1.10 | Front-end software and back-end software
| 1. “Plug-and-play” functionality. The base station must automatically connect and be recognized by the server when switched on. No user actions must be required to enable this, beyond the physical setup and switching on. In the event of interruption of power supply and/or cellular connectivity, the base station must automatically reconnect and be recognized by the server once power supply and/or cellular connectivity is restored. Upon reconnection, the base station must be able to transmit to the server the data that was collected during the interruption. | [Please state how product is compliant to the requirements of this clause.]
| 2. Remote configuration. System administrators must be able to remotely control and change settings through the web portal, including: a. Alarm thresholds b. Measurement and transmission frequencies of the base station c. Alarm recipients and their details (phone numbers, email addresses, etc.) d. Escalation tree for alarm (see Section 4.1.10.4) e. Recipients of periodic emailed reports on the health of the cold chain, f. Format of report (see Section 4.1.11.4) This user-interface must be highly visual and requiring minimal learning and effort (e.g., few mouse “clicks”) to navigate. | [Please reference the precise section / page number in the user manual that explains how to comply with all parts (a - f) in this requirement.]
| 3. Alarm thresholds. The system must allow different alarm settings for various types of excursions, with settings to be made by the system administrator. For example, allow 10 consecutive hours of excursions above 8°C before alarming, allow 1 hour of excursions below -0.5°C before alarming, immediately alarm if excursion reaches 40°C. This type of setting must be possible through the system. | [Please reference the appropriate section in the user manual.]
| 4. Alarm escalation. The system must allow a minimum of 3 layers of alarm escalation: a. The alarm must be sent to the first recipients, as specified by the system administrator. b. If the first recipients of the alarm do not respond within a time selected by the system administrator, then the alarm must be sent to their supervisor, as specified by the system administrator. c. If this second recipient too does not respond within the required time, the alarm must be sent to their supervisor, as specified by the system administrator. | [Please reference the precise section / page number in the user manual that explains how to comply with all parts (a - c) in this requirement.]
| The system administrator must be able to remotely specify (through the web portal) recipients at each level of this escalation tree, their preferred form of alarm (e.g., SMS, voice call, email), and the allowed time of response before the next escalation. | [Please reference the appropriate section in the user manual.]
| 5. Action-oriented alarms. Alarm alerts (e.g., SMSes, voice calls, and emails) must be directly actionable. For example: “Refrigerator number X has been experiencing a freezing excursion for 5 hours.” Escalated alarms to higher-level officials must be similarly actionable. For example: “Refrigerator X in health facility Y has been experiencing a freezing excursion for 1 hour. Please call <phone number> to talk to the relevant healthcare worker at this site.” | [Please reference the appropriate section in the user manual.]
| 6. “Grouping” of sensors. All sensors in the same appliance must be able to be “grouped” together, so that the system is aware that the same appliance is being monitored. This will prevent simultaneous alarms from different sensors for the same appliance. It will also enable users to monitor the cold chain by appliance, not by sensor. | [Please reference the appropriate section in the user manual.]
| 7. Unit of measurement. Temperatures throughout the system must be recorded in degrees Centigrade. Calculation and display of mean kinetic temperature (MKT) must also be possible for each refrigerator/walk-in cold room. | [Please reference the appropriate section in the user manual, e.g. to show temperature and MKT.]
4.1.11 | Web portal
| 1. Multiple layers of access privileges. The web portal must allow the system administrator to set different access privileges and read/write privileges for various users (at least 3 levels of access privileges, customizable by the system administrator). For example, system administrators at the National level must have full privileges for the entire country, while system administrators at the State level must have full privileges, but restricted for their State. Some officials must have read/write privileges, while others must only have read privileges. | [Please reference the appropriate section in the user manual.]
| 2. Web page and map view for cold chain status. The web portal must have a page to provide the remote user with a quick current status of all the equipment in their area of focus. A real-time map view might be available to present this visually. All appliances in alarm state must be highlighted in an obvious manner. | [Please reference the appropriate section in the user manual.]
| 3. Web page for cold chain report. The web portal must have a page to provide the remote user with a summary report of the performance of the cold chain during a selected time period. The user must be able to select the required time period by day, week, month, and year. For the selected time period, the report must show the total time durations of heating and freezing excursions of each appliance. The sites with the greatest time durations of excursions must be displayed at the top of the list, so that they are immediately visible. | [Please reference the appropriate section in the user manual.]
| 4. Periodic emailed report. On a monthly and annual basis, a simple, visual, and action-oriented report must be emailed to relevant officials, as specified by the system administrators. The report should summarize the health of the cold chain being monitored. Suggested information to summarize should include (but not be limited to): a. Since the last report, the total number of heating excursions above 8 °C for longer than 10 consecutive hours b. Since the last report, the total number of freezing excursions below - 0.5 °C for longer than 1 hour c. Since the last report, a list of all the monitored cold chain equipment, in decreasing order of number of heating and freezing excursions (equipment with the most excursions on top, equipment with the least excursions at the bottom). | [Please reference the precise section / page number in the user manual that explains how to comply with all parts (a - c) in this requirement.]
| 5. Raw data and graphs. The user must be able to easily access raw data from the server, and to develop graphs of the data over selected time periods. | [Please reference the appropriate section in the user manual.]
| 6. Exporting files. The web portal must allow data to be exported in Excel/CSV, ASCII, and XML formats and images to be exported in JPEG and PDF formats. | [Please reference the appropriate section in the user manual.]
| 7. Security. Authentication and access control of user and privileges standard not less stringent than US 21 CFR Part 11. | [Please state.]
4.1.12 | Shareable application programming interface (API)
| API's are used in building software applications. In some cases, data from the remote temperature monitoring device may need to be shared with an existing in-country database (e.g., DHIS2, eVIN). In these cases, the API’s must be shareable with the Employer of the system. | [Please state whether this is possible and how to implement this.]
4.2 | Robustness to environmental conditions
4.2.1 | Robustness to ambient temperature range during transport and storage
| -30°C to +55°C with the system components inactivated. | [Please enter the reference for supporting evidence.]
4.2.2 | Robustness to ambient humidity range during transport, storage and use
| 0 to 95% RH. | [Please enter the reference for supporting evidence.]
4.2.3 | Robustness to voltage fluctuations
| The functionality of the system and of the individual system components must not be affected by voltage fluctuations due to issues including (but not limited to) problematic power grid and intense electrical storm activity. Accordingly, the Reseller or Legal Manufacturer must inquire about the quality of power supply and, if necessary, recommend the usage of suitable voltage stabilizers to the procuring agency. If requested by the procuring agency, the Reseller or Legal Manufacturer must be able to source and supply PQS-approved voltage stabilizers. Note: Current specifications* can be found at: http://apps.who.int/immunization_standards/vaccine_quality/pqs_catalogue/c atdocumentation.aspx?id_cat=36 . | [Please state how your RTMD can cope with a problematic grid.]
4.3 | Physical characteristics
4.3.1 | Component dimensions
| Not critical, as units are placed on exterior of cold chain equipment. | [Please state dimensions of all units.]
4.3.2 | Component weight
| Not critical, as units are placed on exterior of cold chain equipment. | [Please state weight of all units.]
4.4 | Interface requirements
4.4.1 | Software compatibility
| • If the software requires an interface with a proprietary spreadsheet program, the list of compatible programs must include all releases of Microsoft Excel currently supported by Microsoft. • The software must be compatible with all Microsoft PC operating systems currently supported by Microsoft in addition to any other operating system. | [Please state which releases of Excel and which MS operating systems to which the software are currently compatible.]
4.5 | Materials
4.5.1 | Ozone depleting chemicals
| During manufacture and assembly of the printed circuit boards and final assembly of the product do not use any substance included in Annex A, B or C of the Montreal Protocol. | [Please submit evidence for compliance.]
4.5.2 | Other restricted materials
| The product and its constituent components, including batteries, must not contain lead, mercury, cadmium, hexavalent chromium, polybrominated biphenyls (PBB) or polybrominated biphenyl ethers (PBDE). | [Please submit evidence for compliance.]
4.5.3 | Packaging
| Materials used for packaging the finished product are to be free of ozone-depleting compounds as defined in the Montreal Protocol. | [Please submit evidence for compliance.]
4.6 | Warranty
| All system components are to be covered by a three year warranty in the event of any component failure. On-site repair and maintenance is to be offered as an option to the purchasing party. All warranty rights are to pass from the Approved Installer to the Employer (typically the system will be installed by an Approved Installer as a component part of a cold room or freezer room installation.) after the system has been commissioned and has been formally accepted by the Employer. Where the Employer is a UN agency, the warranty rights are to pass to the host government. (Some installations will initially be purchased by one of the UN procurement agencies. In this situation, warranty rights must pass to the host government.) | [Please refer to your legally binding warranty document with terms.]
4.7 | Service provision
| The system is to be maintenance-free, apart from routine battery or sensor replacement and re-calibration. If possible, sensors that will not require replacement throughout the lifetime of the device are preferable. | [Please reference the precise section / page number in the user manual which shows how to replace the battery and re-calibrate the sensor(s).]
4.8 | Disposal and recycling
| The manufacturer is to provide information to the buyer on the hazardous materials contained within the system and suggestions for resource recovery/recycling and/or environmentally safe disposal. For the European Union WEEE compliance in accordance with European Union Directive 2002/96/EC is mandatory. | [Please show the information provided to the buyer on hazardous materials and for recycling and safe disposal.]
4.9 | Instructions
| Each unit of the hardware must include a separate user manual and technician’s installation manual in the language most appropriate to the installation site. An English version of all instruction and manuals are required to be supplied for visual inspection during application for PQS certification. | [Please enter the reference for the user manual. Note that it can be supplied on a USB stick.]
4.10 | Training
| Training in the use of the equipment will generally be given by the installer of the cold room/freezer room or refrigerator. The equipment supplier should provide the necessary training materials, and should also be able to provide in-person training to installers and system administrators. | [Please state what training and how it can be made available.]
4.11 | Verification
| Pre-qualification evaluation of sample systems will be carried out in accordance with PQS Verification Protocol E006/TR03-VP.2. Post-tender assessment and field commissioning of systems incorporating pre-qualified system components will also be carried out in accordance with PQS Verification Protocol E006/TR03-VP.2. | [Please enter the reference(s) for the laboratory test report(s).]