## Sheet1
E006 User-programmable temperature data logger submission template (V1.2 24 FEB 2025)
| [Company name] | [Model name] | [Date]
| Dossier submitted by: | [Please enter your name, position within the company, and email address]
| Tested in accordance with PQS verification protocol: | E06/TR05.VP.1
| To verify compliance with PQS specification below: | E06/TR05.1
ASSESSOR’S SUMMARY:
Instructions in [square brackets] are for your guidance. Please delete these instructions and replace with evidence or a precise reference (including page and section numbers) showing compliance with that particular PQS specification requirement. In general a "pass" or "complies" statement alone is not acceptable.
One document may provide evidence for several clause requirements. Please carefully reference the sub-sections of any such document so that it is clear that the clause requirement has been met. This will help speed the pre-qualification process.
Greyed-out cells do not require a response.
NOTE: If it is not clear or doubtful whether a requirement has been met, then this is a non-compliance and pre-qualification might be delayed.
|  |  | REVIEWER’S SUMMARY: |  | REVIEWER’S SUMMARY: |  | REVIEWER’S SUMMARY:
Section 7. | Product dossier | Applicant's submission | Reviewer's remarks [Name & Date] | Applicants response [Name & Date] | Reviewer's remarks [Name & Date] | Applicants response [Name & Date] | Reviewer's remarks [Name & Date]
The legal manufacturer or reseller is to provide WHO with a pre-qualification dossier containing the following:
| Dossier examination fee in US dollars. | [Please contact PQS to arrange payment.]
| General information about the legal manufacturer, including name and address. | [Please enter this information, include for all parties, e.g. manufacturer, reseller, plus other.]
| Factory name | [Please enter this information.]
| Brand name of the product. | [Please enter this information.]
| Full specifications of the product being offered, covering all the requirements set out in this document, including details of product marking and traceability. | [The dossier must have clear traceable references to all information required below.]
| Certified photocopy of Certificate of Traceability and Calibration traceable to an ISO/IEC 17025 accredited testing laboratory, to NIST, or to another internationally recognized standards agency. | [Please enter the reference for the supporting document.]
| Certified photocopies of all type-approvals obtained for the appliance and/or its components, including CE marking and the like. | [Please enter the reference for supporting documents. A simple statement of compliance is not acceptable.]
| Certified photocopies of the legal manufacturer’s ISO 9001 quality system certification. | [Please enter the reference for the supporting document.]
| Where relevant, certified photocopies of the legal manufacturer’s ISO 14001 certification, EMAS registration or registration with an equivalent environmental audit scheme. Conformity with an environmental audit scheme is not mandatory; however, preference will be given to manufacturers who are able to demonstrate compliance with good environmental practice. | [Please enter the reference for the supporting document.]
| Where available, laboratory test report(s) proving conformity with the PQS equipment performance specifications. | [Please enter the reference for supporting document(s) .]
| One sample of the product complete with data connection lead, external sensor(s) (where offered) and software. | [Please send 1 sample to each of the addresses to be provided.]
| Indicative cost of the appliance per unit, per 10 units and per 100 units, EXW (Incoterms 2010) including appliance and solar power system. | [Please state or refer to a document.]
Section 4. | Requirements
4.1 | General
| User-programmable electronic temperature data loggers to be used for study and commissioning purposes throughout the vaccine cold chain. Devices with replaceable and non-replaceable batteries, with or without external sensor leads are included. | [Please describe logger.]
4.2 | Performance:
4.2.1 | Operating temperature range
| Upper limit: +55°C. Lower limit: -30°C. The logger electronics must function over the entire operating temperature range, with or without optional probes. | [Please enter the reference for supporting evidence in the dossier.]
4.2.2 | Accuracy
| ±0.5°C or better within the range -20°C to +20°C for a minimum of 12 months following initial calibration or subsequent re-calibration. ±1.0°C or better after 12 months if re-calibration is not carried out. | [Please enter the reference for the calibration report in the dossier. See also Clause 4.2.9.]
4.2.3 | Resolution
| ±0.2°C or better within the range -30°C to +20°C. | [Please enter the reference for supporting evidence in the dossier.] Note that resolution is the smallest difference that can be measured, not necessarily the smallest increment on the display.
4.2.4 | Power source
| Replaceable or non-replaceable battery. | [Please state whether battery is replaceable or non-replaceable. Also see Clause 4.2.15.]
4.2.5 | Sensor
| Devices with internal and/or external sensors will be considered. (Devices with external sensors may be required for monitoring temperatures at more than one point at a time and for monitoring temperatures in vaccine freezer rooms and freezers where battery and/or LCD performance may be adversely affected.) • Internal sensor: Electronic. • External sensor: Electronic exposed junction device with lead(s) up to 10 metres long. Devices offering more than one external sensor connection will not be excluded. | [Please state whether integrated or remote and if remote, state how the distortion to the refrigerator door seal is to minimized.]
4.2.6 | Memory
| EPROM or equivalent non-volatile solid-state memory device capable of storing a minimum of 1920 temperature readings per sensor. Some applications will require devices with a larger memory. | [Please state type of memory storage and storage capacity.]
4.2.7 | Product response time
| Product response time: T90 20 minutes maximum in accordance with EN12830:1999 | [Please enter the reference for test evidence in the dossier.]
4.2.8 | Unit of measurement
| Unit of measurement: Temperatures must be recorded in degrees Centigrade. A Fahrenheit scale is not required however products with an optional Fahrenheit scale are acceptable. | [Please state.]
4.2.9 | Calibration
| Calibration: Each product is to be covered by a Certificate of Traceability and Calibration. The traceability declaration is to confirm that the measurement standards and instruments used during calibration of the product are traceable to an ISO/IEC 17025 accredited testing laboratory, to NIST, or to another internationally recognized standards agency. The certificate must be accompanied by a copy of the reference instrument calibration certificate. | [Please enter the reference for the calibration report in the dossier. See also Clause 4.2.2.]
4.2.10 | Logging interval
| Logging interval: The user must be able to program the logging interval at equal intervals between one minute and one hour. Products with a wider range of logging intervals will be acceptable. | [Please state exact Section / Page number in the user manual which explains how the user programs the logging interval.]
4.2.11 | Logging start delay
| Logging start delay: The user must be able to program a logging start delay between 0 seconds and 30 days. Products allowing a programmable delay greater than 30 days will be acceptable. | [Please state exact Section / Page number in the user manual which explains how the user programs the logging interval.]
4.2.12 | Alarm settings
| • User programmable upper and lower alarm temperatures and alarm time delays at any temperature in the operating temperature range. | [Please state exact Section / Page number in the user manual which explains how the user programs upper and lower alarm temperatures and alarm time delays.]
| • User programmable cumulative exposure facility for at least one upper temperature setting and one lower temperature setting. | [Please state exact Section / Page number in the user manual which explains how the user programs upper and lower cumulative exposure.]
4.2.13 | Casing
| Casing: Non-corrodible plastics or metal case. | [Please state material.]
4.2.14 | IP rating
| IP rating: Protection of the product not less than IEC 60529: IP54 for devices with replaceable batteries and IP65 for devices with non-replaceable batteries. | [Please enter the reference for supporting IP test report.]
4.2.15 | Battery
| • EITHER: Replaceable battery with a minimum operating life of one year. • OR: Non-replaceable battery with a minimum operating life of one year. | [Please enter the reference for supporting evidence for battery life.]
4.2.16 | Electromagnetic compatibility
| Operation of the device must be unaffected in the normal electromagnetic compatibility environment in which it is intended to work, taking into account disturbance generated by adjacent apparatus which is compliant with relevant ISO, EN, or other internationally recognized standards. Information required to ensure uninterrupted use of the device must be contained in the user instructions. | [Please enter the reference for supporting EMC test report.]
4.2.17 | Data connection leads
| The product must include an RS232 or USB connection lead for linking the device to a PC OR the link between the product and the lead may be via a product-specific reader, provided the reader is an integral part of the cable assembly. | [Please confirm which type of connection leads will be provided.]
4.2.18 | Software
| The data link must be supplied with software on CD or USB stick – refer also to clauses 4.5 and 4.11. Downloadable software is not acceptable because of poor or absent internet connections in many user settings. | [Please confirm that software will be provided on a USD stick (preferable to CD). This software will also need to be provided with the sample sent to the PQS assessor.]
| As a minimum, the software must be capable of allowing the user to program the product to: - set the time zone at initialization either manually or automatically - record user name and job description data - set a logging start delay - set a logging interval - set upper and lower single exposure and cumulative exposure temperature alarm and delay conditions - set the product to stop when the memory is full, or to overwrite data cyclically on a first-in-first-out basis - batch program more than one logger with the same set-up data. | [Please state exact Section / Page number in the user manual which explains how the user programs the device to: • set the time zone at initialization either manually or automatically • record user name and job description data • set a logging start delay • set a logging interval • set upper and lower single exposure and cumulative exposure temperature alarm and delay conditions • set the product to stop when the memory is full, or to overwrite data cyclically on a first-in-first-out basis • batch program more than one logger with the same set-up data.]
| And output downloaded data in the form of temperature graphs and data tables to: - show time periods in seconds, minutes, hours or days to suit the programmed logging interval - show temperatures in degrees centigrade - clearly show the programmed alarm settings - show the time zone set by the user - show the data logger serial number and/or user-set tracking number - show user description data - show statistics, including minimum and maximum temperatures, average temperatures, cumulative time out of range, etc. - the ability to show a composite view of data recorded on multiple data loggers is desirable but not mandatory. - print legibly in A4 format on a monochrome printer. | [Please state exact Section / Page number in the user manual which explains how the user programs the device to: • show time periods in seconds, minutes, hours or days to suit the programmed logging interval • show temperatures in degrees centigrade • clearly show the programmed alarm settings • show the time zone set by the user • show the data logger serial number and/or user-set tracking number • show user description data • show statistics, including minimum and maximum temperatures, average temperatures, cumulative time out of range, etc. • the ability to show a composite view of data recorded on multiple data loggers is desirable but not mandatory. • print legibly in A4 format on a monochrome printer.]
| • Save logger data to a PC. | [Please state exact Section / Page number in the user manual which explains how the user saves logger data to a PC.]
| • Multilingual software, in the languages specified in clause 4.11, is desirable, but not mandatory. | [Please state in which languages the software is available.]
4.3 | Environmental requirements
4.3.1 | Ambient temperature range during transport and storage
| -30°C to +55°C. | [Please enter the reference for supporting evidence.]
4.3.2 | Ambient humidity range during transport, storage and use
| 0 to 95% RH. | [Please enter the reference for supporting evidence.]
4.3.3 | Resistance to electrical storms
| The functionality of the device must not be affected by intense electrical storm activity. | [Please enter the reference for supporting evidence.]
4.3.4 | Impact resistance
| Product to withstand 5 drops from 1 metre onto a concrete floor, with battery in place, without physical damage or loss of calibration. | [Please enter the reference for the supporting test report.]
4.3.5 | Vibration
| Product to withstand 30 minutes on a programmable vibrating table without physical damage or loss of calibration. | [Please enter the reference for the supporting test report.]
4.4 | Physical characteristics
4.4.1 | Overall dimensions
| Not exceeding 100 x 75 x 50mm, excluding data connection lead and any external sensors. | [Please state the dimensions.]
4.4.2 | Weight
| Not critical, provided the product is fully portable. | [Please state the weight.]
4.5 | Interface requirements
4.5.1 | Software compatibility
| • If the software requires an interface with a proprietary spreadsheet program, the list of compatible programs must include all releases of Microsoft Excel currently supported by Microsoft. | [Please state all options for compatible spreadsheet programs.]
| • The software must be compatible with all Microsoft PC operating systems currently supported by Microsoft. | [Please state all options for compatible operating systems.]
4.6 | Human factors
4.6.1 | Activation
| The product may be activated, either by means of a button or switch mounted on the unit, or by software command. | [Please state exact Section / Page number in the user manual that explains how to activate the logger.]
4.6.2 | De-activation
| It should not be possible to de-activate the product inadvertently or deliberately at any time within the logging period, except by software command or by means of a pin-operated switch. | [Please state exact Section / Page number in the user manual that explains how to de-activate the logger and how to avoid inadvertent de-activation.]
4.6.3 | User interface
Products with LED indicator(s) only: The product is to have a minimum of one indicator light. The indicator or indicators must provide the user with the following information by means of combinations of steady or flashing lights:
| • that the product is activated | [Please show a photo or a diagram of this indicator light.]
| • that the battery is functioning | [Please show a photo or a diagram of this indicator light.]
| • whether the temperature of the load has remained within alarm limits since recording commenced OR | [Please show a photo or a diagram of this indicator light.]
| • whether the temperature of the load has strayed outside alarm limits since recording commenced. | [Please show a photo or a diagram of this indicator light.]
Products with LCD displays: The product is to have an LCD display, with or without LEDs, capable of showing the following information:
| • that the product is activated | [Please show a photo or a diagram of this display.]
| • that the battery is functioning | [Please show a photo or a diagram of this display.]
| • the most recently recorded load temperature | [Please show a photo or a diagram of this display.]
| • whether the temperature of the load has remained within alarm limits since recording commenced OR | [Please show a photo or a diagram of this display.]
| • whether the temperature of the load has strayed outside alarm limits since recording commenced. | [Please show a photo or a diagram of this display.]
| • The LCD may show all this information together on the display or the user may be required to access the information by means of a button mounted on the product. Products with more than one sensor must be able to show temperature and alarm data from all connected sensors. | [Please show a photo or a diagram of this display.]
4.6.4 | Mounting device
| Although this is not a mandatory requirement, a means to physically secure the device to the load is desirable. Acceptable options include, but are not confined to: • Fixing eye for tie. • Double-sided tape. • Mounting cleat. | [Please state the mounting options for the device.]
4.6.4 | Audible alarm
| Not required. | [Please state if the device has an audible alarm.]
4.7 | Materials
4.7.1 | Ozone depleting chemicals
| During manufacture and assembly of the printed circuit boards and final assembly of the product do not use any substance included in Annex A, B or C of the Montreal Protocol. | [Please submit evidence for compliance.]
4.7.2 | Other restricted materials
| The appliance and its constituent components, must not contain lead, mercury, cadmium, hexavalent chromium, polybrominated biphenyls (PBB) or polybrominated biphenyl ethers (PBDE). | [Please submit evidence for compliance.]
4.8 | Warranty
| The product is to be covered by a one year warranty in the event of any component failure not caused by mechanical damage. | [Please refer to your legally binding warranty document with terms.]
4.9 | Servicing provision
| Servicing provision: The product is to be maintenance-free, apart from routine battery replacement and re-calibration. | [Please confirm compliance with this PQS specification requirement.]
4.10 | Disposal and recycling
| The manufacturer is to provide information to the buyer on the hazardous materials contained within the system and suggestions for resource recovery/recycling and/or environmentally safe disposal. For the European Union WEEE compliance in accordance with European Union Directive 2002/96/EC is mandatory. | [Please show the information provided to the buyer on hazardous materials and for recycling and safe disposal.]
4.11 | Instructions
| User instructions, including software manual, in Arabic, English, French, Mandarin Chinese, Russian and Spanish. The manual may be in hard copy format or supplied with the software on CD. | [Please enter the reference for the user manual and confirm compliance to the language requirement and state all options for receiving copies of the user instructions (including on a USB stick, or website download).]
4.12 | Training
| No requirement. | [There is no PQS requirement but please state whether training could be available e.g. via video tutorial etc.]
4.13 | Verification
| Verification: In accordance with PQS Verification Protocol E06/TR05.VP.1. | [Please enter the reference(s) for the laboratory test report(s).]
5. | Packaging
| Materials used for packaging the finished product are to be free of ozone- depleting compounds as defined in the Montreal Protocol. | [Please describe the packaging and refer to the user manual section/page number on how this packaging may be disposed / recycled.]