## Sheet1
Application for PQS qualification of a E006 30 day electronic refrigerator temperature logger
| [Company name] | [Model name]
| Dossier submitted by: | [Please enter your name, position within the company, and email address]
| Tested in accordance with PQS verification protocol: | E006/TR06.VP.4 (12 MAY 2022)
| To verify compliance with PQS specification below: | E006/TR06.4 (10 JAN 2022)
SCOPE
This specification describes the performance requirements for a 30 day electronic refrigerator logger. This product is intended for use as the principle means for monitoring storage conditions in vaccine refrigerators in intermediate stores and health facilities. The device may also be used as a secondary back-up device in cold rooms.
Instructions in [square brackets] are for your guidance. Please delete these instructions and replace with evidence or a precise reference (including page and section numbers) showing compliance with that particular PQS specification requirement. In general a "pass" or "complies" statement alone is not acceptable.
One document may provide evidence for several clause requirements. Please carefully reference the sub-sections of any document so that it is clear that the clause requirement has been met. This will help speed the pre-qualification process.
Greyed-out cells do not require a response.
NOTE: If it is not clear or doubtful whether a requirement has been met, then this is a non-compliance and pre-qualification might be delayed.
|  |  | REVIEWER’S SUMMARY: |  | REVIEWER’S SUMMARY: |  | REVIEWER’S SUMMARY:
Section 7. | Product dossier | Applicant's submission | Reviewer's remarks [Name & Date] | Applicants response [Name & Date] | Reviewer's remarks [Name & Date] | Applicants response [Name & Date] | Reviewer's remarks [Name & Date]
The legal manufacturer or reseller is to provide WHO with a pre-qualification dossier containing the following:
| Dossier examination fee in US dollars. | [Please contact PQS to arrange payment.]
| General information about the legal manufacturer, including name and address. | [Please enter this information, include for all parties, e.g. manufacturer, reseller, plus other.]
| Brand name of the product. | [Please enter this information.]
| Full specifications of the product being offered, covering all the requirements set out in this document, including details of product marking and traceability. | [The dossier must have clear traceable references to all information required below.]
| Certified photocopy of Certificate of Traceability and Calibration traceable to an ISO/IEC 17025 accredited testing laboratory, to NIST, or to another internationally recognized standards agency. | [Please enter the reference for the supporting document in the dossier.]
| Certified photocopies of all type-approvals obtained for the appliance and/or its components, including CE marking and the like. | [Please enter the reference for supporting documents in the dossier. A simple statement of compliance is not acceptable.]
| Certified photocopies of the legal manufacturer’s ISO 9001 quality system certification. | [Please enter the reference for the supporting document in the dossier.]
| Where relevant, certified photocopies of the legal manufacturer’s ISO 14001 certification, EMAS registration or registration with an equivalent environmental audit scheme. Conformity with an environmental audit scheme is not mandatory; however, preference will be given to manufacturers who are able to demonstrate compliance with good environmental practice. | [Please enter the reference for the supporting document in the dossier.]
| Laboratory test report(s) proving conformity with the PQS equipment performance specifications. | [Please enter the reference for the supporting document in the dossier.]
| One sample of the product complete with data connection lead, external sensor(s) (where offered) and software. | [Please send 1 sample to each of the addresses to be provided.]
| Indicative cost of the appliance per unit, per 10 units and per 100 units, EXW (Incoterms 2010) | [Please enter the reference for a supporting document in the dossier.]
Section 4. | Requirements
4.1 | General
| Electronic refrigerator logger, with factory-programmed alarms and visual display for monitoring storage conditions in vaccine refrigerators over a 30 day period. The device may also be used as a secondary back-up in cold rooms or freezer rooms. Type 1: +2 to +8°C. Type 2: -25 to -15°C. Type 3: -80 to -70°C. | [Please state Type and how this complies.]
4.2 | Performance
4.2.1 | Logger unit operating temperature range at least
| Type 1: Internal sensor: Upper limit: +50°C. Lower limit: -20°C. Type 1: External sensor: Upper limit: +50°C. Lower limit: -20°C. Type 2: Internal sensor: Upper limit: +50°C. Lower limit: -40°C. Type 2: External sensor: Upper limit: +50°C. Lower limit: -20°C. Type 3: Internal sensor: Upper limit: +50°C. Lower limit: -95°C. Type 3: External sensor: Upper limit: +50°C. Lower limit: -20°C. | [Please enter the reference for supporting evidence in the dossier.]
4.2.2 | Sensor measurement range
| Type 1: Upper limit: +50°C. Lower limit: -20°C. Type 2: Upper limit: +30°C. Lower limit: -40°C. Type 3: Upper limit: +30°C. Lower limit: -95°C. | [Please enter the reference for supporting evidence in the dossier.]
4.2.3 | Accuracy
| Type 1: ±0.5°C or better within the range -20°C to +20°C for a minimum of 12 months following initial calibration or subsequent re-calibration. ±1.0°C or better after 12 months if re-calibration is not carried out. Type 2: ±1.0°C or better within the range -30°C to -5°C. Type 3: ±1.5°C or better within the range -95°C to -30°C. | [Please enter the reference for the calibration report in the dossier. See also Clause 4.2.9.]
4.2.4 | Resolution
| Type 1: ±0.2°C or better within the range -10°C to +20°C. Type 2: ±0.2°C or better within the range -35°C to -5°C. Type 3: ±0.2°C or better within the range -95°C to -30°C. | [Please enter the reference for supporting evidence in the dossier.] Note that resolution is the smallest difference that can be measured, not necessarily the smallest increment on the display.
4.2.5 | Power source
| Non-replaceable battery. | [Please enter the reference for the supporting evidence in the dossier. See also Clause 4.2.16.]
4.2.6 | Sensor
| Electronic. Internal sensor devices: The complete device is to be mounted inside the cabinet in a position which accurately measures the load temperature. External sensor devices: The temperature display unit is to be physically attached to the outside of the cabinet and the sensor lead is to pass through the door seal. The sensor head is to be physically attached to the inside of the refrigerator or freezer in a position which accurately measures the load temperature. | [Please state whether integrated or remote and if remote, state how the distortion to the refrigerator door seal is minimized. (See also Clause 4.2.18.)]
4.2.7 | Product response time
| T90 20 minutes maximum in accordance with EN12830:1999 | [Please enter the reference for test evidence in the dossier.]
4.2.8 | Unit of measurement
| Temperatures must be recorded in degrees Centigrade. | [Please state.]
4.2.9 | Calibration
| Calibration: Each product is to be covered by a Certificate of Traceability and Calibration. The traceability declaration is to confirm that the measurement standards and instruments used during calibration of the product are traceable to an ISO/IEC 17025 accredited testing laboratory, to NIST, or to another internationally recognized standards agency. | [Please enter the reference for the calibration report in the dossier. See also Clause 4.2.3.]
4.2.10 | Logging interval
| The device must measure the storage temperature at intervals not exceeding 10 minutes and log all relevant data at intervals not exceeding one hour. Products with a wider range of logging intervals will be acceptable. | [Please state the logging intervals for the logger.]
4.2.11 | Mode of operation
| The device is to measure and log the temperature inside the refrigerator or freezer cabinet periodically, storing and displaying data for the previous 30 day period. It must be possible to read the maximum and minimum logged temperatures for each day via a ‘history mode’ function. If, at any time during the 30 day cycle the temperature in the cabinet exceeds the high alarm setting or drops below the low alarm setting for the periods described in clause 4.2.13, the device is to display the relevant alarm condition(s). At the end of the 30 day cycle the device is to continue the temperature and alarm monitoring process by incrementally overwriting data older than 30 days. The ‘history mode’ function must be programmed to include a ‘pause period’ so that transient temperatures are not recorded when the device is temporarily removed from the refrigerated environment for reading purposes. The ‘pause period’ operates by suspending the processing the reading of alarm events, maximum/minimum temperatures and duration statistics for a suitable period of time after button press activity. The operation of the internal clock must not be interrupted when the pause function is activated. With the exception of adjusting the device time and date settings, it must not be possible for the end-user to delete or modify logged temperature data and pre-programmed device settings (such as alarm thresholds, logging intervals, etc.) | [Please state exact Section / Page number in the user manual for the following: How to read the maximum and minimum logged temperatures for each day via a ‘history mode’ function. How to set and display high alarm and low alarm for the periods described in clause 4.2.13, the device is to display the relevant alarm condition(s). How to set ‘pause period’ so that transient temperatures are not recorded when the device is temporarily removed from the refrigerated environment for reading purposes.]
4.2.12 | Delayed start option
| A delayed start option may be offered, which, after initial switch-on, allows the device to stabilize to the temperature of the refrigerated environment before it starts recording. The period of delay should not exceed 30 minutes when subjected to the following operating conditions: * Device temperature at time of introduction to refrigerated environment: +43°C. * Temperature of refrigerated environment: +5°C (Type 1), -20°C (Type 2), or -75°C (Type 3). * Air movement in refrigerator: Still air, no circulation fan. | [Please state exact Section / Page number in the user manual for setting the delay function.]
4.2.13 | Alarm
| External sensor devices: Products with an external temperature sensor are to include a high breach and low breach visual alarm. An audible alarm is optional. Internal sensor devices: Devices with an internal temperature sensor are to include a high breach and low breach visual alarm only. | [Please state which type of alarm (as applicable to the type of logger).]
Alarm settings
| Alarm settings: Upper and lower alarm settings must be factory programmed into the device as follows: Type 1: * Low alarm setting: Exposure to a single temperature event of -0.5°C or below for 60 minutes. * High alarm setting: Exposure to a single temperature event of +8°C or above for 10 hours. Type 2: * Low alarm setting: Exposure to a single temperature event of -25°C or below for 60 minutes. * High alarm setting: Exposure to a single temperature event of -15°C or above for 60 minutes. Type 3: * Low alarm setting: Exposure to a single temperature event of -90°C or below for 60 minutes. * High alarm setting: Exposure to a single temperature event of -60°C or above for 60 minutes. | [Please enter the reference for supporting evidence e.g. test verification of these alarm settings.]
4.2.14 | Casing
| Non-corrodible plastics or metal case. | [Please state material.]
4.2.15 | IP rating
| Protection of the product not less than IEC 60529: IP64 | [Please enter the reference for the supporting IP test report.]
4.2.16 | Battery
| Non-replaceable battery with a minimum operating life of two years after a maximum shelf life of one year. | [Please enter the reference for supporting evidence for battery life.]
4.2.17 | Electromagnetic compatibility
| Operation of the device must be unaffected in the normal electromagnetic compatibility environment in which it is intended to work, taking into account disturbance generated by adjacent apparatus which is compliant with relevant ISO, EN, ANSI, or other internationally recognized standards. Information required to ensure uninterrupted use of the device must be contained in the user instructions. | [Please enter the reference for the supporting EMC test report.]
4.2.18 | Sensor lead protection
| External sensor leads are to be of a size and shape which does not damage the door seal or cause an air leak. The lead is to be positively located and protected against mechanical damage where it passes through the seal. | [State how the distortion to the refrigerator door seal is minimized for a remote, sensor. (See also Clause 4.2.6.)]
4.2.19 | PC interface
| It must be possible to access all the data itemized in this specification without connection to a PC or other peripheral device. Additionally, it must be possible to extract all of the data itemized in this specification to a PC or other peripheral device via USB download. Other methods of data download may also be considered, but USB download must be available as standard. | [Please state how the data may be extracted.]
4.2.20 | Type identification
| Type 1, 2 and 3 devices are to be clearly identified in a manner which avoids the risk that the wrong type of device will be used for a specific monitoring application. Acceptable identification includes, but is not confined to, the following information: • Printed identification, including indication of the intended monitoring range as specified in section 4.1. | [Please state how the Type is identified on the device.]
4.3 | Environmental requirements
4.3.1 | Ambient temperature range during transport, storage and use
| -30°C to +55°C. | [Please enter the reference for supporting evidence.]
4.3.2 | Ambient humidity range during transport, storage and use
| 5 to 95% RH. | [Please enter the reference for supporting evidence.]
4.3.3 | Resistance to electrical storms
| The functionality of the device must not be affected by intense electrical storm activity. | [Please enter the reference for supporting evidence.]
4.3.4 | Impact resistance
| Product to withstand 5 drops from 1 meter onto a concrete floor, with battery in place, without physical damage or loss of calibration, when cooled to the following temperatures: Type 1: Internal: to 0°C. Type 1: External: to 0°C. Type 2: Internal: to -25°C. Type 2: External: to 0°C Type 3: Internal: to -80°C. Type 3: External: to 0°C | [Please enter the reference for the supporting test report.]
4.3.5 | Vibration
| Product to withstand 30 minutes on a programmable vibrating table without physical damage or loss of calibration. | [Please enter the reference for the supporting test report.]
4.4 | Physical characteristics
4.4.1 | Overall dimensions
| Dimensions not critical provided volume does not exceed 200 cubic centimetres. | [Please state dimensions.]
4.4.2 | Weight
| Not critical. | [Please state weight.]
4.5 | Interface requirements
4.5.1 | Software compatibility
| • If the software requires an interface with a proprietary spreadsheet program, the list of compatible programs must include all releases of Microsoft Excel currently supported by Microsoft. | [Please list all compatible programs of Microsoft Excel currently supported by Microsoft.]
| • The software must be compatible with all Microsoft PC operating systems currently supported by Microsoft. | [Please state all compatible Microsoft PC operating systems currently supported by Microsoft.]
4.6 | Human factors
4.6.1 | User interface
| • A temperature read-out with visual alarm display is to be integrated into the device. | [Please explain or refer to a precise page and section in the user manual which shows the visual display with temperature and alarm.]
| • The visual alarm is to show unambiguously both high and low temperature breaches and is to show when they occurred during the 30 day cycle. The duration of the exceedance of the preset temperature alarm limits should also be displayed in read-out function. In the case of no alarms, in read-out mode, the device should display both highest and lowest temperatures reached in each day. | [Please explain or refer to a precise page and section in the user manual which shows the visual alarm is to show both high and low temperature breaches and when they occurred during the 30 day cycle.]
| • The display is to be legible by a person with normal vision (with or without glasses) standing in front of the device, both in bright sunlight and in tungsten/fluorescent lighting at 100 lux on the working plane. | [Please explain or refer to a precise page and section in the user manual which shows that the display is legible by a person with normal vision (with or without glasses) standing in front of the device, both in bright sunlight and in tungsten/fluorescent lighting at 100 lux on the working plane.]
| • There is to be a ‘replace battery’ or ‘low battery’ warning display which is triggered when not more than 20% or less than 30 days of battery life remains. | [Please explain or refer to a precise page and section in the user manual which shows a ‘replace battery’ or ‘low battery’ warning display which is triggered when not more than 20% or less than 30 days of battery life remains.]
| • As a battery saving measure, the display may switch off automatically when not required, provided it can be re-activated by the user by means of a push button. | [Please explain or refer to a precise page and section in the user manual which shows how the logger display can be re-activated by the user by means of a push button.]
4.6.2 | Activation
| The product is to be activated by means of a switch. The device is to remain activated until the expiry or removal of the battery. | [Please show all options for activating the logger.]
4.6.3 | De-activation
| It should not be possible to de-activate the product inadvertently. | [Please show how the logger is de-activated and how it is protected from inadvertent de-activation.]
4.6.4 | Mounting device
| • External sensor devices are to have a means for attaching the sensor head to the inside of the cabinet and a means for attaching the device to the outside of the cabinet. The sensor head must be protected from damage and dislodgement in the course of routine use when vaccine is added or removed from the refrigerator. • Internal sensor devices are to be stored with the vaccine. They should be designed so that the device can be attached to a wire shelf in a front opening refrigerator or to a wire basket in a top-opening refrigerator. The method of attachment must allow the device to be easily read whilst still in place. | [Please show method of attachment.]
4.7 | Materials
4.7.1 | Ozone depleting chemicals
| During manufacture and assembly of the printed circuit boards and final assembly of the product do not use any substance included in Annex A, B or C of the Montreal Protocol. | [Please submit evidence for compliance.]
4.7.2 | Other restricted materials
| The product and its constituent components, including batteries, must not contain lead, mercury, cadmium, hexavalent chromium, polybrominated biphenyls (PBB) or polybrominated biphenyl ethers (PBDE). | [Please submit evidence for compliance.]
4.8 | Warranty
| The product (including external sensors) is to be covered by a one year replacement warranty in the event of any component failure not caused by mechanical damage. | [Please refer to your legally binding warranty document with terms.]
4.9 | Servicing provision
| The product is to be maintenance-free. | [Please confirm compliance with this PQS specification requirement.]
4.10 | Disposal and recycling
| The manufacturer is to provide information to the buyer on the hazardous materials contained within the system and suggestions for resource recovery/recycling and/or environmentally safe disposal. For the European Union WEEE compliance in accordance with European Union Directive 2002/96/EC is mandatory. | [Please show the information provided to the buyer on hazardous materials and for recycling and safe disposal.]
4.11 | Instructions
| Instructions: User instructions (including software manual if applicable), in Arabic, English, French, Mandarin Chinese, Russian and Spanish. The user instruction and software manual must be in hard copy format and downloadable via the supplier's website. | [Please enter the reference for the user manual and website link for download.]
4.12 | Training
| Training: No requirement. | [There is no PQS requirement but please state whether training could be available e.g. via video tutorial etc.]
4.13 | Verification
| Verification: In accordance with PQS Verification Protocol E006/TR06.VP.4. | [Please enter the reference(s) for the laboratory test report(s).]
5. | Packaging
| Materials used for packaging the finished product are to be free depleting compounds as defined in the Montreal Protocol. | [Please state the materials used for packaging and recommendations for environmentally safe disposal.]