## Sheet1
E006 Electronic shipping indicator submission template (V1.1 21 FEB 2025)
| Assessment of appliance: | [Brand & Model here] | For dossier review on [TBC]
| Tested in accordance with PQS verification protocol: | E006/TR07.4-VP.4 (26 JAN 2022)
| To verify compliance with PQS specification: | E006/TR07.4 (26 JAN 2022)
ASSESSOR’S SUMMARY:
Before beginning this submission, please check Clause 4.1 and indicate for which versions your application is applicable.
Instructions in [square brackets] are for your guidance. Please delete these instructions and replace with evidence or a precise reference (including page and section numbers) showing compliance with that particular PQS specification requirement. In general a "pass" or "complies" statement alone is not acceptable.
One document may provide evidence for several clause requirements. Please carefully reference the sub-sections of any such document so that it is clear that the clause requirement has been met. This will help speed the pre-qualification process.
Greyed-out cells do not require a response.
NOTE: If it is not clear or doubtful whether a requirement has been met, then this is a non-compliance and pre-qualification might be delayed.
|  |  | REVIEWER’S SUMMARY: |  | REVIEWER’S SUMMARY: |  | REVIEWER’S SUMMARY:
Section 7. | Product dossier | Applicant's submission | Reviewer's remarks [Name & Date] | Applicants response [Name & Date] | Reviewer's remarks [Name & Date] | Applicants response [Name & Date] | Reviewer's remarks [Name & Date]
The legal manufacturer or reseller is to provide WHO with a pre-qualification dossier containing the following:
| Dossier examination fee in US dollars. | [Please contact PQS.]
| General information about the legal manufacturer, including name and address. | [Please enter this information, include for all parties, e.g. manufacturer, reseller, plus other.]
| Unique brand identification reference for the product. | [Please enter this information.]
| Full specifications of the product being offered, covering all the requirements set out in this document, including details of product marking and traceability and documentary evidence of claimed battery life. | [The dossier must have clear traceable references to all information required below.]
| Certified photocopy of Certificate of Traceability and Calibration traceable to an ISO/IEC 17025 accredited testing laboratory, to NIST, or to another internationally recognized standards agency. | [Please enter the reference for the supporting document.]
| Certified photocopies of all type-approvals obtained for the appliance and/or its components, including CE marking and the like. | [Please enter the reference for supporting documents. A simple statement of compliance is not acceptable.]
| Certified photocopies of the legal manufacturer’s ISO 9001 quality system certification. | [Please enter the reference for the supporting document(s) both for the manufacturer and re-seller if applicable.]
| Where relevant, certified photocopies of the legal manufacturer’s ISO 14001 certification, EMAS registration or registration with an equivalent environmental audit scheme. Conformity with an environmental audit scheme is not mandatory; however, preference will be given to manufacturers who are able to demonstrate compliance with good environmental practice. | [Please enter the reference for the supporting document.]
| Laboratory test report(s) proving conformity with the PQS equipment performance specifications. | [Please enter the reference for supporting document(s).]
| Sample(s) of the product complete with data connection lead, external sensor(s) (where offered) and software. | [Please send 1 sample to each of the addresses to be provided.]
| Indicative cost of the appliance per unit, per 10 units and per 100 units, EXW (Incoterms 2010) | [Please state the price for quantity 1, 10, and 100 units.]
Section 4. | Requirements
4.1 | General
| Single use pre-programmed electronic time-temperature data logger with non-replaceable battery to accompany vaccine shipments from the vaccine manufacturer’s warehouse to the receiving country’s primary vaccine store. The logger must be able to display the shipment’s time-temperature exposure without need for downloading to a PC and without need for a separate reading device. Devices that have an additional download function will be acceptable, but a download function is unnecessary, will not routinely be used, and does not form part of this specification. | [Please state how your logger complies.]
The device should be supplied in five versions:
| • Type D: Programmed with alarm settings suitable for the international shipment of COVID-19 vaccine (Comirnaty) and Ebola vaccine (Ervebo). The devices could be probeless or with probe. | [Please state whether you are applying for this version.]
| • Type C: Programmed with alarm settings suitable for the international shipment of COVID-19 vaccine (Corbevax, AstraZeneca, Covishield, CoronaVac, and BIBP), DTP, DT, TT, Td, HepA, HepB, IPV, liquid Hib, influenza seasonal (liquid), JE (liquid), Meningococcal ACYW-135 (liquid), OCV, HPV, PCV (other than Prevenar), Rabies, RV (Liquid and other than Bharat liquid). | [Please state whether you are applying for this version.]
| • Type A/B: Programmed with alarm settings suitable for the international shipment of BCG, COVID-19 vaccine (Janssen, Moderna), Ebola vaccine (Mvabea and Zabdeno), Hib (lyophilized), influenza seasonal (lyophilized), JE lyophilized, measles, MR, MMR, meningococcal A, meningococcal ACYW 135 (lyophilized), OPV, rabies (lyophilized), rotavirus (Bharat liquid and lyophilized – other than RotaTeq), rubella, varicella, and yellow fever vaccines. | [Please state whether you are applying for this version.]
| • Type Rotateq: rotateq only | [Please state whether you are applying for this version.]
| • Type Prevenar: only for Prevenar 7 and 13 -valent conjugated vaccine. | [Please state whether you are applying for this version.]
| Type A/B, type C, type D devices with probes, type rotateq, and type Prevenar may be offered with a minimum 40-day recording period. Probeless type D may be offered with a minimum 20-day recording period. | [Please state which version (20 or 40 day) is being offered.]
| It must be possible to photocopy the logger display as a permanent record of the shipment’s arrival status. A legible copy must be produced using a photocopier, scanner or all-in-one printer. | [Please state whether the display can be copied or scan legibly.]
4.2 | Performance
4.2.1 | Operating temperature range
| • Type A/B, C, Rotateq and Prevnar: Upper limit: +55°C. Lower limit: -30°C. • Type D: Device: Probeless: Upper limit: +30°C. Lower limit: -95°C. • Type D: Device: With probe: Upper limit: +30°C. Lower limit: -30°C. • Type D: Probe: Upper limit: +30°C. Lower limit: -95°C. | [Please enter the reference for supporting evidence in the dossier.]
4.2.2 | Accuracy
| Temperature: • Type A/B, C, Rotateq and Prevnar: ±0.5°C or better within the range 5°C to +25°C; ±1°C within the ranges -20°C to -5°C and +25°C to +55°C. • Type D Probeless: ±1.5°C or better within the range -95°C to -30°C; ±1°C within the ranges -30°C to -5°C; ±0.5°C or better within the range -5°C to +30°C. Time: ± 10 seconds per day or better. Probe sensor accuracy range ±1.5°C or better within the range -90°C to +30°C; ±2°C within the ranges -95°C to 0°C. | [Please enter the reference for the calibration report in the dossier. See also Clause 4.2.9.]
4.2.3 | Resolution
| • ±0.2°C or better within the range -20°C to +55°C for A/B, C, Rotateq and Prevnar types. • ±0.5°C or better over full temperature measurement range for type D. | [Please enter the reference for supporting evidence in the dossier.] Note that resolution is the smallest difference that can be measured, not necessarily the smallest increment on the display.
4.2.4 | Power source
| Non-replaceable battery. | [Please state and also see Clause 4.2.15.]
4.2.5 | Sensor
| Electronic. The devices with sensor must have single sensor (external) and must be serialized as a single system (both the monitor and the probe). | [Please state.]
4.2.6 | Memory
| EPROM or equivalent non-volatile solid-state memory device. | [Please state.]
4.2.7 | Product response time
| For types A/B, C, type D devices with probe, Prevnar and Rotateq devices, T90 10 minutes maximum in accordance with EN12830. For probless type D devices, T90 20 minutes maximum in accordance with EN12830. | [Please enter the reference for supporting evidence in the dossier.]
4.2.8 | Unit of measurement
| Temperatures must be recorded in degrees Centigrade. | [Please state.]
4.2.9 | Calibration
| Each product is to be covered by a Certificate of Traceability and Calibration. The traceability declaration is to confirm that the measurement standards and instruments used during calibration of the product are traceable to an ISO/IEC 17025 accredited testing laboratory, to NIST, or to another internationally recognized standards agency. | [Please enter the reference for the calibration report in the dossier. See also Clause 4.2.2.]
4.2.10 | Logging interval
| The device must measure the storage temperature at intervals not exceeding 10 minutes. As a minimum the device must log the first instance of a time-temperature-violation for each alarm type equalling or exceeding the threshold parameters set out in clause 4.2.12. Devices that can log more than one instance of each type of time-temperature violation will not be excluded. | [Please state how exactly the logger complies.]
4.2.11 | Logging start delay
| 60 minute start delay function after user activation to allow the device to equilibrate with the temperature inside the shipping container before it starts to record temperatures. | [Please state exact Section / Page number in the user manual to set the logging start delay.]
4.2.12 | Alarm settings
The device must be pre-programmed with alarm settings in accordance with the Table on the Figures page of this spreadsheet (Page 6 of the PQS specification):
| • Type D: | [Please state whether you are applying for this version.]
| • Type C: | [Please state whether you are applying for this version.]
| • Type A/B: | [Please state whether you are applying for this version.]
| • Type Rotateq: | [Please state whether you are applying for this version.]
| • Type Prevenar: | [Please state whether you are applying for this version.]
4.2.13 | Casing
| Non-corrodible plastics or metal case. | [Please state the materials used.]
4.2.14 | IP rating
| Protection of the product not less than IEC 60529: IP64 | [Please enter the reference for supporting test report in the dossier.]
4.2.15 | Battery
| Non-replaceable battery capable of powering the device in accordance with the following criteria: • Minimum storage life of 18 months before ‘start’ for A/B, C, Rotateq and Prevnar types, and 12 months for Type D. • Minimum recording period: 40 days for A/B, C, D with probe, Rotateq and Prevnar types, and 20 days for Type D probeless. • Minimum data retention period after ‘stop’: 6 months. | [Please enter the reference for supporting evidence for the battery life.]
4.2.16 | Electromagnetic compatibility
| Operation of the device must be unaffected in the normal electromagnetic compatibility environment in which it is intended to work, taking into account disturbance generated by adjacent apparatus which is compliant with relevant ISO, EN, or other internationally recognized standards. Information required to ensure uninterrupted use of the device must be contained in the user instructions. The device and its power source must conform to all relevant requirements issued by the International Air Transport Authority (IATA) and must not interfere with aircraft electrical or electronic systems. | [Please enter the reference for a supporting test report for EMC compliance.]
4.3 | Environmental requirements
4.3.1 | Ambient temperature range during transport and storage
| -30°C to +55°C with device inactivated for A/B, C, D (probeless and with probe), Rotateq and Prevnar types. | [Please enter the reference for supporting evidence in the dossier.]
4.3.2 | Ambient humidity range during transport, storage and use
| 5 to 95% RH. | [Please enter the reference for supporting evidence in the dossier.]
4.3.3 | Resistance to electrical storms
| The functionality of the device must not be affected by intense electrical storm activity. | [Please enter the reference for supporting evidence in the dossier.]
4.3.4 | Impact resistance
| Product to withstand 5 drops from 1 metre onto a concrete floor, with battery in place, without physical damage or loss of calibration. | [Please enter the reference for supporting test report in the dossier.]
4.3.5 | Vibration
| Product to withstand 30 minutes on a programmable vibrating table without physical damage or loss of calibration. | [Please enter the reference for supporting test report in the dossier.]
4.4 | Physical characteristics
4.4.1 | Overall dimensions
| Not critical provided volume of the device does not exceed 150 cubic centimetres when detached from shipment information card. | [Please state dimensions.]
4.4.2 | Weight
| Not critical. | [Please state weight.]
4.5 | Interface requirements
4.5.1 | Software compatibility
(for devices with additional download function):
| • If the software requires an interface with a proprietary spreadsheet program, the list of compatible programs must include all releases of Microsoft Excel currently supported by Microsoft. | [Please list all compatible programs of Microsoft Excel (currently supported by Microsoft).]
| • The software must be compatible with all Microsoft PC operating systems currently supported by Microsoft. | [Please state all compatible Microsoft PC operating systems (currently supported by Microsoft).]
4.6 | Human factors
4.6.1 | Activation
| The device is to be activated by the sender at the beginning of the recording period by means of a ‘start’ button or switch mounted on the unit. | [Please state exact Section / Page number in the user manual that explains how to activate the logger.]
4.6.2 | De-activation
| The device is to be de-activated by the receiver at the end of the recording period by means of a ‘stop’ button or switch mounted on the unit. If the 'stop' button or switch is not de-activated, the device should automatically default to the de-activated state at the end of the 40 day recording period, as applicable. The 'stop' button or switch should be designed to prevent inadvertent de-activation - for example by contact with a shifting load. | [Please state exact Section / Page number in the user manual that explains how to de-activate the logger and how to avoid inadvertent de-activation.]
4.6.3 | User interface
The device is to have an LCD display screen, with or without LEDs, capable of showing the following information:
| • Activation status. | [Please state exact Section / Page number in the user manual that illustrates activation status.]
| • Post activation battery status, or clearly marked expiry date in the format mm/yyyy. | [Please state exact Section / Page number in the user manual that illustrates post-activation status.]
| • Overall alarm status: whether or not an alarm condition of any kind has occurred since the device was activated. | [Please state exact Section / Page number in the user manual that illustrates alarm status.]
| • Time-temperature alarm status: the status of each of the three time- temperature alarm thresholds specified in clause 4.2.12 at the time when the ‘stop’ button is activated. | [Please state exact Section / Page number in the user manual that illustrates time-temperature alarm status.]
| • Total elapsed transport time in days and hours or in hours measured from device activation to device de-activation. | [Please state exact Section / Page number in the user manual that illustrates total elapsed time status.]
| • Shipment history: A history of the shipment capable of showing details of at least one time-temperature limit violation for each alarm type including the first time-temperature-violation of each alarm type. | [Please state exact Section / Page number in the user manual that explains how to show the shipment history.]
| • The LCD must either show all this information together on a single display screen or the user must be able to access the information on sequential screens by means of a button mounted on the product. In the latter case, the overall status of the indicator (‘OK’, or ‘Alarm’) must be permanently displayed on every screen. Flashing displays are not acceptable because they cannot be photocopied. | [Please state exact Section / Page number in the user manual that explains how to show all the information demanded by this clause.]
| • The display must be capable of being photocopied or transferred wirelessly to a smartphone or computer in order to provide a hardcopy or softcopy record of the status of the device upon arrival. For devices in which data cannot be wirelessly transferred to a phone or computer, the display must not incorporate any flashing or blinking symbols or lights. | [Please state exact Section / Page number in the user manual that explains how this information can be transferred.]
| • Alarm symbols must not be language-dependent and must be easily understood by untrained users. Acceptable symbols include, but are not confined to, the following: Tick’ or ‘OK’ symbol for shipments where no temperature violation has occurred, as graphic in Clause 4.6.3 on Figures page: | [Please state exact Section / Page number in the user manual that illustrates the symbols used on this indicator.]
| ‘Cross’ or ‘Crossed OK’ symbol for shipments where any type of temperature violation has occurred, as graphic in Clause 4.6.3 on Figures page: | [Please state exact Section / Page number in the user manual that illustrates the symbols used on this indicator.]
| • As a battery saving measure, the display may switch off automatically when not required, provided it can be activated by the user by means of a push button. | [Please state whether the display switches off automatically and how it may be activated.]
4.6.4 | Type identification
| All devices are to be clearly identified in a manner which avoids the risk that the wrong type of device will be packed by the sender. Acceptable identification includes, but is not confined to, the following information: • Printed identification, including indication of minimum recording period (40 day for A/B, C, Rotateq and Prevnar types, and 20 day for probeless type D). • Different colored casings. | [Please state how each type of shipping indicator is clearly identified.]
4.6.5 | Shipment information card
| Mount the device on a moisture resistant backing card, using moisture resistant adhesive. The card material must accept indelible markings in ball point pen. The width of the card must be at least the same as the length of the device, subject to a minimum width of 7.5 cm. The length of the card must not exceed 14 cm. The card design must follow the generic format and colours set out in Annex 1 (yellow for Type C and Prevenar®, blue for Type A/B and Rotateq®, and green for Type D). User instructions are to be available either in English, French or Spanish language, as requested by the customer. Text is to be in a high legibility font – minimum 8 point, colour black. | [Please state: - card material - whether it can accept markings with a ball-point pen - dimensions - show a photo of each type including applicable user instructions - size of font.]
4.7 | Materials
4.7.1 | Ozone depleting chemicals
| During manufacture and assembly of the printed circuit boards and final assembly of the product do not use any substance included in Annex A, B or C of the Montreal Protocol. | [Please submit evidence for compliance.]
4.7.2 | Other restricted materials
| The product and its constituent components, including batteries, must not contain lead, mercury, cadmium, hexavalent chromium, polybrominated biphenyls (PBB) or polybrominated biphenyl ethers (PBDE). | [Please submit evidence for compliance.]
4.8 | Warranty
| The product is to be covered by a warranty covering the designed lifetime of the device in the event of any component failure not caused by mechanical damage. | [Please refer to your legally binding warranty document with terms.]
4.9 | Servicing provision
| The product is to be maintenance-free. | [Please confirm compliance with this PQS specification requirement.]
4.10 | Disposal and recycling
| The manufacturer is to provide information to the buyer on the hazardous materials contained within the system and suggestions for resource recovery/recycling and/or environmentally safe disposal. For the European Union WEEE compliance in accordance with European Union Directive 2002/96/EC is mandatory. | [Please show the information provided to the buyer on hazardous materials and for recycling and safe disposal.]
4.11 | Instructions
| Instructions: User instructions, including software manual if applicable, in Arabic, English, French, Mandarin Chinese, Russian and Spanish. The manual may be in hard copy format or supplied with the software on CD. | [Please enter the reference for the user manual and confirm compliance to the language requirement and state all options for receiving copies of the user instructions (including on a USB stick, or website download).]
4.12 | Training
| Training: No requirement. | [There is no PQS requirement but please state whether training could be available e.g. via video tutorial etc.]
4.13 | Verification
| Verification: In accordance with PQS Verification Protocol E006/TR07.VP.4. | [Please enter the reference(s) for the laboratory test report(s).]
5. | Packaging
| Materials used for packaging the finished product are to be free of ozone- depleting compounds as defined in the Montreal Protocol. | [Please describe the packaging and refer to the user manual section/page number on how this packaging may be disposed / recycled.]