## Spec
PQS Submission template for E007 Solar direct drive surplus energy harvest control V1.2 (23 July 2026)
| Assessment of appliance: | [Brand & Model here] | For dossier review on [TBC]
| Tested in accordance with PQS verification protocol | E007/EHC01-VP.1 (01 NOV 2016)
| To verify compliance with PQS specification below | E007/EHC01.1 (02 MAY 2017)
ASSESSOR’S SUMMARY:
Instructions in [square brackets] are for your guidance. Please delete these instructions and replace with evidence or a precise reference (including page and section numbers) showing compliance with that particular PQS specification requirement. In general a "pass" or "complies" statement alone is not acceptable.
One document may provide evidence for several clause requirements. Please carefully reference the sub-sections of any such document so that it is clear that the clause requirement has been met. This will help speed the pre-qualification process.
Greyed-out cells do not require a response.
NOTE: If it is not clear or doubtful whether a requirement has been met, then this is a non-compliance and pre-qualification might be delayed.
|  |  | REVIEWER’S SUMMARY: |  | REVIEWER’S SUMMARY: |  | REVIEWER’S SUMMARY:
Section 7: | Product dossier | Applicant's submission | Reviewer's remarks [Name & Date] | Applicants response [Name & Date] | Reviewer's remarks [Name & Date] | Applicants response [Name & Date] | Reviewer's remarks [Name & Date]
The legal manufacturer or reseller is to provide WHO with a pre-qualification dossier containing the following:
| Dossier examination fee in US dollars. | [Please contact PQS.]
| General information about the legal manufacturer, including name and address. | [Please state or refer to a document.]
| Unique identification reference for the product type. | [Please state.]
| Brand name of the EHC. | [Please state.]
| Full specifications of the EHC being offered, covering all the requirements set out in this document, including details of EHC marking and traceability. | [Please refer to a document giving complete details.]
| Full details of the compatible appliance(s) and certification of appliance conformance with specification E003/FZ03 or E003/RF05. | [Please refer to a document giving complete details.]
| Full details of the compatible Type 2 solar power system and certification of conformance with specification E003/PV01. | [Please refer to a document giving complete details.]
| A comprehensive set of photographs including a three quarter view of the EHC. Additional photographs showing all external surfaces of the EHC, the interior, defined loads, connection points/cables and any special features. | [Please refer to the photos.]
| Certified photocopies of all type-approvals obtained for the EHC and/or its components, including CE marking and the like. | [Please refer to a document.]
| Certified photocopies of the legal manufacturer’s ISO 9001 quality system certification. | [Please refer to a document.]
| Where relevant, certified photocopies of the legal manufacturer’s ISO 14001 certification, EMAS registration or registration with an equivalent environmental audit scheme. Conformity with an environmental audit scheme is not mandatory; however, preference will be given to manufacturers who are able to demonstrate compliance with good environmental practice. | [Please refer to a document.]
| Laboratory test report(s) proving conformity with the product specifications. | [Please refer to a document.]
| Indicative cost of the product per 10 000 units, per 100 000 units and per 1 000 000 units EXW (Incoterms 2015). | [Please state or refer to a document.]
Section 4. | Requirements
4.1 | General
| In all operational modes the EHC prioritizes appliance operation over other loads. | [Please explain how your device prioritizes appliance operation over other loads.]
4.2 | Fail safe operation
| The EHC must be designed so that if for any reason it fails, safe and adequate power is still supplied to the appliance for continuous operation. The EHC must include a simple and intuitive mechanism to be disconnected or bypassed completely to allow the appliance to continue to operate correctly in the event of EHC failure. The EHC must survive common control fault conditions. Such conditions are as follows but not limited to: connection errors including reversed polarity, voltage mismatch and wiring connections in any sequence possible, high temperature operation, power input over-current conditions and short circuit (appliance, EHC, load and solar power system). In the event of EHC failure, the appliance will continue to operate correctly as prequalified and remain safely connected to the solar power system. | [Please explain how your device complies with each part of this clause.]
4.3 | Energy harvest
| The EHC must manage surplus solar electricity that can be controlled, diverted to and consumed by other loads. | [Please explain how your device manages surplus solar electricity.]
| An EHC must be specified by manufacturers for either undefined loads selected by the user or for defined loads provided by the manufacturer. In both cases the EHC and appliance will be tested only with a laboratory supplied variable direct current (DC) load designed to test the limits of the system. | [Please define the loads or range of loads that can be powered by your EHC system.]
| The energy harvest (in watt-hours per day on average) and the energy harvest efficiency have no set criteria but will be measured and reported. Manufacturers are to state the minimum daily quantity of harvested energy (Wh/day) available based on the Day/Night test at the minimum solar radiation reference period. | [Please state the energy consumption (in watt-hours per day on average) and the energy harvest efficiency and reference the test(s) where this is confirmed.] [Please state the minimum daily quantity of harvested energy (Wh/day) available based on the Day/Night test at the minimum solar radiation reference period).]
| Manufacturer to provide clear examples of energy harvest available for users (e.g., X quantity of Y device in operation for Z hours per day). | [Please provide clear examples of energy harvest available for users (e.g., X quantity of Y device in operation for Z hours per day).]
4.4 | Compatibility and compliance with relevant references
| The EHC must be compatible with at least one WHO PQS prequalified appliance and the appliance’s required Type 2 solar power system per E003/PV01. When possible it is recommended that load batteries comply with E003/PV01. | [Please state at least one compatible WHO PQS prequalified appliance and the appliance’s required Type 2 solar power system (per E003/PV01).]
4.5 | Appliance operation
| The SDD appliance power switch (as required in PQS E003 RF05.4, section 4.5.2) must simultaneously disconnect both the EHC and the SDD appliance to prevent accidental or intentional power diversion to secondary load(s) through the EHC when the SDD appliance is switched off. | [Please refer to the test where this is demonstrated.]
| Manufacturer to certify that EHC design prevents simultaneous initial starting of appliance and load. Manufacturer to certify that EHC design prevents loads with a battery from powering the appliance under any operating condition. | [Please certify that EHC design prevents simultaneous initial starting of appliance and load.] [Please certify that EHC design prevents loads with a battery from powering the appliance under any operating condition.]
| Certifications to include a description of the method(s) used to prevent simultaneous initial starting of appliance and loads and methods to prevent load battery from powering the appliance under any operating condition. | [Please state how your device prevents simultaneous initial starting of appliance and loads and methods to prevent load battery from powering the appliance under any operating condition.]
4.6 | Appliance compatibility
| The EHC must be compatible with at least one WHO PQS prequalified appliance. | [Please state at least one compatible WHO PQS prequalified appliance.].
4.7 | Solar power system compatibility
| The EHC must be compatible with the specific solar power system(s) required by the specific appliance(s) with which the EHC is prequalified. All EHCs are operated on direct current (DC) electricity generated by the solar power system. | [Please give example(s) of specific solar power system(s) as required by specific appliance(s) with which the EHC is prequalified.]
| The energy harvest system cannot result in a solar power system larger than that required by the SDD appliance alone. | [Please confirm energy harvest system does not effect the size of the solar array.].
4.8 | Load compatibility
| Legal Manufacturer/Reseller are to clearly indicate if the EHC is designed for use with a defined load or an undefined load. Legal Manufacturer/Reseller are to clearly note any incompatible load types including external battery charge regulator type (e.g. maximum power point tracking, pulse width modulated, series, shunts). It is recommended that all EHCs with load outputs of 12 V DC or higher be compatible with all common solar battery charge regulator types. | [Please state if the EHC is designed for use with a defined load or an undefined load and show where this is indicated.] [Please state any incompatible load types and show where this is indicated.]
4.9 | Available energy harvest and energy efficiency
| Manufacturer to state the minimum average daily watt hours (Wh/day) of energy available to users based on the day/night test at the minimum solar radiation reference period and provide clear examples for users (e.g., X quantity of Y device in operation for Z hours per day). | [Please state the minimum average daily watt hours (Wh/day) of energy available to users (based on the day/night) and refer to the test where this is demonstrated.]
| An EHC with a defined load will report only the energy used for the defined load. Any additional energy will not be reported. Results will be published for the minimum available Wh/day and overall energy efficiency (estimated as the available Wh/day divided by the possible average surplus solar electricity as measured in the day/night appliance testing). | [Please state the energy used for the defined load and refer to test evidence.] [Please state the minimum available Wh/day and overall energy efficiency and the test evidence for this.]
4.10 | Electrical safety rating
| Legal Manufacturer/Reseller to certify compliance to IEC 60335-1, IEC 60335-2- 24 and IEC 60364-1. | [Please refer to a document that shows evidence for electrical safety compliance.]
4.11 | Electromagnetic compatibility
| Legal Manufacturer/Reseller to certify compliance with the requirements of the latest edition of IEC 61000-6-1 and IEC 61000-6-3. | [Please refer to a document that shows evidence for EMC compliance.]
4.12 | Restricted materials
| The EHC and its constituent components must not contain lead, mercury, cadmium, hexavalent chromium, polybrominated biphenyls (PBB) or polybrominated diphenyl ethers (PBDE). The only exception will be for loads requiring a recyclable lead acid battery. | [Please state any potentially toxic materials both in use and disposal.]
4.13 | Maintenance and servicing provision
| The EHC and load(s) cannot impact the day-to-day operation of the appliance or increase preventive appliance maintenance. The EHC is to be designed to achieve a no-maintenance life of no less than 10 years. | [Please explain how the design of your device does not impact daily appliance operation and how it is designed to achieve a 10-year lifespan.]
4.14 | Essential spare parts
| Based on EHC design and requirements the type and quantity of spare parts, basic installation tools/supplies, and user and technician maintenance manuals (see 4.16 Instructions) must be determined and agreed upon in advance of order placement. As a minimum each EHC is to be supplied with 10 spare fuses of all fuse size and types used in the EHC and, if supplied, used by any defined load(s). The spare fuses are to be attached within or on the EHC and, if supplied, the defined load(s). | [Please refer to a diagram or photo of where the 10 spare fuses are attached on the EHC and, if supplied, the defined load(s).]
| Legal Manufacturer/Resellers are to publish a list of spare parts recommended for purchases of 10 and 50 EHCs and, if supplied, defined load(s). Legal Manufacturer/Reseller must ensure supply of spare parts for a minimum of 5 years from the time of cessation of the last production of equipment. Spare parts are to be provided in kit form for storage in appropriate quantities at National or Sub-national level in the purchasing country, as agreed with the purchasing agency. | [Please refer to the published list of spare parts.] [Please confirm availability of spare parts for said duration.]
4.15 | Disposal and recycling
| The Legal Manufacturer/Reseller is to provide information to the buyer on the hazardous materials contained within the EHC and, if supplied, the defined load(s) with suggestions for resource recovery/recycling and/or environmentally safe disposal. For Legal Manufacturer/Resellers from the European Union WEEE compliance in accordance with European Union Directive 2002/96/EC is mandatory. | [Please state any potentially toxic materials both in use and disposal.]
4.16 | Instructions and labelling
| Each EHC must include a separate user manual and technician installation manual in Arabic, English, French, Mandarin Chinese, Russian and Spanish. An English version of all instructions and manuals are required to be supplied at time of laboratory testing. Instructions to include easy to understand visuals whenever possible to avoid reliance on text. | [Please confirm the manual will be in the requested languages and will cover the requested topics.]
The user manual must include the following information:
| Health and safety guidance | [Please refer the user manual section.]
| Basic operations description | [Please refer the user manual section.]
| Correct handling to avoid damage to the EHC | [Please refer the user manual section.]
| Loads that are compatible and not compatible | [Please refer the user manual section.]
| Minimum average daily Wh of energy available to user based on the minimum solar radiation reference period and clear examples for users (e.g., X quantity of Y device in operation for Z hours per day) | [Please refer the user manual section.]
| Preventive maintenance tasks (e.g., daily, weekly, and monthly) | [Please refer the user manual section.]
| End-of-life resource recovery and recycling procedures. | [Please refer the user manual section.]
4.17 | Training
| If requested, all Legal Manufacturers/Resellers are required to have the capability of providing in-person training in the countries where their product is deployed. Training may be delivered by the Legal Manufacturer/Reseller or a manufacturer representative. The Legal Manufacturer/Reseller must provide user training guidance to enable installation technicians to train users. In addition, supporting video material supplied on DVD and/or on-line can be supplied to assist the instructor when delivering on-site user training. | [Please advise if training is available.]
4.18 | Human factors
| An EHC should be simple to understand, simple to install, simple to use and require minimal maintenance. The EHC should be repairable on-site or replaceable on-site. | [Please refer the user manual section that explains: Installation; Maintenance; Possible repairs.]
4.19 | Warranty
| EHC warranty to be a minimum of five years. | [Please refer to the warranty statement.]
4.20 | Ambient temperature range during transport and storage
| Ambient temperature range during transport and storage to be -30°C to +70°C when the product is inactivated. | [Please state the ambient temperature range during transport and storage.]
4.21 | Ambient humidity range during transport, storage and use
| Ambient humidity range during transport, storage and use to be 5% to 95% RH, non-condensing. | [Please state the ambient humidity range during transport and storage.]
4.22 | Protection against dust and water ingress
| A stand-alone EHC that is not part of the appliance design could be installed in dusty environments. The design of the enclosure should ensure that the EHC is not damaged by dust penetration or splashing water. A minimum of IP54 rating is required. This IP rating requirement does not apply for an EHC that is integrated into the appliance design (i.e., within the body of the appliance). | [Please refer to a document that shows evidence for IP compliance.]
4.23 | Verification
| In accordance with PQS Verification Protocol E007/EHC01.1-VP.1. | [Please refer the test report.]
5. | Packaging
| Materials used for packaging the finished EHC are to be free of ozone-depleting compounds as defined in the Montreal Protocol. | [Please state the materials used for packaging the finished product and whether they are free of CFC compounds as defined in the Montreal Protocol.]
| The packaging is to be of a sturdy export quality and of a commercial standard that will provide adequate protection of the goods for carriage by air, sea and/or road to final destinations worldwide, including remote locations under adverse climatic and storage conditions and high humidity. The packaging is to be not less than 17 kN edge crush resistance with minimum 60% remaining with 90% humidity at a temperature of +70°C (tropical conditions). | [Please explain how the packaging complies with this clause.]
The general specification of shipping containers will be subject to agreement with the individual procurement agencies.