## Spec
PQS Submission template for E007 Solar Direct Drive Basic Energy Harvest Control devices V1.0 (8 November 2022)
| Assessment of appliance: | [Brand & Model here] | For dossier review on [TBC]
| Tested in accordance with PQS verification protocol | E007/EHC 02-VP.1 (22 September 2021)
| To verify compliance with PQS specification below | E007/EHC 02.1 (22 September 2021)
General guidance to applicant: Your application for PQS pre-qualification should contain a dossier of supporting documentation in the order as listed in Section 7 of E007/EHC02.1 and below. Except for the grey cells, there must be a considered response against each clause often with a reference to a supporting document – statements such as “Comply” or “Pass” are not sufficient. Please clearly reference these documents and complete the applicant’s response column with as much detail as possible referencing the particular page and section in a supporting document. Keeping to this guidance will help speed your application.
[Prompts in square brackets should be replaced with appropriate wording or NA (Not Applicable).] Grey cells do not require a response as the text is for information.
NOTE: If it is not clear or doubtful whether a requirement has been met, then this is a non-compliance.
NOTE TO APPLICANT: Please complete the VP sheet as well as the Spec sheet - thank you.
| ASSESSOR’S CONCLUSION |  | Assessed by [Name of assessor]
Section 7: | Requirement | Applicant’s response | Assessor’s comment | Applicant’s response | Assessor’s comment | Applicant’s response | Assessor’s comment
The legal manufacturer or reseller must provide WHO with a prequalification dossier containing the following:
| Dossier examination fee in US dollars. | [Please contact PQS.]
| General information about the legal manufacturer, including name and address. | [Please state or refer to a document.]
| Unique identification reference for the Basic EHC type. | [Please state.]
| Brand name of the Basic EHC. | [Please state.]
| Full specifications of the Basic EHC being offered, covering all the requirements set out in this document, including details of Basic EHC marking and traceability. | [Please refer to a document giving complete details.]
| Full details of the compatible appliance(s) and certification of appliance conformance with specification E003/FZ03 or E003/RF05. | [Please refer to a document giving complete details.]
| Full details of the compatible Type 2 solar power system and certification of conformance with specification E003/PV01. | [Please refer to a document giving complete details.]
| A comprehensive set of photographs including a three quarter view of the Basic EHC. Additional photographs showing all surfaces of the Basic EHC, the interior, USB-A port connection points/cables and any special features. | [Please refer to the photos.]
| Certified photocopies of all type-approvals obtained for the EHC and/or its components, including CE marking and the like. | [Please refer to a document.]
| Certified photocopies of the legal manufacturer’s ISO 9001 quality system certification. | [Please refer to a document.]
| Where relevant, certified photocopies of the legal manufacturer’s ISO 14001 certification, EMAS registration or registration with an equivalent environmental audit scheme. Conformity with an environmental audit scheme is not mandatory; however, preference will be given to manufacturers who are able to demonstrate compliance with good environmental practice. | [Please refer to a document.]
| Laboratory test report(s) proving conformity with the equipment specifications. | [Please refer to a document.]
| Indicative cost of the prequalified SDD appliance including solar power system and optional Basic EHC (cost per unit, per 10 units and per 100 units, EXW (Incoterms 2015). | [Please state or refer to a document.]
Section 4. | Requirement | Applicant’s response | Assessor’s comment
4.1 | General
In all operational modes the Basic will not cause an adverse impact on appliance.
4.2 | Fail safe operation | [Please explain how your device complies with each part of this clause.]
| The Basic EHC must be designed so that if for any reason it fails, safe and adequate power is still supplied to the appliance for continuous operation. The Basic EHC must include a simple and automatic means (requiring no user intervention) to be disconnected or bypassed completely to allow the appliance to continue to operate correctly in the event of Basic EHC failure. Hot swap repair options may be offered by legal manufacturers. The Basic EHC must survive common control fault conditions. Such conditions are as follows but are not limited to: connection errors including reversed load polarity, load voltage mismatch and wiring connections in any sequence possible, high temperature operation, power input over-current conditions and short circuit at the appliance, within the Basic EHC, at load and solar power system. In the event of Basic EHC failure, the appliance will continue to operate correctly as prequalified and remain safely connected to the solar power system. | [Explain or demonstrate how the EHC is designed so that if it fails for any reason, safe and adequate power is still supplied to the appliance for continuous operation.] [Explain how the EHC can be automatically disconnected or bypassed completely to allow the appliance to continue to operate correctly in the event of Basic EHC failure.] [Explain how the EHC is designed to survive common control fault conditions. (e.g. connection errors including reversed load polarity, load voltage mismatch and wiring connections in any sequence possible, high temperature operation, power input over-current conditions and short circuit at the appliance, within the Basic EHC, at load and solar power system.)]
4.3 | Energy harvest
The Basic EHC must manage harvestable solar electricity that can be converted, limited, controlled, diverted to and consumed by other loads.
| A Basic EHC and appliance will be tested only with a laboratory supplied fixed direct current (dc) load designed to test the limits of the system. The energy harvest (in watt-hour/average day) must not exceed 60 Wh/day and the energy harvest efficiency must be greater than 80% and will be measured and reported. | [Please state the energy consumption (in watt-hours per day on average) and the energy harvest efficiency and reference the test(s) where this is confirmed.]
4.4 | Compatibility and compliance with relevant references
| The Basic EHC must be compatible with at least one WHO PQS prequalified appliance and the appliance’s required Type 2 solar power system per E003/PV01. | [Please state at least one compatible WHO PQS prequalified appliance and the appliance’s required Type 2 solar power system (per E003/PV01).]
4.5 | Appliance operation
| Appliance/Basic EHC power on/off switch to simultaneously switch off power to both appliance and Basic EHC. | [Please refer to the test where this is demonstrated.]
| Manufacturer to certify that Basic EHC design prevents loads with a battery from powering the appliance under any operating condition. Certifications to include a description of the method(s) used to prevent load battery from powering the appliance under any operating condition. | [Please certify that EHC design prevents loads with a battery from powering the appliance under any operating condition.]
4.6 | Appliance compatibility
| All Basic EHC must be integrated in the SDD appliance cabinet and must be compatible with at least one WHO PQS prequalified SDD appliance. The location of the USB-A port cannot interfere with regular appliance operations or routine appliance maintenance and must be accessible without requiring the removal of any appliance parts. | [Please state at least one compatible WHO PQS prequalified appliance.] [Please demonstrate how the USB port cannot interfere with appliance operation and maintenance.]
4.7 | Solar power system compatibility
| All the specific solar power system(s) required by the specific appliance(s) the Basic EHC is prequalified with. All Basic EHC are operated on direct current (dc) electricity generated by the solar power system. | [Please give example(s) of specific solar power system(s) as required by specific appliance(s) with which the EHC is prequalified.]
The energy harvest system cannot result in a solar power system larger than that required by the SDD appliance alone.
4.8 | Load compatibility
| Basic EHC are limited to 5 V dc loads through one USB A port where the total daily output is limited to 60 Wh/day. It must be clearly labelled that the USB-A port is 5 Vdc and primarily intended for RTMD power. | [Please show the Basic EHC is limited to a single USB-A port, and is labeled as 5 V dc.]
| Legal manufacturer to certify that the USB-A connector shall be of the ruggedized version rated to 10,000 cycles as certified by EIA-364-09. Legal manufacturer/reseller are to clearly note any incompatible load types. | [Please provide documentation verifying the USB-A port is rated to 10,000 cycles or greater.]
4.9 | Available energy harvest and energy efficiency
| The USB-A port output must not exceed 60 Wh/day. Manufacturer to state the minimum average daily watt hour (Wh/average day) of energy available to user based on the measured results from WHO/PQS/E003 SDD appliance day/night test at the minimum solar radiation reference period of 3.5 kWh/m2 day and provide clear examples for users (e.g., X quantity of Y device in operation for Z hour per day). Results will be published for the total available Wh/average day as measured in the Day/night appliance testing. | [Please state the minimum average daily watt hours (Wh/day) of energy available to users (based on the day/night) and refer to the test where this is demonstrated.]
4.10 | Electrical safety rating
| Legal Manufacturer/Reseller to certify compliance to IEC 60335-1, IEC 60335-2- 24 and IEC 60364-1. | [Please refer to a document that shows evidence for electrical safety compliance.]
4.11 | Electromagnetic compatibility
| Legal Manufacturer/Reseller to certify compliance with the requirements of the latest edition of IEC 61000-6-1 and IEC 61000-6-3. | [Please refer to a document that shows evidence for EMC compliance.]
4.12 | Restricted materials
| The EHC and its constituent components must not contain lead, mercury, cadmium, hexavalent chromium, polybrominated biphenyls (PBB) or polybrominated diphenyl ethers (PBDE). | [Please state any potentially toxic materials both in use and disposal.]
4.13 | Maintenance and servicing provision
| The Basic EHC is to be designed to achieve a no-maintenance life of not less than 10 years. If necessary, any preventive maintenance required by the Basic EHC and load(s) cannot create an adverse impact on the appliance causing it to not operate correctly. | [Please explain how the design of your device does not impact daily appliance operation and how it is designed to achieve a 10-year lifespan.]
4.14 | Essential spare parts
| Based on Basic EHC design and requirements the type and quantity of spare parts, basic installation tools/supplies, user and technician maintenance manuals (see clause 4.16 Instructions), must be determined and agreed upon in advance of order placement. As a minimum each Basic EHC is to be supplied with 10 spare fuses of all fuse size and type used in the Basic EHC. The spare fuses are to be attached within or on the appliance. | [Please refer to a diagram or photo of where the 10 spare fuses are attached on the EHC and, if supplied, the defined load(s).]
| Legal manufacturer/resellers are to publish a list of spare parts recommended for purchases of 10 and 50 SDD appliances with a Basic EHC. Legal manufacturer/reseller must ensure supply of spare parts for a minimum of 5 years from the time of cessation of the last production of equipment. Spare parts are to be provided in kit form for storage in appropriate quantities at National or Sub-national level in the purchasing country, as agreed with the purchasing agency. | [Please refer to the published list of spare parts.] [Please confirm availability of spare parts for said duration.]
4.15 | Disposal and recycling
| The Legal Manufacturer/Reseller is to provide information to the buyer on the hazardous materials contained within the EHC and, if supplied, the defined load(s) with suggestions for resource recovery/recycling and/or environmentally safe disposal. For Legal Manufacturer/Resellers from the European Union WEEE compliance in accordance with European Union Directive 2002/96/EC is mandatory. | [Please state any potentially toxic materials both in use and disposal.]
4.16 | Instructions and labelling
Each SDD with an integrated, Basic EHC shall include a separate section dedicated to the Basic EHC use and operations in the applicable manuals and documentation. These sections shall be found in the user manual and technician installation manual and be provided in Arabic, English, French, Mandarin, Chinese, Russian and Spanish. An English version of all instructions and manuals are required to be supplied at time of laboratory testing. Instructions to include easy to understand visuals whenever possible to avoid reliance on text.
The user manual must include the following information specific to the Basic EHC:
| Health and safety guidance | [Please refer the user manual section.]
| Basic operations description | [Please refer the user manual section.]
| Correct handling to avoid damage to the EHC | [Please refer the user manual section.]
| Loads that are compatible and not compatible | [Please refer the user manual section.]
| Minimum average daily Wh of energy available to user based on the minimum solar radiation reference period and clear examples for users (e.g., X quantity of Y device in operation for Z hours per day) | [Please refer the user manual section.]
| End-of-life resource recovery and recycling procedures. | [Please refer the user manual section.]
4.18 | Human factors
| Basic EHC should be simple to understand, simple to use and require minimal or no maintenance. The Basic EHC should be repairable on-site or replaceable on-site. | [Please refer the user manual section that explains: Installation; Maintenance; Possible repairs.]
4.19 | Warranty
| Basic EHC warranty to be a minimum of five years. | [Please refer to the warranty statement.]
4.20 | Ambient temperature range during transport and storage.
| Ambient temperature range during transport and storage to be -30°C to +70°C when the product is inactivated. | [Please state the ambient temperature range during transport and storage.]
4.21 | Ambient humidity range during transport, storage and use
| Ambient humidity range during transport, storage and use to be 5% to 95% RH, non-condensing. | [Please state the ambient humidity range during transport and storage.]
4.22 | Protection against dust and water ingress
| The design of the enclosure should ensure that the Basic EHC is not damaged by dust penetration or splashing water. | [Please describe how this requirement is met.]
4.23 | Verification
| In accordance with PQS Verification Protocol E007/EHC02-VP.1. | [Please refer the test report.]
5. | Packaging
| Materials used for packaging the finished EHC are to be free of ozone-depleting compounds as defined in the Montreal Protocol. | [Please state the materials used for packaging the finished product and whether they are free of CFC compounds as defined in the Montreal Protocol.]

## VP
| ASSESSOR’S CONCLUSION |  | Assessed by [Name of assessor]
| Acceptance criterion | Applicant’s response | Assessor’s comment | Applicant’s response | Assessor’s comment | Applicant’s response | Assessor’s comment
5.3.4 Test 2: Excess load and solar radiation variation
| The appliance remains in correct operation and sustains acceptable vaccine storage temperature range and/or water pack freezer temperature range per applicable E003 specifications. Energy harvest continues and load is limited to the legal manufacturers specification of a single output value at 5 V dc (2.5 watts up to a maximum of 5 watts). EHC efficiency is calculated equal to or greater than 80%. | [Test report should show all data, including load output power at various array operating voltages and efficiency greater than 80%]
| No obvious or quantifiable damage to Basic EHC and appliance and the Basic EHC and appliance continue to operate correctly. The Basic EHC automatically returns to correct operation when the excessive load is removed. | [Test report should show all data]
5.3.5 Test 3: Connection sequence
| After all possible connection sequences have been tested the Basic EHC and the appliance does operate correctly. | [Test report should show all connection squences tested and the results.]
5.3.6 Test 4: Reverse polarity connections
| At the end of the reverse polarity testing and after polarity connections are all correctly reconnected the appliance cooling system and Basic EHC now operate correctly. | [Test report should show all data]
5.3.7 Test 5: Short circuit
| After short circuit is established the Basic EHC must not create an adverse impact for the appliance cooling system operation. After short circuit is corrected the Basic EHC must demonstrate that the cooling system will operate correctly, the Basic EHC is undamaged, will operate correctly and restores energy harvest diversion to the fixed DC load. | [Test report should show all data]