## Spec
E010 Submission to PQS for approval: Needle cutter for health waste management
| Compliance with PQS specification: | E010/NC01.1 | Brand & Model:
| As shown by testing in accordance with PQS verification protocol: | E010/NC01-VP.1
Instructions in [square brackets] are for your guidance. Please delete these instructions and replace with evidence e.g. a precise reference (including page and section numbers) showing compliance with that particular PQS specification requirement. In general a "pass" or "complies" statement alone is not acceptable - evidence for compliance must be provided.
One document may provide evidence for several clause requirements. Please carefully reference the sub-sections of any such document so that it is clear that the clause requirement has been met. This will help speed the pre-qualification process.
Greyed-out cells are for information and do not require a response.
NOTE: If it is not clear or doubtful whether a requirement has been met, then this is a non-compliance and pre-qualification might be delayed.
| Please also complete the VP sheet. |  | REVIEWER’S SUMMARY: |  | REVIEWER’S SUMMARY: |  | REVIEWER’S SUMMARY:
Section 7. | Product dossier | Applicant's submission | Reviewer's remarks [Name & Date] | Applicants response [Name & Date] | Reviewer's remarks [Name & Date] | Applicants response [Name & Date] | Reviewer's remarks [Name & Date]
The legal manufacturer or reseller is to provide WHO with a pre-qualification dossier containing the following:
| Dossier examination fee in US dollars. | [Please arrange payment with WHO PQS.]
| General information about the legal manufacturer, including name and address. | [Please state.]
| Unique identification reference for the device type. | [Please state.]
| Full specifications of the product being offered, covering all the requirements set out in this document, including details of product marking and traceability. | [Please list the document(s) provided in the submission for pre-qualification).]
| Certified photocopies of all type-approvals obtained for the product, including CE marking and the like. (If applicable) and the like. | [Please refer to certified copies of all type approvals.]
| Certified photocopies of the legal manufacturer’s ISO 9001: 2000 quality system certification. | [Please refer to a document.]
| Where available, laboratory test report(s) proving conformity with the product specifications. | [Please refer to a test report where the product has been tested in accordance with PQS verification protocol WHO/PQS/E010/NC01-VP.1.]
| Ten complete samples of the device, including spare parts and instructions. | [Please send sample(s) to the address(es) to be provided.]
| Indicative cost of the product per unit, per 100 units and per 1,000 units EXW (Incoterms 2000). | [Please state or refer to a document.]
Section 4. | Requirements
4.1 | General
A manually operated needle cutter comprises a cutting device and a needle container which allows health workers to make used syringes safe and harmless at the point of use immediately after administering an injection. Where not intended for stationary applications, the device should be easily portable. The device must be safe, easy and convenient to use, easily cleaned, affordable and reliable. Needles or needle residues are stored until the needle container is filled up to its designed maximum fill line, at which point the container must be removed, capped, and either disposed of or emptied. In the case of disposable needle cutter devices with integral containers and cutting assemblies, it should be ensured that the entire device is disposed of properly.
4.2 | Performance
4.2.1 | Needle size
| The device should disable wet or dry needles, 10-76 mm in length and 18-28 gauge in diameter. | [Please state the type and size of needles the device can disable.]
4.2.2 | Needle/syringe type
| The device should disable all ISO compliant syringe/needle combinations. | [Please refer to document or test report showing compliance to this requirement.]
4.2.3 | Needle insertion
| All needles in the size range specified in clause 4.2.1 should insert easily into the device, with little or no force. | [Please refer to document or test report showing compliance to this requirement.]
4.2.4 | Cycle time
| Needle removal or cutting devices should have a maximum cycle time per needle not exceeding 5 seconds. | [Please refer to document or test report showing compliance to this requirement.]
4.2.5 | Needle entry geometry
| The needle aperture must be designed so that the needle can be inserted into the port at any angle lying within a 60 degree cone whose apex is centred on the aperture. | [Please refer to document or test report showing compliance to this requirement.]
4.2.6 | Complete cutting
| The cutting blade configuration should ensure that the needle, needle hub, or syringe nozzle is completely cut or sheared. Incomplete shearing or other modes of disabling the needle, such as crimping or bending, are not allowed. | [Please describe how the needle is cut.]
4.2.7 | Self-clearing mechanism
| The cutting mechanism must be self-clearing. Syringe or needle remnants remaining in the device must not impair its operation. | [Please describe how the device clears the needle after cutting.]
4.2.8 | Needle container attachment
| The needle container, if separable, must attach securely to the device so that tipping or dropping it does not separate the container from the cutting assembly. Attachment of the needle container to, and subsequent removal from the cutting assembly should be safe, clean and easy. There must be no risk of needle stick injury during these operations. | [Please describe or provide photos of the needle container and how it is attached to the cutting device.]
4.2.9 | Operating life
| • Non-disposable devices: must withstand at least 100,000 cycles of operation, and require no major maintenance or user intervention, other than cleaning and lubrication, no more frequently than once every 10,000 cycles of operation. • Disposable devices: must withstand a minimum of 300 cycles of operation, prior to being discarded. Preferably, the number of cycles of operation should match the maximum capacity of the needle container. | [Please refer to document or test report showing compliance to this requirement.]
4.2.10 | Splatter
| During or after normal use of the device, there should be no detectable contamination of: • Exposed skin, mucous membrane, or clothing of the operator. • Work surfaces or other surfaces adjacent to and surrounding the device. • The outer surfaces of the device which are accessible to the user, with the exception of the needle entry target area. | [Please refer to document or test report showing compliance to this requirement (VP test 6).]
4.2.11 | Drop test (complete device)
| The performance and safety of the cutting assembly must not be compromised by dropping from a height of 1,000mm onto a smooth concrete surface in accordance with the test method in IEC 60068-2-32. In devices with a removable needle container, the container must NOT be detached. | [Please refer to document or test report showing compliance to this requirement.]
4.2.12 | Drop test (needle container only)
| The container when full of needles and with the closure device engaged, should be dropped 100 times onto a smooth concrete surface from a height of 1000mm. No needles must fall out of the container. Not more than one needle must pierce any of the sides. The container must not be seriously damaged by the test. | [Please refer to document or test report showing compliance to this requirement.]
4.2.13 | Tilt angle
| The device must not tip over, whether empty or full, when placed on a 15 degree non-slip plane with its short axis parallel to the line of tilt in general accordance with the test method in AS 4031:1992, Appendix D. | [Please refer to document or test report showing compliance to this requirement.]
4.2.14 | Leak-proof
| The device must not leak any liquid contents when placed in the upright position at any angle between 0 and 15 degrees. | [Please refer to document or test report showing compliance to this requirement.]
4.2.15 | Needle escape prevention
| The cutting assembly must be designed to prevent the migration of cut needles from the needle container into the needle aperture. | [Please describe or provide evidence of compliance to this requirement.]
4.2.16 | Cutting device closure mechanism
| If the device is intended to be carried with the needle container attached, the needle aperture must have a closure mechanism to prevent needles from falling out of the attached needle container in any orientation of the assembly. | [If applicable, please describe closure mechanism.]
4.2.17 | Needle container closure mechanism
| If the needle container is intended to be detachable, it must have a secure closure mechanism that prevents spillage of sharps after detachment from the cutting assembly, whatever the orientation of the container. Preferably the closure mechanism should engage automatically upon removal of the full container from the cutting assembly. The container must pass the dropping, toppling and leakage tests described in BS7320:1990, Appendix D and Appendix E. | [Please refer to document or test report showing compliance to this requirement.]
4.2.18 | Needle container puncture resistance
| The needle container must pass the penetration resistance test in BS7320:1990, Appendix C. | [Please refer to document or test report showing compliance to this requirement.]
4.2.19 | Needle container capacity
| The needle container must hold at least 150 nbr. 20mm needles, and/or needle remains, without affecting operation of the device. | [Please refer to document or test report showing compliance to this requirement.]
4.2.20 | Needle container capacity indication
| The needle container must be translucent enough to allow the user visually to detect the level of needles in the container. The sides of the container must be clearly marked with a fill line in accordance with BS7320:1990. | [Please describe or provide photos of the needle container and capacity indicator.]
4.2.21 | Tamper proofing
| The needle aperture must be designed to prevent any part of the hand from entering. Needles should not protrude from the needle container when it is filled up the level of the fill line. | [Please describe or provide photos showing compliance to this requirement.]
4.3 | Environmental requirements
4.3.1 | Operating environment
| The performance of the device must not be compromised by exposure to continuous ambient conditions of 43°C and 90% relative humidity for a period of one week when the needle container is in any condition between empty and full. | [Please refer to document or test report showing compliance to this requirement.]
4.3.2 | Chemical resistance
| The device should be resistant to saline solution and to mild chemical cleaning agents, including diluted bleach. | [Please refer to document or test report showing compliance to this requirement.]
4.3.3 | Bio-hazard marking
| The needle container must be clearly marked with the international bio-hazard warning not less than 35mm diameter, printed in black or red, on each of the front and back faces of the box. | [Please provide photos of the container showing the bio-hazard markings.]
4.4 | Physical characteristics
4.4.1 | Overall dimensions
| If intended to be portable, the device must be compact and have minimal protrusions. It must be transportable over long distances on foot by the lowest quartile of female operator without inconvenience and with minimal dismantling. | [Please state device dimensions.]
4.4.2 | Weight
| If intended to be portable, the empty device, complete with empty needle container, should weigh a maximum of 750 grams. | [Please state device weight.]
4.5 | Interface requirements
4.5.1 | Disposal
| The needle container or integral cutter and container, if disposable, must be able to fit into a protected needle pit with a 10 cm inner diameter entry tube. | [If applicable, please describe how the container meets this requirement.]
4.6 | Human factors
4.6.1 | Generally
| The product must be useable by the widest practicable range of active health workers, regardless of age, gender, size or minor disability, including long-sighted and short-sighted people without glasses, in accordance with the general principles laid out in ISO 20282-1: 2006. | [Please describe how the device meets this requirement.]
4.6.2 | Skill level
| It must be possible for health workers to operate the device with minimal training. | [Please describe how the device meets this requirement.]
4.6.3 | User position
| The device must be comfortable to operate by 5th to 95th percentile adults in standing and seated positions with the device resting on a firm surface. | [Please refer to document or test report showing compliance to this requirement (VP test 13).]
4.6.4 | Handedness
| The device must be equally useable by left and right handed health workers. | [Please refer to document or test report showing compliance to this requirement (VP test 13).]
4.6.5 | Activation force
| The maximum force required to cut a standard (21 g) needle, needle hub, or syringe nozzle should not exceed 67 N. | [Please refer to document or test report showing compliance to this requirement (VP test 5).]
4.6.6 | Repetitive use
| The alignment of the cutting mechanism handles should avoid ulnar deviation and should be designed to prevent discomfort or the occurrence of repetitive strain injuries when the device is routinely used by a single operator for 200 cycles per day. | [Please describe how the device meets this requirement.]
4.6.7 | Pinch points
| Normal use should not result in pinching of the operator’s hands. | [Please provide photos or describe how the device meets this requirement.]
4.6.8 | Smoothness of operation
| Complete needle removal or destruction must be achieved with a single smooth hand movement. | [Please describe how the device meets this requirement.]
4.6.9 | Hand to needle distance
| The distance from the needle to the hand holding or operating the needle cutter must exceed 50 mm while operating the device. | [Please provide evidence of compliance to this requirement.]
4.6.10 | Blade edge protection
| The device’s cutting blades must not expose the user to cut hazards, either with or without the needle container connected. | [Please provide photos or describe how the device meets this requirement.]
4.6.11 | Cleaning
| External parts and reusable internal parts accessible to the user must be cleanable with standard mild cleaning agents. | [Please describe how the device meets this requirement.]
4.7 | Materials
4.7.1 | Generally
| The materials used must be selected to minimize surface degradation or corrosion arising from repeated use up to the specified minimum number of operating cycles, when the device is cleaned and lubricated in accordance with the manufacturer’s recommendations. | [Please describe or refer to a test report how the device meets this requirement.]
4.8 | Warranty
| The device must be warranted to meet all physical and performance requirements defined in this specification over the relevant operating life as specified in clause 4.2.9. | [Please refer to manufacturer's warranty statement.]
4.9 | Servicing provision
| • Non-disposable devices: The only maintenance required during the design life of the device should be consumable part replacement, regular cleaning and lubrication. The minimum life cycle of consumable parts should be 25,000 removals, cuts or destruction cycles. Three additional sets of consumable parts should be provided with the device, together with product-specific service tools, if required. Used consumable parts should not be re-furbished but must be disposed of in the needle container. • Disposable devices: Must be maintenance free. | [Please describe how device meets this requirement.]
4.10 | Disposal and recycling
| Device must be able to be disposed of in the medical waste stream in accordance with the prevailing government approved and regulated waste disposal practices applicable in the country of use. | [Please state hox box should be properly disposed of after final use.]
4.11 | Instructions
| User and maintenance instructions must be available in Arabic, English, French, Mandarin Chinese, Russian and Spanish. Labelling on the device should include clear pictorial instructions. | [Please provide instructions (as applicable) on the assembly, use, and maintenance of the device. Please confirm the instructions are available in the requested languages.]
4.12 | Training
Printed user-operations and maintenance instructions specifically directed at the health
5. | centre staff shall be pictorial. All key information is to be summarized on a single
pictogram sheet that may be fixed onto the appliance. The sheet is to be sufficiently
| Training on the assembly, use and maintenance (if any) of the device will be provided by the health care programme when the device is first introduced, and subsequently during supervisory visits. | [Although not required, please state if there is any training available (videos, online, etc.).]
4.13 | Verification
| In accordance with PQS Verification Protocol E10/NC01-VP | [Please refer to the test report detailing results in accordance with WHO/PQS/E10/NC01-VP and see "VP Sheet" in this submission template.]
5.0 | Packaging
| Materials used for packaging the finished product are to be free of ozonedepleting compounds as defined in the Montreal Protocol. The general specification of shipping containers will be subject to agreement with the individual procurement agencies. | [Please state how the packing meets this requirement and refer to recommendations for correct disposal.]

## VP
E010 Submission to PQS for approval: Needle cutter for health waste management
| Compliance with PQS specification: | E010/NC01.1 | Brand & Model:
| As shown by testing in accordance with PQS verification protocol: | E010/NC01-VP.1
Instructions in [square brackets] are for your guidance. Please delete these instructions and replace with evidence or a precise reference (including page and section numbers) showing compliance with that particular PQS specification requirement. In general a "pass" or "complies" statement alone is not acceptable.
One document may provide evidence for several clause requirements. Please carefully reference the sub-sections of any such document so that it is clear that the clause requirement has been met. This will help speed the pre-qualification process.
Greyed-out cells do not require a response.
NOTE: If it is not clear or doubtful whether a requirement has been met, then this is a non-compliance and pre-qualification might be delayed.
Please ensure the Spec sheet is completed.
Section 7. | VP Clause | Applicant's submission | Assessor's remarks [Name & Date] | Applicants response [Name & Date] | Assessor's remarks [Name & Date] | Applicants response [Name & Date] | Assessor's remarks [Name & Date]
5.1 | Evidence of conformity assessment
| Products may carry the CE mark and/or equivalent internationally accepted evidence of conformity assessment, but this is not a mandatory requirement. | [Please state if product carries the CE or other mark.]
5.2 | Number of samples
| The Legal Manufacturer or Reseller must supply the testing laboratory with a full duplicate set of the Product Dossier already supplied to WHO in accordance with the requirements of specification clause 7. A minimum of ten samples of the product are required. Provide one sample of the packaging described in specification clause 5. | [Please provide the testing laboratory a draft dossier for checking and amendment. This will help to ensure that the dossier is complete.]
5.3 | Test procedure
5.3.1 | Test 1: Type examination
| The test report must clearly show that the requirements of each section and sub-section in the type examination are individually and clearly met. | [Please refer to the test report and see the specification sheet in this template where precise references to each specification requirements are required. (Test 1 requirements refer to many specification clause requirements. See E010/NC01.1.)]
5.3.2 | Test 2: Needle cutting performance
Number of samples: One. • Step 1: Assemble a number of needle/syringe combinations in order to include at least ten of each of the following attributes: o Needles of the following gauges: 18, 23, and 28 o Needles of the following lengths: 10 mm, 25 mm, and 76 mm. o Syringes of the following types: fixed luer, luer-lock, luer-slip and snap-on. • All syringes should be compliant with ISO 7786 or 8537 and needles with ISO 7864. • Step 2: Insert one of each of the assembled needle/syringe combinations into the needle aperture at various angles within 30 degrees of vertical (a 60 degree cone around the aperture). • Step 3: Process five of each of the assembled needle/syringe combinations using the device. • Step 4: Inspect needle and syringe remnants. Verify that needles or needle hubs are completely cut or sheared. • Step 5: Inspect needle cutter to ensure cutting pathway is clear following each cut to verify self-clearing mechanism. • Step 6: Aspirate water with the remaining five of each of the assembled needle/syringe combinations and expel liquid. • Step 7: Repeat Step 3 - Step 5 with wet needles and syringes.
| • Acceptance criteria: - All needle types should insert easily into the needle aperture, with little or no force. - A needle inserted at any angle within a 60 degree cone of the aperture should be able to enter. - Needles of all sizes and types tested should be disabled. - Needles should be completely cut or sheared. - Syringe or needle remnants remaining in the device must not impair its operation. • Rejection criteria: Failure to meet any of the acceptance criteria | [Please refer to a precise section in the test report showing the results of the test.]
5.3.3 | Test 3: Needle penetration test
Number of samples: Sufficient needle containers to prepare the 12mm x 12mm test samples required by BS 7320: 1990, Appendix C. • Step 1: Follow the method of the test describe in BS 7320: 1990, Appendix C, except that the hypodermic needles used should be 23 gauge x 25mm needles as fitted to AD syringes to ISO 7886 - part 3. Record the measured force for each penetration.
| • Acceptance criteria: The average of forces needed to penetrate samples taken from each position must not be less than 15 N, and the minimum force required to penetrate any sample taken from any position must not be less than 12.5 N. • Rejection criterion: Failure to meet either or both of the acceptance criteria. | [Please refer to a precise section in the test report showing the results of the test.]
5.3.4 | Test 4: Cycle time
Number of samples: One. • Step 1: Assemble ten syringes with a luer-lock hub and 23 gauge needles. Syringes should be compliant with ISO 7786 or 8537 and needles with ISO 7864. Measure the time needed for a single person to completely process all ten needles sequentially using the device. Record the time elapsed. • Step 2: Repeat test and record the average of the results.
| • Acceptance criteria: The maximum time allowed to process ten needles sequentially must be less than or equal to 50 seconds. The mean time to cut one needle must be less than or equal to 5 seconds. • Rejection criterion: Failure to meet the acceptance criteria. | [Please refer to a precise section in the test report showing the results of the test.]
5.3.5 | Test 5: Activation force and operating life
Number of samples: One. • Step 1: Assemble 20 syringes with a luer-lock hub and 23 gauge needles. Syringes should be compliant with ISO 7786 or 8537 and needles with ISO 7864. • Step 2: Process 10 of the assembled syringes and measure the force required to cut each syringe/needle with the device. Record the mean force for the ten tests. • Step 3: Set up device to repeatedly cut 23 gauge needles or a length of 23 gauge stainless steel tubing (type 304). • Step 4: Operate the needle cutter 500 times. (If the needle cutter is disposable, and reaches the maximum fill level prior to 500 cycles, then stop at the maximum fill level.) • Step 5: Process the remaining 10 syringe/needles assembled in step 1. Measure the force required to cut each syringe/needle with the device. Record the mean force for the ten tests.
| • Acceptance criteria: Activation force before and after 500 cycles of operation (or the number of cycles required to fill an integral container) must not exceed 67N. • Rejection criteria: Failure to meet any of the acceptance criteria. | [Please refer to a precise section in the test report showing the results of the test.]
5.3.6 | Test 6: Splatter test
Number of samples: One. • Step 1: Operator should wear a clean white coat. Assemble ten syringes of each of the following types: fixed luer, luer-lock, luer-slip. Fit with 23 gauge needles. Syringes should be compliant with ISO 7786 or 8537 and needles with ISO 7864. Make a strong solution of food coloring in a vivid color. Draw up a food coloring solution with each syringe and expel liquid. Place a clean, 1 meter square sheet of white paper on the work surface centered underneath the device. • Step 2: Process all thirty syringe and needles with the device. After each one, inspect the operator’s skin and clothing, the outside surfaces of the device, and the white paper for visible splattering of dye.
| • Acceptance criteria: No visible spots of dye should be present on the paper, the operator, or the outside surfaces of the device other than the needle entry area. • Rejection criterion: Failure to meet the acceptance criteria. | [Please refer to a precise section in the test report showing the results of the test.]
5.3.7 | Test 7: Capacity
Number of samples: One. • Step 1: Assemble 150 syringes with 20mm needles, 23 gauge, or the manufacturer’s stated maximum capacity number of needles. • Step 2: Process all needles with the device. If the needles ‘stack’, shake the container gently. • Step 3: Close the needle aperture.
| • Acceptance criteria: The maximum capacity must equal or exceed 150 needles or the manufacturer’s stated maximum capacity without passing the maximum limit line on the container. No needles must penetrate the container and there must be no other visible damage or distortion to the container. The closure mechanism must work correctly, and the needles must not affect the operation of the device. • Rejection criterion: Failure to meet any of the acceptance criteria. | [Please refer to a precise section in the test report showing the results of the test.]
5.3.8 | Test 8: Drop test (complete device)
Number of samples: One. • Step 1: Perform drop test on complete device with needle container attached according to IEC 60068-2-32: 1975: (Procedure 1). Drop device twice onto a smooth concrete surface from a height of 1000 mm. • Step 2: Inspect the device for visible damage. • Step 3: Process ten 23 gauge needles with each device.
| • Acceptance criteria: After dropping from a height of 1000mm, the device should not be seriously damaged and the needle container must not be separated from the cutting assembly. Needles must be easily and completely disabled when device is subsequently tested. • Rejection criterion: Failure to meet any of the acceptance criteria. | [Please refer to a precise section in the test report showing the results of the test.]
5.4.9 | Test 9: Drop test (separable needle container only)
Number of samples: One. • Step 1: Load the needle container with needles to its indicated fill line. Close and seal the aperture. • Step 2: Place the loaded container in a large tumble box arranged so that each drop is from a height of 1000mm. Inspect the box after 5, 10, 50 and 100 drops and record signs of deterioration or needle piercing. Stop the test if inspection shows that more than one needle has penetrated.
| • Acceptance criteria: After 100 drops, no needles should have fallen out of the container; the needle container should not be seriously damaged, the lid should remain secure, and no more than one needle should have penetrated the container walls. • Rejection criterion: Failure to meet any of the acceptance criteria. | [Please refer to a precise section in the test report showing the results of the test.]
5.4.10 | Test 10:Stability and spillage test
Number of samples: One. • Step 1: Prepare an adjustable plane tilt-test apparatus as described in AS 4031-1992, Appendix D. • Step 2: Prepare an empty device. Do not close the sharps aperture in any of the subsequent tests. • Step 3: Set the tilting plane at 15° to the horizontal. Place the device on a sloping plane so that its short axis is parallel to the line of tilt and the needle aperture is on the downward side of the slope. If not possible to meet both of these conditions simultaneously, then Steps 4-9 below should be performed under each condition separately. • Step 4: With the tilting plane at 15° and the empty device placed as described above, note whether the device remains standing or whether it topples. • Step 5: Fill the container to half its maximum capacity with needles. Return the device to the 15° plane oriented as described in Step 3 and note whether the device remains standing or whether it topples. • Step 6: Increase the angle of the plane until the box just topples. Note the toppling angle and record whether any needles spill or protrude from the device. Carry out this procedure a total of 10 times, gently shaking the container between each test to re-centre the contents. • Step 7: Fill the container to its fill line with needles. Return the device to the 15° plane oriented as described in Step 3 and note whether the device remains standing or whether it topples.
• Step 8: Increase the angle of the plane until the box just topples. Note the toppling angle and record whether any needles spill or protrude from the device. Carry out this procedure a total of 10 times, gently shaking the container between each test to re-centre the contents. • Step 9: Fill the container to 5% of its total volume with water. Return the device to the 15° plane oriented as described in Step 3 and note whether the device remains standing or whether it topples. Observe whether any liquid leaks from the device.
| • Acceptance criteria: The box must not topple in any of the four trials at 15°. No part of any needle must spill or protrude from the sharps aperture after any of the Step 6 or Step 8 trials. No liquid must leak from the device during the Step 9 trials. • Rejection criteria: Failure to meet any of the acceptance criteria. | [Please refer to a precise section in the test report showing the results of the test.]
5.4.11 | Test 11: Environmental resistance
Number of samples: Two. • Step 1: Prepare a normal saline solution with 9 grams sodium chloride per liter of water. Submerse one device in saline solution for one minute. Remove from solution and invert to drain liquid. Load device with needles to its maximum capacity. • Step 2: Leave the other device empty. Close the apertures in accordance with the pictorial instructions on the device. • Step 3: Set a test chamber temperature to 43°C and 90% relative humidity. Place the empty and loaded devices inside for one week. • Step 4: Remove the devices from the test chamber and record their condition. • Step 5: Process ten luer lock hub syringes with 23 gauge needles with each device. (Syringes should be compliant with ISO 7786 or 8537 and needles with ISO 7864.)
| • Acceptance criteria: No visible corrosion or rusting of needle cutter. Moving parts must work smoothly. Needles must be easily and completely disabled. • Rejection criterion: Failure to meet either or both of the acceptance criteria. | [Please refer to a precise section in the test report showing the results of the test.]
5.4.12 | Test 12: Cleaning resistance
Number of samples: Three. • Step 1: Clean devices according to manufacturer’s instructions. Clean one device with a solution of 1:10 5% hypochlorite bleach to water. Clean another device with a solution of soap and water. Clean the third device with alcohol. • Step 2: Leave devices for one day in the ambient testing environment of the laboratory. Record ambient conditions. • Step 3: Inspect devices for corrosion and damage and record their condition. • Step 4: Process ten luer lock hub syringes with 23 gauge needles with each device. (Syringes should be compliant with ISO 7786 or 8537 and needles with ISO 7864.)
| • Acceptance criteria: No visible corrosion or degradation of needle cutter or container. Moving parts must work smoothly. Needles must be easily and completely disabled. • Rejection criterion: Failure to meet either or both of the acceptance criteria. | [Please refer to a precise section in the test report showing the results of the test.]
5.4.13 | Test 13: Human factors
Number of samples: One • Step 1: Operate device in while in a standing position with device resting on a firm surface such as a counter top. • Step 2: Operate device while in a seated position with device resting on a firm surface such as a table. • Step 3: Operate device with the right hand. • Step 4: Operate device with the left hand.
| • Acceptance criteria: The device must be considered to be comfortable and safe to operate by 5th to 95th percentile adults in standing and seated positions. The device must be equally useable by left and right handed individuals. • Rejection criterion: Failure to meet any of the acceptance criteria. | [Please refer to a precise section in the test report showing the results of the test.]