## Spec
E010 Submission to PQS for approval: Safety box for the disposal of used sharps
| Compliance with PQS specification: | E010/SB01.1 | Brand & Model:
| As shown by testing in accordance with PQS verification protocol: | E010/SB01-VP.1
Instructions in [square brackets] are for your guidance. Please delete these instructions and replace with evidence e.g. a precise reference (including page and section numbers) showing compliance with that particular PQS specification requirement. In general a "pass" or "complies" statement alone is not acceptable - evidence for compliance must be provided.
One document may provide evidence for several clause requirements. Please carefully reference the sub-sections of any such document so that it is clear that the clause requirement has been met. This will help speed the pre-qualification process.
Greyed-out cells are for information and do not require a response.
NOTE: If it is not clear or doubtful whether a requirement has been met, then this is a non-compliance and pre-qualification might be delayed.
| Please also complete the VP sheet. |  | REVIEWER’S SUMMARY: |  | REVIEWER’S SUMMARY: |  | REVIEWER’S SUMMARY:
Section 7. | Product dossier | Applicant's submission | Reviewer's remarks [Name & Date] | Applicants response [Name & Date] | Reviewer's remarks [Name & Date] | Applicants response [Name & Date] | Reviewer's remarks [Name & Date]
The legal manufacturer or reseller is to provide WHO with a pre-qualification dossier containing the following:
| Dossier examination fee in US dollars. | [Please arrange payment with WHO PQS.]
| General information about the legal manufacturer, including name and address. | [Please state.]
| Unique identification reference for the device type. | [Please state.]
| Full specifications of the product being offered, covering all the requirements set out in this document, including details of product marking and traceability. | [Please list the document(s) provided in the submission for pre-qualification).]
| Certified photocopies of all type-approvals obtained for the product, including CE marking and the like. | [Please refer to certified copies of all type approvals.]
| Certified photocopies of the legal manufacturer’s ISO 9001 2000 quality system certification. | [Please refer to a document.]
| Where available, laboratory test report(s) proving conformity with the product specifications. | [Please refer to a test report where the product has been tested in accordance with PQS verification protocol WHO/PQS/E010/SB01-VP.1.]
| One sample of each size and type of box. One sample of the instruction insert. | [Please send sample(s) to the address(es) to be provided.]
| Indicative cost of the product per 100, 1,000 and 10,000 units, EXW (Incoterms 2000). | [Please state or refer to a document.]
Section 4. | Requirements
4.1 | General
Sharps safety box of 1.25, 2.5, 5, 10, 15 or 20 litre nominal storage capacity intended safely and efficiently to contain, transport and store used sharps until final destruction, safe disposal or recycling.
4.2 | Performance
4.2.1 | Functionality
| The safety box must safely contain contaminated sharps: • at the point of use; • during temporary storage; • during handling and transport to the point of treatment and final disposal. | [Please state how the box will safely contain contaminated sharps at the point of use, during storage, and during transport.]
4.2.2 | Shipping and storage volume before use
| Boxes must be supplied flat-packed or nested to minimize shipping and storage volume. | [Please state or provide photos showing how boxes can minimize storage volume.]
4.2.3 | Nominal capacity
| Boxes must accept no less than 20 nbr. 0.5ml AD syringes per nominal litre of storage capacity. This capacity is to be achieved when syringes are dropped in randomly, needle first, with 25mm unsheathed non-retractable needles attached and plungers fully depressed. No syringe must protrude from the container or above the fill line and the box must be capable of being correctly and permanently closed without any risk of needle-stick injury. | [Please refer to document or test report showing compliance to this requirement.]
4.2.4 | Maximum capacity:
| The maximum capacity is allowed to exceed the nominal capacity of 20 syringes per nominal litre provided all the conditions of clause 4.2.3 still apply. | [Please refer to document or test report showing compliance to this requirement.]
4.2.5 | Sharps aperture
| Boxes must be fitted with a sharps aperture, capable of receiving syringes and needle assemblies of all standard sizes up to and including 20 ml, together with other sharps. It must be possible to close and seal the aperture at any time between empty and full to maximum capacity. The closure mechanism must pass the test for security of attachment of aperture closure devices described in BS 7320:1990. | [Please refer to document or test report.]
4.2.6 | Handles
| Boxes must be supplied with a handle or other lifting device which allows the container to be carried safely with one hand. The lifting device must be positioned above the fill line, must not obstruct access to the sharps aperture, and must be sufficiently robust to ensure that it does not to break during use and during transport to the disposal site. It must remain attached to the box when the box is filled with sharps to its maximum capacity and tested in accordance with BS 7320:1990. | [Please describe or provide photos of the box handles showing how it meets the specification requirements.]
4.2.7 | Colour
| Boxes can be the colour of unbleached sulphate board, or non-chlorine bleached white, or yellow. | [Please state or provide photos of box colour.]
4.2.8 | Bio-hazard marking
| Boxes must be clearly marked with the international bio-hazard warning not less than 50mm diameter, printed in black or red on each of the front and back faces of the box. | [Please provide photos showing bio-hazard marking.]
4.2.9 | Fill line
| The maximum recommended fill line must be clearly marked on all vertical faces of the box, in black or red. | [Please provide photos showing maximum fill line.]
4.2.10 | Resistance to penetration
| The average of forces needed to penetrate samples taken from each position must not be less than 15 N, and the minimum force required to penetrate any sample taken from any position must not be less than 12.5 N. | [Please refer to document or test report showing penetration resistance.]
4.2.11 | Resistance to damage during drops from height
| Boxes must pass the drop test described in E10/SB01-VP. After 100 drops, no syringe should have fallen out of the container; the box should not be seriously damaged, and no more than one needle should have penetrated the container walls. | [Please refer to document or test report showing resistance to damage from being dropped.]
4.2.12 | Stability and spillage
| Boxes must not tip over when placed on a 15 degree non-slip plane with its short axis parallel to the line of tilt in general accordance with the test method in AS 4031:1992, Appendix D. If overturning occurs, the arrangement of the sharps aperture should minimize the risk that sharps are spilled. | [Please refer to document or test report showing compliance to this requirement.]
4.3 | Environmental requirements
4.3.1 | Temperature resistance
| Cardboard boxes, filled to their maximum capacity, must be able to resist temperatures of up to 170°C for periods up to 30 minutes without spilling any part of the load. | [Please refer to document or test report showing compliance to this requirement.]
4.3.2 | Water resistance
| Boxes, filled to their maximum capacity, must be able to withstand 48 hours at 43°C and 90% relative humidity in 5 mm of water, without spilling any part of the load. | [Please refer to document or test report showing compliance to this requirement.]
4.4 | Physical characteristics
4.4.1 | Overall dimensions
| Assembled box dimensions should be selected to accommodate the full range of sharps, as defined in clause 3, and to allow effective filling of the box. | [Please state box dimensions and describe or how the box meets the specification requirements.]
4.4.2 | Minimum dimensions
| The minimum height from the bottom of the container to the fill line must be no less than 150mm for 2.5 litre boxes and 230mm for other sizes. | [Please state this dimension.]
4.4.3 | Sharps aperture dimensions
| 38 mm diameter, or 38mm width and breadth. Larger apertures are allowed, but must be fitted with an engineered protective feature – for example a flange on a plastic safety box. | [Please state aperture dimension.]
4.4.4 | Weight
| No specific restriction, consistent with keeping shipping weight to a minimum. | [Please state box weight.]
4.5 | Interface requirements
| External dimensions should be chosen to allow the box to fit within the treatment loading mechanisms. | [Please describe how box dimensions meet this requirement.]
4.6 | Human factors
4.6.1 | Sharps aperture marking
| The aperture must be clearly visible against the colour of the container. | [Please describe or provide photo showing aperture marking.]
4.6.2 | Tamper-proofing
| To reduce the risk of needlestick injury, the lowest point of the sharps aperture must be at least 50 mm above the maximum recommended fill line marked on the exterior of the box. | [Please describe, provide photos, or refer to document that confirms compliance with this requirement.]
4.6.3 | Handling
| It must be possible to carry the box in one hand without spillage of contents and without risk of needle stick injury, both before and after final closure of the sharps aperture. | [Please describe, provide photos, or refer to document that confirms compliance with this requirement.]
4.7 | Materials
| The following materials are permitted: • Bio-degradable cardboard-based materials – post-consumer recycled material is preferred; • Other bio-degradable board materials. • Non-toxic inks, glues and dyes. • Hard recyclable plastic (plastic containers should not be incinerated). • Metal. | [Please state box materials.]
| If incineration is the final treatment option, the following materials are not permitted: • Materials which are not bio-degradable. • Materials which emit ozone depleting substances as defined in the Montreal Protocol; • Materials which generate toxic emissions during incineration at any temperature between 650°C and 1,200°C; • Materials which release gases with a high global warming potential. | [Please state if box is compliant with this requirement.]
4.8 | Warranty
| 100% of boxes are to remain physically intact and satisfactory for use when used in compliance with this performance specification. | [Please refer to manufacturer's warranty statement.]
4.9 | Servicing provision
| The product is a consumable item with no maintenance requirement. | [Please describe how device meets this requirement.]
4.10 | Disposal and recycling
| Boxes are disposed of after a single use cycle if made of cardboard. If made of plastic or metal, they are typically taken to a treatment site to be reused, recycled or disposed of. | [Please state how box should be properly disposed of after final use.]
4.11 | Instructions
| In addition to the international bio-hazard symbol, clear pictorial instructions without writing are to be printed on two sides of the container showing: • How to assemble the box. • How to use the box as a container for contaminated sharps; • Syringe disposal direction (needle down). • How to close the sharps aperture when the box is full. | [Please provide photos showing compliance to this requirement.]
4.12 | Training
Printed user-operations and maintenance instructions specifically directed at the health
5. | centre staff shall be pictorial. All key information is to be summarized on a single
pictogram sheet that may be fixed onto the appliance. The sheet is to be sufficiently
| Training on the assembly, use and disposal of safety boxes will be provided by the health care programme when the boxes are first introduced, and subsequently during supervisory visits | [Although not required, please state if there is any training available (videos, online, etc.).]
4.13 | Verification
| In accordance with PQS Verification Protocol E10/SB01-VP | [Please refer to the test report detailing results in accordance with WHO/PQS/E10/SB01-VP and see "VP Sheet" in this submission template.]
5.0 | Packaging
| Recyclable cardboard is to be used. | [Please state how the packing meets this requirement and refer to recommendations for correct disposal.]

## VP
E010 Submission to PQS for approval: Safety box for the disposal of used sharps
| Compliance with PQS specification: | E010/SB01.1 | Brand & Model:
| As shown by testing in accordance with PQS verification protocol: | E010/SB01-VP.1
Instructions in [square brackets] are for your guidance. Please delete these instructions and replace with evidence or a precise reference (including page and section numbers) showing compliance with that particular PQS specification requirement. In general a "pass" or "complies" statement alone is not acceptable.
One document may provide evidence for several clause requirements. Please carefully reference the sub-sections of any such document so that it is clear that the clause requirement has been met. This will help speed the pre-qualification process.
Greyed-out cells do not require a response.
NOTE: If it is not clear or doubtful whether a requirement has been met, then this is a non-compliance and pre-qualification might be delayed.
Please ensure the Spec sheet is completed.
Section 7. | VP Clause | Applicant's submission | Assessor's remarks [Name & Date] | Applicants response [Name & Date] | Assessor's remarks [Name & Date] | Applicants response [Name & Date] | Assessor's remarks [Name & Date]
5.1 | Evidence of conformity assessment
| Products may carry the CE mark and/or equivalent internationally accepted evidence of conformity assessment, but this is not a mandatory requirement. | [Please state if product carries the CE or other mark.]
5.2 | Number of samples
| The Legal Manufacturer or Reseller must supply the testing laboratory with a full duplicate set of the Product Dossier already supplied to WHO in accordance with the requirements of specification clause 7. A minimum of ten sample(s) of theproduct are required. If the product is available in more than one of the sizes described in specification clause 4.1, provide ten sample(s) of each size. Provide one sample of the packaging described in specification clause 5. | [Please provide the testing laboratory a draft dossier for checking and amendment. This will help to ensure that the dossier is complete.]
5.3 | Test procedure
5.3.1 | Test 1: Type examination
| The test report must clearly show that the requirements of each section and sub-section in the type examination are individually and clearly met. | [Please refer to the test report and see the specification sheet in this template where precise references to each specification requirements are required. (Test 1 requirements refer to many specification clause requirements. See E010/SB01.1.)]
5.3.2 | Test 2: Ease of assembly and filling
Number of samples: Two. Assessment panel: Three people should carry out this assessment independently. • Step 1: Assemble the box using the instructions printed on the pictorial instructions provided. Record the time taken and comment on the operation. • Step 2: Hold uncapped AD syringes by their finger holds, with needles pointing downwards. Drop 20 syringes per nominal litre of container capacity, into the box. If syringes ‘stack’, shake the container gently. • Step 3: Continue to load syringes as described in Step 2 until the fill line is reached or when the insertion of an additional syringe requires heavy shaking or other special attention. Close the sharps aperture. • Step 4: Record the mean number of syringes accommodated by the two test samples. Report this figure as the maximum capacity.
| • Acceptance criterion: The box must be easy to assemble using the pictorial instructions only. The maximum capacity must equal or exceed 20 syringes per nominal litre of container capacity. No needles must penetrate the container and there must be no other visible damage or distortion to the container. The closure mechanism must work correctly when the box is filled to its maximum capacity. • Rejection criterion: Significant difficulty in assembling the container and/or failure to accommodate at least 20 syringes per nominal litre of capacity and/or needle penetration or other visible damage or distortion and/or failure of the closure mechanism to work correctly. | [Please refer to a precise section in the test report showing the results of the test.]
5.3.3 | Test 3: Needle penetration test
Number of samples: Sufficient boxes to prepare the 12mm x 12mm test samples required by BS 7320: 1990, Appendix C. • Step 1: Follow the method of the test describe in BS 7320: 1990, Appendix C, except that the hypodermic needles used should be 23 gauge x 25mm needles as fitted to AD syringes to ISO 7886 - part 3. Record the measured force for each penetration.
| • Acceptance criteria: The average of forces needed to penetrate samples taken from each position must not be less than 15 N, and the minimum force required to penetrate any sample taken from any position must not be less than 12.5 N. • Rejection criterion: Failure to meet either or both of the acceptance criteria. | [Please refer to a precise section in the test report showing the results of the test.]
5.3.4 | Test 4: Drop test
Number of samples: One. • Step 1: Assemble and load the box to its maximum capacity as described in Test 2. Close and seal the aperture in accordance with the pictorial instructions on the box. • Step 2: Place the loaded container in a large tumble box arranged so that each drop is from a height of 800mm. Inspect the box after 5, 10, 50 and 100 drops and record signs of deterioration or needle piercing. Stop the test if inspection shows that more than one needle has penetrated.
| • Acceptance criteria: After 100 drops, no syringe should have fallen out of the container; the box should not be seriously damaged, and no more than one needle should have penetrated the container walls. • Rejection criterion: Failure to meet any of the acceptance criteria. | [Please refer to a precise section in the test report showing the results of the test.]
5.3.5 | Test 5: Stability and spillage test
Number of samples: One. • Step 1: Prepare an adjustable plane tilt-test apparatus as described in AS 4031-1992, Appendix D. • Step 2: Assemble the box. Do not close the sharps aperture in any of the subsequent tests. • Step 3: Set the tilting plane at 15° to the horizontal. Place the box on the sloping plane so that its short axis is parallel to the line of tilt and the sharps aperture is on the downward side of the slope3. Note whether the box remains standing or whether it topples. • Step 4: Fill the container to half its nominal capacity with 0.5ml AD syringes, adopting the procedure described in 5.3.2, Step 2. Repeat the Step 3 test. • Step 5: Increase the angle of the plane until the box just topples. Note the toppling angle and record whether any syringes spill or protrude from the box. Carry out this procedure a total of 10 times, gently shaking the box between each test to re-centre the contents. • Step 6: Fill the container to its fill line with 0.5ml syringes, again adopting the procedure described in 5.3.2, Step 2. Repeat the Step 3 test. • Step 7: Repeat the Step 5 test a further 10 times.
| • Acceptance criteria: The box must not topple in any of the three trials at 15°. No part of any syringe must spill or protrude from the sharps aperture after any of the Step 5 or Step 7 trials. • Rejection criteria: Failure to meet either or both of the acceptance criteria. | [Please refer to a precise section in the test report showing the results of the test.]
5.3.6 | Test 6: Temperature resistance test
Applicability: Cardboard boxes only. Number of samples: One. • Step 1: Assemble and load the box as described Test 2. • Step 2: Place the loaded container in a test chamber at 170°C for 30 minutes. • Step 3: Remove and inspect the container.
| • Acceptance criteria: No spillage of the contents. Distortion of the container is acceptable. • Rejection criteria: Spillage of contents. | [Please refer to a precise section in the test report showing the results of the test.]
5.3.7 | Test 7: Water resistance test
Number of samples: One. • Step 1: Assemble and load the box to its maximum capacity as described in Test 2. Close and seal the aperture in accordance with the pictorial instructions on the box. • Step 2: Set a test chamber temperature to 43°C and 90% relative humidity. Place the loaded box in a tray of water 5mm deep for 48 hours. Maintain the water level throughout this period. • Step 3: Remove the box from the water tray and record its condition. • Step 4: Shake the box strongly 20 times, using an up and down motion. Record any spillage or needle penetration.
| • Acceptance criteria: No spillage of contents or needle penetration. • Rejection criterion: Failure to meet either or both of the acceptance criteria. | [Please refer to a precise section in the test report showing the results of the test.]
5.3.8 | Test 8: Handle strength test
Number of samples: One. • Step 1: Carry out a test for the security and strength of the handles in accordance with BS 7320:1990, Appendix A.
| • Acceptance criteria: Handle(s) retain integrity. No evidence of separation of handle(s) from the container. | [Please refer to a precise section in the test report showing the results of the test.]
5.4.9 | Test 9: Aperture closure test
Number of samples: One. • Step 1: Carry out a test for security and attachment of the aperture closure device in accordance with BS 7320:1990, Appendix B.
| • Acceptance criteria: No parts detached. | [Please refer to a precise section in the test report showing the results of the test.]