## Checklist
| PQS performance verification protocol Tool |  |  |  | WHO/PQS/E006/TR03-VPT.2 Original: English Distribution: General
TITLE: Programmable remote temperature and event monitoring systems
| Product Verification protocol Tool: | E006/TR03-VPT.2
| Product Verification protocol: | E006/TR03-VP.2
| Applies to specification ref(s): | E006/TR03.2
| Date of origin: | 2015-09-02
| Date of last revision: | 2015-09-02
Name of Supplier: |  |  |  |  | Result of Visit:
Name of System:
Date of visit:
Conforms to Specifications?
Category (Specification clause) | Sub-Category | Requirement | Yes | No | Comments:
System characteristics | Generic description of system including minimum system setup and details of expansion options | Description of system
| List of submitted system components, including software | List of components
Inclusion of identification of system components | Code (unique identifier) | Code
| Brand/model | Brand/model numbers
| Legal Manufacturer or Reseller | Name of Legal Manufacturer or Reseller
| List of OEMs of all sourced components | List of OEMs
| Mode of operation | E.g., cellular or satellite or Bluetooth
| Country of origin | Country of origin of Legal Manufacturer or Reseller
| Conformity assessment markings | List of markings
Hardware
Performance characteristic of sensors (4.1.1) | Temperature sensor operating range | Upper limit: +50 C, Lower limit: -30 C
| Temperature sensor accuracy | ±0.5 C or better, within the range -30 C to +20 C
| Temperature sensor resolution | ±0.2 C or better, within the range -30 C to +20 C
| Temperature sensor response time | T90 within 20 minutes maximum
| Temperature sensor type | Electronic
| "Door-open" event sensor operating parameters | "Door-open" event must be identified when the door panel is not fully seated in the closed position. User-programmable delay function for event
| Voltage sensor operating parameters | To monitor all combinations of single phase or three phase AC voltage and frequency
| Sensor lead length | Can be supplied as required by procuring agent
| Sensor fixings | Supplied with fixings suitable for permanent attachment to inside skin of a cold room or freezer room or to the inside skin of a vaccine refrigerator or freezer
| Sensor IP rating | Not less than IP64
Performance characteristic of logging units, where logging units are a part of the system (4.1.2) | Logging unit operating parameters | Can incorporate integrated temperature sensors or can be connected to external sensors
| Logging unit power source | EITHER rechargeable, replaceable battery backup with 3 year operational life and the following charge capacities: 2 days for Segment 1, 5 days for Segment 2, and 7 days for Segment 3 OR rechargeable, non-replaceable battery with minimum 7 years operational life and minimum 3 year warranty on battery and same charge capacities as above OR non-rechargeable, non-replaceable battery with minimum 7 years operational life and minimum 3 years warranty on battery
| Logging unit memory | EPROM or equivalent non-volatile solid-state memory device capable of storing a minimum of 30 days worth of data for each sensor, assuming a 10-minute logging interval
| Logging interval | User-programmable logging interval, starting with a minimum interval of ten minutes
| Logging unit casing | Non-corrodible and robust material (e.g., plastics or treated metal)
| Logging unit fixings | Supplied with fixings suitable for permanent attachment to inside skin of a cold room or freezer room or to the inside skin of a vaccine refrigerator or freezer
| Logging unit IP rating | Not less than IP64
| Connection to base station | Data transmission to the base station established automatically, without requiring any manual adjustment or fine-tuning other than plugging in sensors and cables, as necessary. If radio frequency links are used, must be in accordance with ETSI EN 300-220 or the ZigBee standard. In the event of interruption of power supply, the logging unit must be able to automatically reconnect with the base station once power supply is resumed
Performance characteristic of base station (4.1.3) | Base station channels | Minimum number of temperature sensors to be supported by each base station (including the sensors from the logging units) must be: 8 sensors for Segment 1, 4 sensors for Segment 2, 2 sensors for Segment 3 For Segment 1, at least 2 door sensors must also be supported in addition to temperature sensors
| Base station memory | EPROM or equivalent non-volatile solid-state memory capable of storing a minimum of 30 days worth of sensor data from all connected channels, assuming a 10-minute logging interval Configured to overwrite data cyclically on a first-in-first-out basis when the memory is full In case of trouble with transmitting data to the server, the base station must store all data internally. Upon resumption of connectivity, all stored data must be automatically transmitted to the server
| Base station power source | EITHER rechargeable, replaceable battery backup with 3 year operational life and the following charge capacities: 2 days for Segment 1, 5 days for Segment 2, and 7 days for Segment 3 OR rechargeable, non-replaceable battery with minimum 7 years operational life and minimum 3 year warranty on battery and same charge capacities as above OR non-rechargeable, non-replaceable battery with minimum 7 years operational life and minimum 3 years warranty on battery
| Base station data connection capability | Must include at least USB connection capability for downloading data to a PC
| Base station IP rating | Not less than IP50
Performance characteristic of audio-visual alarm and response/mute button (4.1.4) | Alarm and mute button location | Must be at the same site as the device, and must be capable of being positioned in an easily reachable location for the user
| Alarm sound intensity | 70dB(A) at a distance of one meter from the sounder. The pattern of the signal is to be an intermittent pulse
| Alarm light intensity | Flashing red colored light rated 100 mcd diffused
| Alarm mode of operation | Audio-visual alarm to be triggered whenever an alarm event occurs for a preset time period Alarm is to be canceled by: i) physically pressing "mute" button, ii) sending a remote response (e.g., through SMS) to the base station, or iii) through the web portal. Each of these responses is to be recorded by the system and used to calculate "time of user response", which is to be transmitted to the server and displayed in the web portal
Performance characteristic of modem (4.1.5) | Modem power source | Must utilize same power source as base station
| Modem functionality | Transmit data periodically from the base station to the server Receive any remote updates, setting changes, or commands Issue immediate alerts when alarms are triggered. E.g., by sending SMSes to pre-programmed mobile phone numbers or audio calling and leaving a recorded voice message. These alerts may be issued either directly by the modem or indirectly through the server Escalate the alert to higher-level decision-makers (e.g., through SMSes, audio calls, or emails) if the first alert is not responded to within a specified time limit
| Modem connectivity | Due to likely issues with connectivity (e.g., cellular connectivity issues), remote alerts and data transmission to server must automatically be possible through basic GSM standards. When available, higher-speed and more cost-efficient transmission (e.g., 3G) must be automatically utilized
Performance characteristic of display (4.1.6) | Display information | Must display at least the following information: a) That the device is activated, b) Battery charge level indicator OR battery replacement warning indicator activated when no less than 20% of battery life remains, c) The most recently read temperature for each sensor connected directly or indirectly to the base station, d) Whether or not the component is in an alarm condition, e) Date, time, and cellular signal strength are optional information that could be included
Performance characteristic of calibration (4.1.7) | System component calibration certificate(s) | System components are to be covered by a Certificate of Traceability and Calibration. Producer-issued calibration certificates are acceptable along with a general NIST-traceable certificate for the series of logging units/base stations. However, in the case of producer-issued certification, the testing laboratory must be accredited with ISO/IEC 17025
Performance characteristic of power leads (4.1.8) | Power leads | Must be supplied with a factory-sealed plug connector to suit the mains electrical socket standard in the country of intended use
Performance characteristic of circuit design for electromagnetic compatibility (4.1.9) | Circuit design | Operation of the system and of the individual system components must be unaffected in the normal electromagnetic compatibility environment in which the system is intended to work
Software
Performance characteristic of front-end and back-end software (4.1.10) | “Plug-and-play” functionality | Base station must automatically connect and be recognized by server when switched on. No user actions must be required to enable this, beyond the physical connections and switching on In case of interruption of power supply or cellular connectivity, base station must automatically reconnect and be recognized by the server once power supply or cellular connectivity is restored. Upon reconnection, the base station must be able to transmit to the server the data that was collected during the interruption
| Remote configuration capability | System administrator must be able to remotely control and change the following settings through the web portal: a) Alarm thresholds, b) Measurement and transmission frequencies of the base station, c) Alarm recipients and their details (phone numbers, email addresses, etc.), d) Escalation tree for alarm, e) Recipients of periodic emailed reports on the health of the cold chain, f) Format of report User-interface must be highly visual and requiring minimal learning and effort (e.g., few mouse “clicks”) to navigate
| Alarm threshold | Must allow different alarm settings for various types of excursions, with settings to be made by the system administrator. For example, allow 10 consecutive hours of excursions above 8 °C before alarming, allow 1 hour of excursions below -0.5 °C before alarming, immediately alarm if excursion reaches 40 °C
| Alarm escalation features | Minimum 3 levels of alarm escalations: 1) Alarm sent to first recipients, 2) If first recipients do not respond within specified time period, alarm escalated to second recipients, 3) If second recipients do not respond within specified time period, alarm escalated to third recipients System administrator must be able to remotely specify (through the web portal) recipients at each level of this escalation tree, their preferred form of alarm (e.g., SMS, voice call, email), and the allowed time of response before the next escalation
| Action-oriented alarm features | Alarm alerts (e.g., SMSes, voice calls, and emails) must be directly actionable. For example: “Refrigerator number X has been experiencing a freezing excursion for 5 hours.”
| “Grouping” of sensor capability | All sensors in the same appliance must be able to be “grouped” together, so that the system is aware that the same appliance is being monitored
| Unit of measurement | Must be recorded in degrees Celsius. Calculation and display of mean kinetic temperature (MKT) must also be possible
Performance characteristic of web portal (4.1.11) | Multiple layers of access privileges | Web portal must allow system administrator to set different access privileges and read/write privileges for various users (at least 3 levels of access privileges, customizable by the system administrator) For example, system administrators at the National level must have full privileges for the entire country, while system administrators at the State level must have full privileges, but restricted only to their State. Some officials must have read/write privileges, while others must only have read privileges
| Web page and map view for cold chain status | Web portal must have a page to provide the remote user with a quick current status of all the equipment in their area of focus. A real-time map view might be available to present this visually. All appliances in alarm state must be highlighted in an obvious manner
| Web page for cold chain report | Web portal must have a page to provide the remote user with a summary report of the performance of the cold chain during a selected time period The user must be able to select the required time period by day, week, month, and year. For the selected time period, the report must show the total time durations of heating and freezing excursions of each appliance. The sites with the greatest time durations of excursions must be displayed at the top of the list, so that they are immediately visible
| Periodic emailed report capability | On a monthly and annual basis, simple, visual, and action-oriented report must be emailed to relevant officials, as specified by system administrators. Suggested information to summarize in the report should include (but not be limited to): a) Since the last report, the total number of heating excursions above 8 °C for longer than 10 consecutive hours, b) Since the last report, the total number of freezing excursions below -0.5 °C for longer than 1 hour, c) Since the last report, a list of all the monitored cold chain equipment, in decreasing order of number of heating and freezing excursions (equipment with the most excursions on top, equipment with the least excursions at the bottom)
| Raw data and graphs capability | User must be able to easily access raw data from the server, and to develop graphs of the data over selected time periods
| Exporting files capability | Web portal must allow data to be exported in Excel/CSV, ASCII, and XML formats and images to be exported in JPEG and PDF formats
| Security features | Authentication and access control of user and privileges standard not less stringent than US 21 CFR Part 11
| Shareable application programming interface (API) | Supplier must be capable of sharing API with procuring agency, if this is a requirement
Robustness to environmental conditions (4.2.1 to 4.2.3) | Robustness to ambient temperature range | From -30°C to +55°C with the system components inactivated
| Robustness to ambient humidity range | 0 to 95% RH
| Robustness to voltage fluctuations | If requested by the procuring agency, the Reseller or Legal Manufacturer must be able to source and supply PQS-approved voltage stabilizers
Interface requirements (4.4.1) | Software compatibility | List of compatible programs must include all releases of Microsoft Excel currently supported by Microsoft Software must be compatible with all Microsoft PC operating systems currently supported by Microsoft
Materials and construction (4.5.1 to 4.5.3) | Ozone depleting chemicals | During manufacture and assembly of the printed circuit boards and final assembly of the product do not use any substance included in Annex A, B or C of the Montreal Protocol
| Other restricted materials | The product and its constituent components, including batteries, must not contain lead, mercury, cadmium, hexavalent chromium, polybrominated biphenyls (PBB) or polybrominated biphenyl ethers (PBDE)
| Packaging materials | Free of ozone-depleting compounds as defined in the Montreal Protocol
Warranty (4.6) | Warranty | All system components are to be covered by a three year warranty in the event of any component failure. On-site repair and maintenance is to be offered as an option to the purchasing party
Servicing provision (4.7) | Servicing provision | System is to be maintenance-free, apart from routine battery or sensor replacement and re-calibration
Disposal and recycling (4.8) | Disposal and recycling | Manufacturer is to provide information to the buyer on the hazardous materials contained within the system and suggestions for resource recovery/recycling and/or environmentally safe disposal
Instructions (4.9) | Manuals for users and technician installation | Each unit of the hardware must include a separate user manual and technician’s installation manual in the language most appropriate to the installation site. English version must be inspected by Evaluator for PQS approval
Training (4.10) | Training | Equipment supplier should provide the necessary training materials to the installer, and should also be able to provide in-person training to installers and system administrators