## Didier Fr (hidden)
Output: WHO Core PMM Indicators |  |  | Input: Country data collection
Indicator | Definition | Data elements required | ODK Question | ODK Question | ODK Answer |  | ODK Answer | notes
Functional status | Functioning = operating between 2-8 degrees during period. Device functionality categorization as per inventory and subsequent data analysis (metadata) | NUMERATOR: total number of functional (working) monitored devices (brand/model) DENOMINATOR: total number of monitored devices (brand/model) | Is the refrigerator working properly? | Le réfrigérateur fonctionne-t-il correctement ? | Oui | Non | Working/ Non functional | Monthly data collection from health facility. Surveillance Officer inputs data into ODK2 data collection tool.
Frequency of heat alarms per month | Heat alarm: 10 consecutive hours of continuous temperature excursions above +8°C | NUMERATOR: Total number of heat alarm events recorded in a month (brand/model) DENOMINATOR: N/A | How many heat alarms with over 10 hour duration above 8°C show in the temperature log for the past month? | Combien d'alarme haute température (à plus de 10°C) a-t-il été enregistrée au cours du dernier mois ? | 1 à 60 |  | Numeric response | Monthly data collection from health facility. Surveillance Officer inputs data into ODK2 data collection tool.
Frequency of freeze alarms per month | Freeze alarm: 1 hour of continuous temperature excursions below -0.5°C | NUMERATOR: Total number of freeze alarm events recorded in a month (brand/model) DENOMINATOR: N/A | How many freeze alarms with over 1 hour duration below -0.5°C show in the temperature log for the past month? | Combien d'alarme basse température (inférieure à -O,5°C) a-t-il été enregistrée au cours du dernier mois ? | 1 à 60 |  | Numeric response | Monthly data collection from health facility. Surveillance Officer inputs data into ODK2 data collection tool.
Frequency of heat alarms with duration of 48h or more |  | NUMERATOR: Total number of heat alarm events with duration of more than 48h recorded in a month (brand/model) DENOMINATOR: N/A | How many heat alarms with 48h or longer duration show in the temperature log for the past month? | Combien d'alarmes ayant duré 48h ou plus ont été enregistré au cours du dernier mois ? | 0 à 15 |  | Numeric response | Monthly data collection from health facility. Surveillance Officer inputs data into ODK2 data collection tool.
Failure Type | Diagnosed root cause of failure by component (e.g. fried thermostat, damaged compressor, leaking evaporate line, etc.) | NUMERATOR: Total number of failure types identified (brand/model) DENOMINATOR: Total number of deployed devices evaluated (brand/model) | Why did the refrigerator fail? | Pourqoui le réfrigérateur ne fonctionne t-il pas ? |  |  | Drop down + other with free text | Triggered when refrigerator is non functional. Surveillance Officer does root cause of failure analysis using ODK2 data collection tool.
Build quality | Issues related to corrosion, component damage, degradation, etc. that negatively effects the performance or useful lifetime of the device | NUMERATOR: Total number of build quality issues reported (brand/model) DENOMINATOR: N/A | Have there been any issues related to build quality in the past month? | Y a-t-il eu un problème lié au défaut de fabrication du réfrigérateur ? | Oui | Non | 1) Yes/No 2) if yes drop down list + other and free text | Monthly data collection from health facility. Surveillance Officer inputs data into ODK2 data collection tool.
Si oui,les quels
Voltage stabilizer quality | Issues related to voltage stabilizer failure | NUMERATOR: Total number of voltage stabilizer failures (brand/model) DENOMINATOR: Total number of voltage stabilizers deployed (brand/model) | 1) Is the voltage stabilizer working? 2) Has the voltage stabiliser been replaced in the past month? | Le stabilisateur de tension fonctionne t-il ? | Oui | Non | 1) Yes/ No, 2) if no, why? drop down list + other and free text 1) Yes/No 2) If yes, why? drop down list + other and free text | Monthly data collection from health facility (only where a voltage stabiliser is used). Surveillance Officer inputs data into ODK2 data collection tool.
Si Non,les causes
|  |  |  | Le satbilisateur de tension a-t-il éré remplacé au cours du dernier mois ? | Oui | Non
Si oui,pour quelles raisons
Warranty claims | # of warranty claims | NUMERATOR: # of warranty claims (brand/model) DENOMINATOR: total number of monitored devices under warranty (brand/model) | Has a warranty claim been made? | L'équipement est-il sous garantie du fabricant ? | Oui | Non | Yes/ No | Monthly data collection from health facility (only where the refrigerator is under warranty). Surveillance Officer inputs data into ODK2 data collection tool.
Time to repair | Product performance, quality or safety defects or failures that qualify for warranty coverage | NUM: end time of non-functional status minus start time of warranty claim. DEN: N/A' | Number of days from warranty claim until repaired | La garantie couvre jusqu'à quelle période ? |  |  | Numeric calculation | Data collection tool to automatically calculate the time from warranty claim to repair.
Spare parts consumption | # of spare parts consumed per month | NUMERATOR: Total number of specific spare parts consumed (brand/model) DENOMINATOR: N/A | How many spare parts have been consumed in the past month? | Combien de pièces de rechange ont été consommées au cours du dernier mois? |  |  | Count of parts from taxonomy/list | Monthly data collection from health facility. Surveillance Officer inputs data into ODK2 data collection tool.
Note on C11: Mahunde, Didier: The denominator can't be the number of refrigerator with voltage stabiliszers deployed ?

## English
WHO/PQT/VAX/IMD/PMM Core Indicators V1.0 |  |  |  | 2024-06-01
0, avenue Appia – CH-1211 Geneva 27 – Switzerland – Tel central +41 22 791 2111 – Fax central +41 22 791 3111 www.who.int
PMM Core Indicators
Output: WHO Core PMM Indicators |  |  | Input: Country data collection
Indicator | Definition | Data elements required | ODK Question | ODK Answer | notes
Functional status | Functioning = operating between 2-8 degrees during period. Device functionality categorization as per inventory and subsequent data analysis (metadata) | NUMERATOR: total number of functional (working) monitored devices (brand/model) DENOMINATOR: total number of monitored devices (brand/model) | Is the refrigerator working properly? | Working/ Non functional | Monthly data collection from health facility. Surveillance Officer inputs data into ODK2 data collection tool.
Frequency of heat alarms per month | Heat alarm: 10 consecutive hours of continuous temperature excursions above +8°C | NUMERATOR: Total number of heat alarm events recorded in a month (brand/model) DENOMINATOR: N/A | How many heat alarms with over 10 hour duration above 8°C show in the temperature log for the past month? | Numeric response | Monthly data collection from health facility. Surveillance Officer inputs data into ODK2 data collection tool.
Frequency of freeze alarms per month | Freeze alarm: 1 hour of continuous temperature excursions below -0.5°C | NUMERATOR: Total number of freeze alarm events recorded in a month (brand/model) DENOMINATOR: N/A | How many freeze alarms with over 1 hour duration below -0.5°C show in the temperature log for the past month? | Numeric response | Monthly data collection from health facility. Surveillance Officer inputs data into ODK2 data collection tool.
Frequency of heat alarms with duration of 48h or more |  | NUMERATOR: Total number of heat alarm events with duration of more than 48h recorded in a month (brand/model) DENOMINATOR: N/A | How many heat alarms with 48h or longer duration show in the temperature log for the past month? | Numeric response | Monthly data collection from health facility. Surveillance Officer inputs data into ODK2 data collection tool.
Failure Type | Diagnosed root cause of failure by component (e.g. fried thermostat, damaged compressor, leaking evaporate line, etc.) | NUMERATOR: Total number of failure types identified (brand/model) DENOMINATOR: Total number of deployed devices evaluated (brand/model) | Why did the refrigerator fail? | Drop down + other with free text | Triggered when refrigerator is non functional. Surveillance Officer does root cause of failure analysis using ODK2 data collection tool.
Build quality | Issues related to corrosion, component damage, degradation, etc. that negatively effects the performance or useful lifetime of the device | NUMERATOR: Total number of build quality issues reported (brand/model) DENOMINATOR: N/A | Have there been any issues related to build quality in the past month? | 1) Yes/No 2) if yes drop down list + other and free text | Monthly data collection from health facility. Surveillance Officer inputs data into ODK2 data collection tool.
Voltage stabilizer quality | Issues related to voltage stabilizer failure | NUMERATOR: Total number of voltage stabilizer failures (brand/model) DENOMINATOR: Total number of voltage stabilizers deployed (brand/model) | 1) Is the voltage stabilizer working? 2) Has the voltage stabiliser been replaced in the past month? | 1) Yes/ No, 2) if no, why? drop down list + other and free text 1) Yes/No 2) If yes, why? drop down list + other and free text | Monthly data collection from health facility (only where a voltage stabiliser is used). Surveillance Officer inputs data into ODK2 data collection tool.
Warranty claims | # of warranty claims | NUMERATOR: # of warranty claims (brand/model) DENOMINATOR: total number of monitored devices under warranty (brand/model) | Has a warranty claim been made? | Yes/ No | Monthly data collection from health facility (only where the refrigerator is under warranty). Surveillance Officer inputs data into ODK2 data collection tool.
Time to repair | Product performance, quality or safety defects or failures that qualify for warranty coverage | NUM: end time of non-functional status minus start time of warranty claim. DEN: N/A' | # of days from warranty claim until repaired | Numeric calculation | Data collection tool to automatically calculate the time from warranty claim to repair.
Spare parts consumption | # of spare parts consumed per month | NUMERATOR: Total number of specific spare parts consumed (brand/model) DENOMINATOR: N/A | How many spare parts have been consumed in the past month? | Count of parts from taxonomy/list | Monthly data collection from health facility. Surveillance Officer inputs data into ODK2 data collection tool.
Note on C12: Mahunde, Didier: The denominator can't be the number of refrigerator with voltage stabiliszers deployed ?

## Didier (hidden)
Output: WHO Core PMM Indicators |  |  | Input: Country data collection
Indicator | Definition | Data elements required | ODK Question | ODK Answer | notes
Functional status | Functioning = operating between 2-8 degrees during period. Device functionality categorization as per inventory and subsequent data analysis (metadata) | NUMERATOR: total number of functional (working) monitored devices (brand/model) DENOMINATOR: total number of monitored devices (brand/model) | Is the refrigerator working properly? | Working/ Non functional | Monthly data collection from health facility. Surveillance Officer inputs data into ODK2 data collection tool.
Frequency of heat alarms per month | Heat alarm: 10 consecutive hours of continuous temperature excursions above +8°C | NUMERATOR: Total number of heat alarm events recorded in a month (brand/model) DENOMINATOR: N/A | How many heat alarms with over 10 hour duration above 8°C show in the temperature log for the past month? | Numeric response | Monthly data collection from health facility. Surveillance Officer inputs data into ODK2 data collection tool.
Frequency of freeze alarms per month | Freeze alarm: 1 hour of continuous temperature excursions below -0.5°C | NUMERATOR: Total number of freeze alarm events recorded in a month (brand/model) DENOMINATOR: N/A | How many freeze alarms with over 1 hour duration below -0.5°C show in the temperature log for the past month? | Numeric response | Monthly data collection from health facility. Surveillance Officer inputs data into ODK2 data collection tool.
Frequency of heat alarms with duration of 48h or more |  | NUMERATOR: Total number of heat alarm events with duration of more than 48h recorded in a month (brand/model) DENOMINATOR: N/A | How many heat alarms with 48h or longer duration show in the temperature log for the past month? | Numeric response | Monthly data collection from health facility. Surveillance Officer inputs data into ODK2 data collection tool.
Failure Type | Diagnosed root cause of failure by component (e.g. fried thermostat, damaged compressor, leaking evaporate line, etc.) | NUMERATOR: Total number of failure types identified (brand/model) DENOMINATOR: Total number of deployed devices evaluated (brand/model) | Why did the refrigerator fail? | Drop down + other with free text | Triggered when refrigerator is non functional. Surveillance Officer does root cause of failure analysis using ODK2 data collection tool.
Build quality | Issues related to corrosion, component damage, degradation, etc. that negatively effects the performance or useful lifetime of the device | NUMERATOR: Total number of build quality issues reported (brand/model) DENOMINATOR: N/A | Have there been any issues related to build quality in the past month? | 1) Yes/No 2) if yes drop down list + other and free text | Monthly data collection from health facility. Surveillance Officer inputs data into ODK2 data collection tool.
Voltage stabilizer quality | Issues related to voltage stabilizer failure | NUMERATOR: Total number of voltage stabilizer failures (brand/model) DENOMINATOR: Total number of voltage stabilizers deployed (brand/model) | 1) Is the voltage stabilizer working? 2) Has the voltage stabiliser been replaced in the past month? | 1) Yes/ No, 2) if no, why? drop down list + other and free text 1) Yes/No 2) If yes, why? drop down list + other and free text | Monthly data collection from health facility (only where a voltage stabiliser is used). Surveillance Officer inputs data into ODK2 data collection tool.
Warranty claims | # of warranty claims | NUMERATOR: # of warranty claims (brand/model) DENOMINATOR: total number of monitored devices under warranty (brand/model) | Has a warranty claim been made? | Yes/ No | Monthly data collection from health facility (only where the refrigerator is under warranty). Surveillance Officer inputs data into ODK2 data collection tool.
Time to repair | Product performance, quality or safety defects or failures that qualify for warranty coverage | NUM: end time of non-functional status minus start time of warranty claim. DEN: N/A' | # of days from warranty claim until repaired | Numeric calculation | Data collection tool to automatically calculate the time from warranty claim to repair.
Spare parts consumption | # of spare parts consumed per month | NUMERATOR: Total number of specific spare parts consumed (brand/model) DENOMINATOR: N/A | How many spare parts have been consumed in the past month? | Count of parts from taxonomy/list | Monthly data collection from health facility. Surveillance Officer inputs data into ODK2 data collection tool.
Nombres d'heures en temperatures anormales haute
Nombres d'heures en temperatures anormales basse
There are somes questions can be collected once for all equipment like us: |  | One collect data for all equipment
L'equipement a-t-il ete fourni avec les cles pour la maintenance ?
Votre equipemet est-fourni avec plus de deux cles pour ourvir la portière ?
La cle de votre equipement peut-il etre utilisé pour ouvrir les autres équipements ?
La temperature,les performances,la duree de l'autonomie,les alarmes sont-ils possibles avec votre équipement
Votre équipement affiche t-il la température interne ?
La lecture ou l'écran est-il lisible facilement sans beaucoup des peines ?
Votre équipement dispose -t-il des roues en dessous qui facilite son deplacement ?
Le numéro de serie est-il facilement visible sans beaucoup trop d'acrobaties ?
En cas de panne ou probleme de frourniture d'enérgie,votre equipement a t-il un écran d'affichage du nombre d'heures restant qu'il gardera la température dans les plages normales recommandées par l'OMS?
Les panniers sont-ils suffisants et adaptés et bien dimensionnés à l'equipement ?
Les panniers ont ils des grands espaces qui font que les flacons des vaccins y rangés tombent facilement au fond du refrigérateur ?
Le manuel ou guide d'utilisation est-il fourni avec l'equipement ?
Y a-t-il une procédure opératoire affichée au réfrigerateur montrant comment assurer la maintenance de l'équipement ?
Note on C9: Mahunde, Didier: The denominator can't be the number of refrigerator with voltage stabiliszers deployed ?

## French
Résultat: Indicateurs de base du PMM de l’OMS |  |  | Entrée: collecte de données par pays
Indicateur | Definition | Éléments de données requis | Question ODK | Réponse ODK | Remarques
Etat de fonctionnalité | Fonctionnement opérationnelle entre 2-8 degrés pendant la période. Catégorisation des fonctionnalités de l'appareil en fonction de l'inventaire et analyse ultérieure des données (métadonnées) | NUMERATEUR: nombre total d'equipements surveillés fonctionnels (en fonctionnement) (marque / modèle) DENOMINATEUR: nombre total d'equipements surveillés (marque / modèle | Le réfrigérateur fonctionne-t-il correctement? | Fonctionnel / non fonctionnel | Collecte mensuelle de données auprès des structures de santé. Le responsable de la surveillance entre les données dans l'outil de collecte de données ODK2.
Fréquence des alarmes superieures (de chaleur) par mois | Alarme de chaleur (alarme haute): 10 heures consécutives d'excursions continue de température haute (supérieure à) à + 8 ° C | NUMERATEUR: Nombre total d'événements d'alarme de chaleur (haute) enregistrés au cours d'un mois (marque / modèle) DENOMINATEUR: N / A | Combien d'alarmes de haute température (supérieure à 8°C et ayant durée 10 heures) enregistré dans le releve de température du mois dernier? | Reponse numerique | Collecte mensuelle de données auprès des structures de santé. Le responsable de la surveillance entre les données dans l'outil de collecte de données ODK2. 5 alarmes de chaleur ou plus par mois = non fonctionnel
Fréquence des alarmes inferieures (de gel) par mois | Alarme gel (alarme basse): 1 heure d’excursion continue en-dessous de -0,5 ° C | NUMERATEUR: Nombre total d’événements d’alarme gel (basse) enregistrés dans un mois (marque / modèle) DENOMINATEUR: N / A | Combien d'alarmes de basse température (-0,5°C ayant dur au moins 1 heure) apparaissent/enregistrées dans le releve de température du mois dernier? | Réponse numérique | Collecte mensuelle de données auprès des structures de santé. Le responsable de la surveillance entre les données dans l'outil de collecte de données ODK2. 1 ou plusieurs alarmes de gel dans un mois = non fonctionnel
Fréquence des alarmes superieures (de chaleur) avec une duree superieure a 48 heures |  | NUMERATOR: Total number of heat alarm events with duration of more than 48h recorded in a month (brand/model) DENOMINATOR: N/A Nombre total des alarmes de chaleurs avec une duree de plus de 48h enregistres dans un mois | Combien d'alarmes de haute température d'une durée de 48h ou plus apparaissent dans le releve de température du mois dernier ? | Réponse numérique | Collecte mensuelle de données auprès des structures de santé. Le responsable de la surveillance entre les données dans l'outil de collecte de données ODK2. 1 ou plusieurs alarmes de chaleur d'une durée de 48h ou plus = non fonctionnel
Type de panne | Diagnostic de la cause fondamentale de défaillance par composant (par exemple, thermostat frit, compresseur endommagé, conduite d’évaporation qui fuit, etc.) | NUMERATEUR: Nombre total de types de défaillance identifiés (marque / modèle) DÉNOMINATEUR: Nombre total d'appareils déployés évalués (marque / modèle) " |  | Liste déroulante + autre texte libre | Déclenché lorsque le réfrigérateur est non fonctionnel. Le responsable de la surveillance effectue l’analyse de la cause fondamentale de la défaillance à l’aide de l’outil de collecte de données ODK2
Qualité de fabrication | Problèmes liés à la corrosion, aux dommages des composantes, à la dégradation, etc. qui affectent négativement la performance ou la durée de vie utile de l'appareil | NUMERATEUR: Nombre total de problèmes de qualité de fabrication signalés (marque / modèle) DENOMINATEUR: N / A | Y at-il eu des problèmes liés à la qualité de la fabrication (ou Défaut de fabrication) au cours du dernier mois? | 1) Oui / Non 2) si oui liste déroulante + autre texte libre " | Collecte mensuelle de données auprès des structures de santé. Le responsable de la surveillance entre les données dans l'outil de collecte de données ODK2.
Qualite du Stabilisateur de tension | Problèmes liés à la défaillance du stabilisateur de tension | NUMERATEUR: Nombre total de défaillances du stabilisateur de tension (marque / modèle) DENOMINATEUR: Nombre total de stabilisateurs de tension déployés (marque / modèle) | 1) Le stabilisateur de tension fonctionne-t-il? 2) Le stabilisateur de tension a-t-il été remplacé le mois dernier? | 1) Oui / Non 2) si oui liste déroulante + autre texte libre " 1) Oui / Non 2) si oui liste déroulante + autre texte libre " | Collecte mensuelle de données auprès des structures de santé (uniquement lorsqu'un stabilisateur de tension est utilisé). Le responsable de la surveillance entre les données dans l'outil de collecte de données ODK2.
Réclamations de garantie | Nombre de réclamations de garantie | NUMÉRATEUR: Nombre de réclamations portant sur la garantie (marque / modèle) DÉNOMINATEUR: nombre total d'appareils surveillés sous garantie (marque / modèle) | Nombre de jours entre la demande de garantie et la réparation | Oui/Non | Collecte mensuelle de données auprès de structures de santé (uniquement lorsque le réfrigérateur est sous garantie). Le responsable de la surveillance entre les données dans l'outil de collecte de données ODK2.
Temps de reparation | Défauts ou défaillances liés a la performance, à la qualité ou à la sécurité des produits sous garantie | NUM: Heure de fin de l'état de non fonctionnalite moins l'heure de début de la demande de garantie. DEN: N / A ' | Nombre de jours entre la demande de garantie et la réparation | Cacul Numerique | Outil de collecte de données permettant de calculer automatiquement le délai entre la demande de garantie et la réparation