## Follow-up
WHO/PQT/VAX/IMD/PMM Questionnaire V1.0 |  |  |  |  | 2024-06-01
0, avenue Appia – CH-1211 Geneva 27 – Switzerland – Tel central +41 22 791 2111 – Fax central +41 22 791 3111 www.who.int
PMM follow-up questionnaire
Section A: remote or on-site | Core user can answer Section A questions with sentinel’s direction if sentinel is completing PMM sentinel survey remotely. Sentinel should answer if they are on site. Technician should verify answers during TFA visit but does not necessarily need to recomplete the questionnaire. If the failure cause is known after completing Section A, record the component that failed and the failure cause(s) in the Failure Reporting section below. If the failure cause is not known after completing Section A, continue on to Section B, or note that a TFA visit is needed.
Section B: on-site only | Sentinel or technician must complete this section. Sentinel can answer if on site for the PMM sentinel survey; otherwise, this part can be delayed until a technician arrives for TFA and can answer the questions then. If done by a sentinel, then technician should verify answers during TFA visit but does not necessarily need to recomplete the questionnaire. If the failure cause is known after completing Section B, record the component that failed and the failure cause(s) in the Failure Reporting section below. If the failure cause is not known after completing Section B, note that a TFA visit is needed.
[7 empty] | CCE identification | Input |  | Input |  | Input
[7 empty] | Refrigerator ID: | enter here |  | enter here |  | enter here
[7 empty] | Brand and model: | enter here |  | enter here |  | enter here
[7 empty] | Health facility: | enter here |  | enter here |  | enter here
[7 empty] | Failure reporting | Input | Other (please specify) | Input | Other (please specify) | Input | Other (please specify)
[7 empty] | Component: | enter here |  | enter here |  | enter here
[7 empty] | Failure cause #1: | enter here |  | enter here |  | enter here
[7 empty] | Failure cause #2: | enter here |  | enter here |  | enter here
[7 empty] | Information for technician | Input |  | Input |  | Input
[7 empty] | TFA visit needed: | Yes/No |  | Yes/No |  | Yes/No
[7 empty] | Enter any repairs attempted, or concerns for technician: | enter here |  | enter here |  | enter here
[7 empty] | Replacement part(s) needed: | enter here |  | enter here |  | enter here
Section | # | Applicable indicators | Skip logic | Question | Field type | Procedure for sentinel if on site | Procedure for sentinel if remote | Input | Other (please specify) | Input | Other (please specify) | Input | Other (please specify)
A | 1 | temperature alarms |  | Where is the FT2 (or FT2 sensor) placed inside the refrigerator? | list | Ask user where FT2/sensor is typically placed inside of the refrigerator. Also ask if FT2/sensor is always left in the same place or if it has been removed recently for any reason or if it gets removed sometimes to access vaccines. Visually verify placement of FT2/sensor within the refrigerator and note if different than user's answer. If the FT2/sensor has been removed for any reason or gets removed semi-regularly (to access a lower basket for example), select other and provide information. | Ask user where FT2/sensor is typically placed inside of the refrigerator. Also ask if FT2/sensor is always left in the same place or if it has been removed recently for any reason or if it gets removed sometimes to access vaccines. If the FT2/sensor has been removed for any reason or gets removed semi-regularly (to access a lower basket for example), select other and provide information.
A | 2 | temperature alarms |  | Do any temperature alarms appear to have been caused by the removal of FT2 (or FT2 sensor) from refrigerator? | list | Look through the FT2 data for periods of time without readings or periods during which FT2/sensor appears to have been removed from refrigerator (i.e., abrupt or extended periods with readings near ambient temperature). If any such gaps in the data can be identified within the month and align with the timing of one or more alarms, then answer “Yes.” If it is clear that all alarms occurred while the FT2 was in the refrigerator, select “No.” If there are gaps or inconsistencies in the data for which the cause seems unclear, select “Unknown.” | Look through the FT2 data for periods of time without readings or periods during which FT2/sensor appears to have been removed from refrigerator (i.e., abrupt or extended periods with readings near ambient temperature). If any such gaps in the data can be identified within the month and align with the timing of one or more alarms, then answer “Yes.” If it is clear that all alarms occurred while the FT2 was in the refrigerator, select “No.” If there are gaps or inconsistencies in the data for which the cause seems unclear, select “Unknown.”
A | 3 | temperature alarms |  | How often is the refrigerator cleaned? | list | Ask user. | Ask user.
A | 4 | temperature alarms |  | How often is the refrigerator drained? | list | Ask user. | Ask user.
A | 5 | temperature alarms |  | Is the refrigerator in a location exposed to direct sunlight, rainwater, or other harsh conditions at any times or seasons? | list | Ask user and visually verify to the extent possible. Check for roof leaks that may expose the equipment to rainwater. | Ask user.
A | 6 | temperature alarms |  | Does the refrigerator have at least 2 cm of space for ventilation on all sides? | boolean | Use a ruler, if available, or estimate the amount of space on all sides of the refrigerator, including the bottom. Be sure that the electronics compartment vents are not obstructed by anything. | Instruct the user to use a ruler, if available, or estimate the amount of space on all sides of the refrigerator. Ask user to verify that the electronics compartment vents are not obstructed by anything.
A | 7 | temperature alarms |  | What is the current reading on the external temperature display in ˚C (if applicable)? | numeric | Record the temperature reading currently shown on the external display. | Ask the user to read the temperature currently shown on the external display.
A | 8 | temperature alarms |  | Is there currently any water collected in the bottom of the refrigerator? | boolean | Look inside the refrigerator for any water pooling at the bottom of the vaccine compartment. | Instruct user to look inside the refrigerator for any water pooling at the bottom of the vaccine compartment.
A | 9 | temperature alarms | if yes to 8 | What is the height of the water in cm? | numeric | Use a ruler, if available, or estimate the height of the water collected at the bottom of the vaccine compartment. | Instruct the user to use a ruler, if available, or estimate the height of the water collected at the bottom of the vaccine compartment.
A | 10 | temperature alarms |  | Are baskets or shelves used to store vaccines? | boolean | Look inside the vaccine compartment and observe how vaccines are stored. If all vaccines are on a shelf or in a suspended basket, answer yes. Otherwise, answer no. | Instruct the user to look inside the vaccine compartment and observe how vaccines are stored. If all vaccines are on a shelf or in a suspended basket, answer yes. Otherwise, answer no.
A | 11 | temperature alarms; build quality |  | Does the gasket have any visible damage, such as rips, deadhesion, or mold? | boolean | Check the gasket for any visible signs of damage. | Instruct the user to check the gasket for any visible signs of damage.
A | 12 | temperature alarms; build quality | if yes to 11 | Describe observation. | string | Enter details about gasket damage observed. | Ask the user for details about gasket damage observed.
A | 13 | temperature alarms; build quality |  | Is there any corrosion visible on the refrigerator exterior, such as on hinges, hardware, or vents? | boolean | Check the refrigerator exterior for any signs of corrosion, particularly on the hinges, hardware, and vents. | Instruct the user to check the refrigerator exterior for any signs of corrosion, particularly on the hinges, hardware, and vents.
A | 14 | temperature alarms; build quality | if yes to 13 | Describe observation. | string | Enter details about corrosion observed. | Ask the user for details about corrosion observed.
A | 15 | temperature alarms |  | Is the condensation drain leaking, blocked, or missing the plug? | boolean | Check the drain that removes any water that collects inside the vaccine compartment. See if the drain is leaking outside of the refrigerator, or if the drain is blocked by something inside the vaccine compartment that is preventing the removal of water. Also verify that the plug for the drain is in place. | Instruct the user to check the drain that removes any water that collects inside the vaccine compartment. Ask user if the drain is leaking outside of the refrigerator, or if the drain is blocked by something inside the vaccine compartment that is preventing the removal of water. Also ask user to verify that the plug for the drain is in place.
A | 16 | temperature alarms | if yes to 15 | Describe observation. | string | Enter details about condensation drain problems. | Ask the user for details about condensation drain problems.
A | 17 | temperature alarms |  | Does the power cable and/or plug show any signs of damage from rodents, strain, or other causes? | boolean | Check the refrigerator power cable and the plug at the end of this cable for signs of damage from rodents, strain (i.e., the cable is stretched with no slack), or other causes. | Instruct the user to check the refrigerator power cable and the plug at the end of this cable for signs of damage from rodents, strain (i.e., the cable is stretched with no slack), or other causes.
A | 18 | temperature alarms | if yes to 17 | Describe observation. | string | Enter details about power cable/plug damage observed. | Ask the user for details about power cable/plug damage observed.
A | 19 | temperature alarms |  | Has the refrigerator been making any loud noises or producing abnormal smells during the past 30 days? | list | Ask user. | Ask user.
A | 20 | temperature alarms | if yes to 19 | Describe observation. | string | Enter details about noises and/or smells observed during the visit. Ask the user for details about noises and/or smells observed over the past 30 days. | Ask the user for details about noises and/or smells observed over the past 30 days.
A | 21 | temperature alarms | solar only | Where is the solar array located? | list | Ask user. Visually verify location of the solar array. | Ask user.
A | 22 | temperature alarms | solar only | In the past 3 months (or since the last PMM check), has anyone been here to repair or work on the solar system, or has repair or maintenance been requested for any reason? | boolean | Ask user. Verify with repair logs if available. | Ask user.
A | 23 | temperature alarms | if yes to 22 | Why? What system/area? What action was taken? What parts were replaced (if any)? | string | Ask user. Verify with repair logs if available. | Ask user.
A | 24 | temperature alarms | solar only | How often are the solar panels cleaned? | list | Ask user. | Ask user.
B | 25 | build quality |  | What is the main component associated with the first reported build quality issue in the sentinel survey? | list | If at least one build quality issue was reported in the most recent sentinel survey, investigate the first issue that was reported. Read the sentinel survey and ask the user more detailed questions as needed to identify the main component associated with the first build quality issue. If possible, visually verify the component indicated by the user to avoid inadvertently misidentifying components. For example, a user may indicate that pipes within the electronics compartment are corroded, but the sentinel’s visual examination finds that the filter dryer is corroded instead of pipes. | Section B cannot be completed remotely.
B | 26 | build quality |  | What is the main component associated with the second reported build quality issue in the sentinel survey? | list | If at least two build quality issues were reported in the most recent sentinel survey, investigate the second issue that was reported. Read the sentinel survey and ask the user more detailed questions as needed to identify the main component associated with the second build quality issue. If possible, visually verify the component indicated by the user to avoid inadvertently misidentifying components. For example, a user may indicate that pipes within the electronics compartment are corroded, but the sentinel’s visual examination finds that the filter dryer is corroded instead of pipes. | Section B cannot be completed remotely.
B | 27 | build quality |  | What is the main component associated with the third reported build quality issue in the sentinel survey? | list | If at least three build quality issues were reported in the most recent sentinel survey, investigate the third issue that was reported. Read the sentinel survey and ask the user more detailed questions as needed to identify the main component associated with the third build quality issue. If possible, visually verify the component indicated by the user to avoid inadvertently misidentifying components. For example, a user may indicate that pipes within the electronics compartment are corroded, but the sentinel’s visual examination finds that the filter dryer is corroded instead of pipes. | Section B cannot be completed remotely.
B | 28 | temperature alarms; build quality |  | Are the door hinges loose or misaligned? | boolean | Check the hinges to see if they are loose and if the door has any play due to the hinges. Use a ruler to measure the distance between the refrigerator door edge and the refrigerator cabinet edge in two places (i.e. one measurement near the hinged side of the door and one near the free side of the door). The difference between these two measurements is the door alignment gap. If the door alignment gap is zero, the door is not misaligned. If the door alignment gap is nonzero, provide the gap measurement in the following question. | Section B cannot be completed remotely.
B | 29 | temperature alarms; build quality | if yes to 28 | Door alignment gap (in mm)? | numeric | See procedure for question 28. | Section B cannot be completed remotely.
B | 30 | temperature alarms |  | Does the access cover show signs of tampering? | list | Check the access cover for signs of tampering (unauthorized access). | Section B cannot be completed remotely.
B | 31 | temperature alarms |  | Does the power LED on the refrigerator turn on when the equipment is receiving power (if applicable)? | list | Check power LED when equipment is receiving power – for AC systems, any time when power is not out; for DC systems, check during the day when solar system is receiving power. | Section B cannot be completed remotely.
B | 32 | temperature alarms |  | Is the ice liner intact (i.e. not missing any water packs and appears to be full) and in the correct place? | list | An ice liner is an internal lining within the refrigerator surrounding the storage that is filled with ice, cold water, or other coolant. Some equipment has readily accessible ice liners while some does not. If there is a user manual available or any user instructions located on the refrigerator that have to do with the ice liner, these should be consulted for information on the ice liner. Even if the ice liner is accessible, it can be disruptive to access it. With experience this may become easier, but it may be difficult to answer this question effectively if operating remotely. Some ice liners consist of water packs. For example, many B Medical refrigerators have panels of approximately 6 water packs that can be accessed by removing a square plastic cover. For this type of design, if the panels are missing, any water packs are missing, or any water packs are not full, the ice liner is not intact. If a panel has been moved from its intended position, the ice liner is not in the correct place. Some ice liners have a user-accessible reservoir marked with a visible fill line for water. If this is the case, check that the height of water or ice meets the fill line. If it does not, the ice liner is not intact. Some other equipment may have completely sealed and inaccessible ice-liners, and some may require the use of tools to access. For this effort, these liners should be left alone and “unknown” selected. | Section B cannot be completed remotely.
B | 33 | temperature alarms |  | What is the thermostat setpoint? | string | Record the temperature that the thermostat is currently set to. If a manual is available it may have details on checking the set point. If the thermostat is located in the electronics compartment, access the electronics compartment by removing the access cover/vent with appropriate tools. If the sentinel and user are not experienced enough with the specific equipment to do this, enter “unknown.” In some models, the thermostat may be contained in the same component as the external display and can be accessed with certain keypad sequences. If it cannot be accessed by the sentinel, enter “unknown” and the technician will need to answer during TFA. Some equipment may have a dial thermostat control that does not have a numeric set point or has an arbitrary scale of settings, for example, between 1 and 7. If the setpoint is not a temperature, enter “unknown” and also note the current dial setting. | Section B cannot be completed remotely.
B | 34 | temperature alarms |  | Are there any signs of corrosion, burns, shorts, or heavily accumulated dirt in the electronics compartment? | boolean | Access the electronics compartment by removing the access cover/vent with appropriate tools. Look for any signs of corrosion, burns, shorts, or a large amount of dirt that may be hindering performance of refrigerator components. In particular, check the piping, filter dryer, capillaries, fan, and compressor for corrosion, burns, damage, etc. | Section B cannot be completed remotely.
B | 35 | temperature alarms | if yes to 34 | Describe the problem, location/component(s), and severity. | string | Enter details about the problem observed. | Section B cannot be completed remotely.
B | 36 | temperature alarms |  | Are there any signs of leaks in the electronics compartment? | boolean | Access the electronics compartment by removing the access cover/vent with appropriate tools. Check pipes, capillaries, and the cable seal around the thermostat sensor where it enters the vaccine compartment. Look for signs of leaks, moisture, etc. | Section B cannot be completed remotely.
B | 37 | temperature alarms | if yes to 36 | Describe the problem, location/component(s), and severity. | string | Enter details about the leak(s) observed. | Section B cannot be completed remotely.
B | 38 | temperature alarms |  | Is the fan in the electronics compartment working? | boolean | Access the electronics compartment by removing the access cover/vent with appropriate tools. Observe if the fan is rotating to provide airflow when the compressor is running. If not, the fan is not working. If the compressor is not running, open the refrigerator door and wait 10 minutes to get the compressor running, and then check if the fan is working. | Section B cannot be completed remotely.
B | 39 | temperature alarms |  | Are there any indications of illicit repairs in the electronics compartment? | boolean | Access the electronics compartment by removing the access cover/vent with appropriate tools. Check for indications of illicit repairs, such as snipped wires, wires or connections that appear to have been added post-manufacturing, incorrect or added components, etc. If available, a circuitry diagram for the refrigerator may be helpful in identifying illicit repairs. | Section B cannot be completed remotely.
B | 40 | temperature alarms | if yes to 39 | Describe the observation and the location/component(s) involved. | string | Enter details about the repair(s) observed. | Section B cannot be completed remotely.
B | 41 | voltage stabilizer |  | Rated operating range of voltage stabilizer (volts)? | string | Check the manufacturer label, other markings on the power quality device, and/or the device manual for the range of acceptable input voltage. If unknown, leave blank. | Section B cannot be completed remotely.
B | 42 | voltage stabilizer |  | What is the input voltage reading (volts)? | numeric | This question only applies to voltage stabilizers that display the input voltage they are currently receiving from the power source. Record the input voltage displayed (in volts). | Section B cannot be completed remotely.
B | 43 | voltage stabilizer |  | What is the output voltage reading (volts)? | numeric | This question only applies to voltage stabilizers that display the output voltage they are currently delivering to the refrigerator. Record the output voltage displayed (in volts). | Section B cannot be completed remotely.
B | 44 | voltage stabilizer |  | If the voltage stabilizer was replaced within the month, what is the make and model of the new device? | string | Find the make and model labeled on the voltage stabilizer. If the voltage stabilizer was not replaced within the month, enter "N/A." If the voltage stabilizer was replaced within the month but the make and model cannot be identified, enter "unknown." | Section B cannot be completed remotely.
B | 45 | voltage stabilizer |  | If the voltage stabilizer was replaced within the month, is it the same make and model that was previously used with the refrigerator? | list | Reference inventory data collected at the start of PMM to determine if the new voltage stabilizer is of the same make and model as the stabilizer previously used with the refrigerator. | Section B cannot be completed remotely.
B | 46 | voltage stabilizer |  | If the voltage stabilizer was not replaced within the month, was the voltage stabilizer provided with the refrigerator or purchased separately? | list | Ask the user if the voltage stabilizer was delivered with the refrigerator when the refrigerator was first installed, or if the device was purchased separately from the refrigerator by health facility or district/etc. | Section B cannot be completed remotely.
B | 47 | voltage stabilizer |  | Is there any visible physical damage to the voltage stabilizer? | boolean | Check the voltage stabilizer for any signs of damage, such as corrosion or dents. | Section B cannot be completed remotely.
B | 48 | voltage stabilizer |  | Are there any problems with connections to the voltage stabilizer and/or the power cable? | list | Check the connections of the voltage stabilizer to the refrigerator power cable and the power source (i.e. power outlet). Are any connections loose? Are any cables damaged? Do any plugs or receptacles appear damaged (burns, dents, etc.)? Are the refrigerator plug and power quality device receptacle mismatched (i.e. plug types not compatible)? Any other apparent issues? | Section B cannot be completed remotely.
B | 49 | voltage stabilizer |  | If a voltage stabilizer for the refrigerator is available on site but is not currently in use, should this device be put into use with the refrigerator? | boolean | Ask the user if a voltage stabilizer was delivered with or purchased for this refrigerator. If a voltage stabilizer for the refrigerator is available on site but not being used, ask the user why it is not in use. The answer may indicate that the power quality device is damaged or the receptacle on the device is not compatible with the refrigerator power cable (in which case, answer no). If the answer indicates that the power quality device could be put into use safely, answer yes and connect the refrigerator power cable to the power quality device, plug in the power quality device, and turn it on (if needed). Some voltage stabilizers may have a delay before starting after plugged in (~5-10 mins). | Section B cannot be completed remotely.
B | 50 | temperature alarms | solar only | Is the solar system grounded? | boolean | Check for a cable connection from the solar array to a ground rod that is submerged in the earth within a ground pit. If such connection and ground rod exist, then the solar system is grounded. Otherwise, answer no. | Section B cannot be completed remotely.

## Data validation
Question | Component | Failure cause #1 | Failure cause #2 | TFA | 1 | 2 | 3 | 21 | 25 | 26 | 27 | 30 | 32 | 46 | 48
| basket | break | abuse | Yes | bottom of vaccine compartment | yes | weekly | ground | Not applicable – no build quality issues reported | Not applicable – no second build quality issue reported | Not applicable – no third build quality issue reported | yes | Yes, the ice liner is intact and in the correct place | provided with refrigerator | loose connection
| battery | build quality | factory concealed | No | walls of vaccine compartment | no | monthly | roof | Casing | Casing | Casing | no | No, the ice liner is not intact and/or not in the correct place | purchased separately | cable damage
| battery terminal | corrosion | factory observable |  | on shelf or basket | unknown | yearly | pole | Base | Base | Base | unable to determine | This refrigerator is not designed to operate with an ice liner | unable to determine | plug/receptacle damage
| battery voltmeter | damage | impact |  | inside of door |  | never | other (please specify) | Access cover | Access cover | Access cover |  | Unknown/no information |  | refrigerator plug type and voltage stabilizer receptacle type mismatched
| burner/boiler | degradation | insect |  | other (please specify) |  | other (please specify) |  | Vents | Vents | Vents |  |  |  | other (please specify)
| cabinet | disintegration | installation | [5 empty] | Condensation drainage | Condensation drainage | Condensation drainage
| capacitor | hot spot | misuse | [5 empty] | Power cable(s) | Power cable(s) | Power cable(s)
| capillary tube | incomplete | rodent | [5 empty] | Handle | Handle | Handle
| circuit breaker | insufficient capacity | shipping handling | [5 empty] | Gasket | Gasket | Gasket
| combiner | lack of maintenance | shipping sea freight | [5 empty] | Alignment | Alignment | Alignment
| compressor | leak | shipping trucking | [5 empty] | Hinges | Hinges | Hinges
| compressor electronic unit | missing | shipping storage | [5 empty] | Latch/clasp | Latch/clasp | Latch/clasp
| condenser | misplacement | sunlight | [5 empty] | Lock/key | Lock/key | Lock/key
| control panel | open circuit | water | [5 empty] | Keypad | Keypad | Keypad
| coupler system | orientation | other (please specify) | [5 empty] | External temperature display | External temperature display | External temperature display
| display | outage | unknown | [5 empty] | Power switch | Power switch | Power switch
| door | power quality | [6 empty] | Power LED | Power LED | Power LED
| drier | power tampering | [6 empty] | Ice liner/water packs | Ice liner/water packs | Ice liner/water packs
| energy harvest control | shading | [6 empty] | Shelving/baskets | Shelving/baskets | Shelving/baskets
| evaporator | short circuit | [6 empty] | Vaccine compartment walls | Vaccine compartment walls | Vaccine compartment walls
| fan | soiling | [6 empty] | Vaccine compartment bottom | Vaccine compartment bottom | Vaccine compartment bottom
| firmware | tampering | [6 empty] | Grate | Grate | Grate
| flue | unauthorized use | [6 empty] | Drain | Drain | Drain
| flue baffle | ultraviolet degradation | [6 empty] | Thermostat sensor/external display sensor | Thermostat sensor/external display sensor | Thermostat sensor/external display sensor
| freezer compartment | voltage (low) | [6 empty] | Sensor mount or seal | Sensor mount or seal | Sensor mount or seal
| fuse | other (please specify) | [6 empty] | Base plate | Base plate | Base plate
| gasket | unknown | [6 empty] | Fan | Fan | Fan
| handle | [7 empty] | Compressor | Compressor | Compressor
| heater | [7 empty] | Compressor controller/starting device | Compressor controller/starting device | Compressor controller/starting device
| hinge | [7 empty] | Piping | Piping | Piping
| hinge cover | [7 empty] | Thermostat controller | Thermostat controller | Thermostat controller
| holdover gauge | [7 empty] | Internal condenser | Internal condenser | Internal condenser
| indicator light | [7 empty] | External condenser | External condenser | External condenser
| interconnect (electrical) | [7 empty] | Filter dryer | Filter dryer | Filter dryer
| lamp (absorption) | [7 empty] | Capillary | Capillary | Capillary
| lid | [7 empty] | Starting capacitor | Starting capacitor | Starting capacitor
| monitoring device | [7 empty] | Solar panels | Solar panels | Solar panels
| mounting hardware | [7 empty] | Solar array support structure/mounting | Solar array support structure/mounting | Solar array support structure/mounting
| on/off switch | [7 empty] | Solar battery | Solar battery | Solar battery
| phase change material (PCM) | [7 empty] | Solar wiring | Solar wiring | Solar wiring
piping
power adapter
power cable
power cable connector
remote temperature monitoring device (RTMD)
refrigerant
removable insulation
safety valve
seal (sealant)
sensor
SD card
shelf
SIM card
software
solar array
solar array cable
solar cell
solar module
solar support structure
solar power system
starter relay
status indicator - audible alarm
status indicator - autonomy gauge
status indicator - door opening
status indicator - holdover gauge
status indicator - LED
status indicator - voltage
theft deterrent fastener
thermal storage
thermocouple
thermometer
thermostat
thermostat control card
thermostat sensor lead
thermostat wiring
30-DTR
transformer
ventilation grill
voltage stabilizer
water pack
wicks
wiring
wiring connections
wiring terminals
other (please specify)
unknown